## 1_Raw_Training
Concentration (pg/mL) | n_obs (beads) | AEB_raw | k_train (positive beads) | P_train | mu_train (TCS) | mu_train_4pl
0 | 29568 | 0.0029 | 85.623 | 0.0028958 | 0 | 0
0 | 33529 | 0.0037 | 123.8281 | 0.0036932 | 0 | 0
0.0328 | 31110 | 0.0255 | 783.2758 | 0.025178 | 0.18992 | 0.15194
0.0328 | 30257 | 0.0268 | 800.1181 | 0.026444 | 0.18992 | 0.15194
0.0984 | 27297 | 0.0655 | 1730.6558 | 0.063401 | 0.56977 | 0.45582
0.0984 | 33491 | 0.0687 | 2223.5773 | 0.066393 | 0.56977 | 0.45582
0.296 | 23171 | 0.1807 | 3830.497 | 0.16531 | 1.714 | 1.3712
0.296 | 35706 | 0.197 | 6384.5672 | 0.17881 | 1.714 | 1.3712
0.888 | 31713 | 0.5426 | 13280.2921 | 0.41876 | 5.1419 | 4.1135
0.888 | 30509 | 0.5826 | 13471.4161 | 0.44156 | 5.1419 | 4.1135
2.664 | 31098 | 1.6441 | 25090.2908 | 0.80681 | 15.4256 | 12.3405
2.664 | 30145 | 1.7458 | 24884.5369 | 0.82549 | 15.4256 | 12.3405
8 | 22988 | 4.4611 | 22722.4966 | 0.98845 | 46.3231 | 37.0585
8 | 22866 | 4.8785 | 22692.0259 | 0.99239 | 46.3231 | 37.0585
24 | 28320 | 13.2759 | 28319.9514 | 1 | 138.9692 | 111.1754
24 | 28760 | 12.3785 | 28759.879 | 1 | 138.9692 | 111.1754

## 2_Raw_Test
Test sample # | test_n (beads) | test_AEB_raw | test_k (positive beads)
1 | 29071 | 0.2366 | 6125
2 | 33328 | 0.27244 | 7948
3 | 28472 | 2.6172 | 26393.3188
4 | 28456 | 2.5644 | 26265.9644
5 | 20960 | 0.058197 | 1185
6 | 24727 | 0.058413 | 1403
7 | 25139 | 0.35136 | 7448
8 | 27922 | 0.36993 | 8634
9 | 22721 | 0.10277 | 2219
10 | 22127 | 0.10608 | 2227
11 | 24019 | 0.12617 | 2847
12 | 20192 | 0.12558 | 2383
13 | 22124 | 0.33779 | 6342
14 | 21902 | 0.34098 | 6328
15 | 25262 | 0.093416 | 2253
16 | 27985 | 0.09299 | 2485
17 | 26420 | 2.396 | 24013.6774
18 | 23602 | 2.2728 | 21170.4048
19 | 27587 | 0.33734 | 7899
20 | 26023 | 0.34979 | 7681
21
22
23
24
25 | 18699 | 16.7462 | 18698.999
26 | 18713 | 16.345 | 18712.9985
27 | 23020 | 0.41737 | 7855
28 | 24709 | 0.42142 | 8497
29 | 16844 | 18.3559 | 16843.9998
30 | 16167 | 18.7763 | 16166.9999
31 | 20085 | 3.4546 | 19450.3166
32 | 20342 | 3.5734 | 19771.1731
33 | 19173 | 5.8473 | 19117.6329
34 | 19937 | 6.32 | 19901.1147
35 | 23344 | 9.983 | 23342.922
36 | 24012 | 9.974 | 24010.8811
37 | 17073 | 6.9533 | 17056.6866
38 | 17136 | 6.8398 | 17117.6587
39 | 24776 | 2.0099 | 21455.8667
40 | 21719 | 2.2728 | 19481.5269
41 | 24918 | 11.4706 | 24917.74
42 | 24142 | 11.6721 | 24141.7941
43 | 24784 | 3.9126 | 24288.5909
44 | 22893 | 4.225 | 22558.1978
45 | 23404 | 8.4397 | 23398.9419
46 | 26483 | 9.3405 | 26480.675
47 | 23871 | 7.7491 | 23860.7081
48 | 16792 | 7.7964 | 16785.0946
49 | 24276 | 0.72913 | 12567
50 | 31450 | 0.74396 | 16504
51 | 22859 | 1.0918 | 15187
52 | 22096 | 1.159 | 15162
53 | 20070 | 0.38103 | 6359
54 | 21347 | 0.36036 | 6459
55 | 21735 | 0.30048 | 5641
56 | 28879 | 0.3094 | 7685
57 | 26687 | 10.0901 | 26685.8928
58 | 31241 | 9.986 | 31239.5617
59 | 25211 | 3.6447 | 24552.2522
60 | 28192 | 3.8179 | 27572.521
61 | 30914 | 1.3125 | 22593.7576
62 | 31997 | 1.2657 | 22972.7534
63 | 30768 | 0.49113 | 11940
64 | 29375 | 0.48994 | 11378
65 | 19332 | 0.16466 | 2935
66 | 33952 | 0.16166 | 5068
67 | 27176 | 0.060328 | 1591
68 | 25337 | 0.061404 | 1509
69 | 28955 | 0.022457 | 643
70 | 29381 | 0.02105 | 612
71 | 27738 | 0.004046 | 112
72 | 31775 | 0.0039417 | 125

## 3_Training_Predictions
Concentration (pg/mL) | AEB_observed | AEB_pred_4PL | resid_4PL | P_pred_TCS | AEB_pred_TCS | resid_TCS
0 | 0.0029 | 0.0033 | -0.0004 | 0.003712 | 0.0037189 | -0.00081889
0 | 0.0037 | 0.0033 | 0.0004 | 0.003712 | 0.0037189 | -1.8891e-05
0.0328 | 0.0255 | 0.024292 | 0.0012079 | 0.019984 | 0.024304 | 0.0011964
0.0328 | 0.0268 | 0.024292 | 0.0025079 | 0.019984 | 0.024304 | 0.0024964
0.0984 | 0.0655 | 0.066399 | -0.00089901 | 0.051736 | 0.065473 | 2.712e-05
0.0984 | 0.0687 | 0.066399 | 0.002301 | 0.051736 | 0.065473 | 0.0032271
0.296 | 0.1807 | 0.19316 | -0.012462 | 0.1413 | 0.18948 | -0.0087829
0.296 | 0.197 | 0.19316 | 0.0038377 | 0.1413 | 0.18948 | 0.0075171
0.888 | 0.5426 | 0.5702 | -0.027603 | 0.36209 | 0.56101 | -0.018411
0.888 | 0.5826 | 0.5702 | 0.012397 | 0.36209 | 0.56101 | 0.021589
2.664 | 1.6441 | 1.6712 | -0.027138 | 0.73847 | 1.6756 | -0.031495
2.664 | 1.7458 | 1.6712 | 0.074562 | 0.73847 | 1.6756 | 0.070205
8 | 4.4611 | 4.714 | -0.25289 | 0.98205 | 5.0244 | -0.56327
8 | 4.8785 | 4.714 | 0.16451 | 0.98205 | 5.0244 | -0.14587
24 | 13.2759 | 11.9439 | 1.332 | 0.99999 | 15.0657 | -1.7898
24 | 12.3785 | 11.9439 | 0.43463 | 0.99999 | 15.0657 | -2.6872

## 4_Fit_Parameters
Parameter | Value | Note
A_4pl (fixed from blank) | 0.0033 | Lower asymptote
B_4pl (slope) | 1.0025
C_4pl (midpoint) | 359.2456 | pg/mL × scale
D_4pl (upper asymptote) | 50.6425
κ_tcs (depletion) | 8.2265 | Mid-conc identifiable
b_tcs (background) | 0.003712 | Non-specific binding
Mw (Da) | 26000 | IL-6 molecular weight
V (assay volume) | 0.000125 | Liters
N_total (beads) | 500000 | TCS normalization

## 5_Bootstrap_CI
Model | Parameter | Point | 2.5% | 97.5% | Width | Relative width
4PL | B | 1.0025 | 0.98581 | 1.0209 | 0.035071 | 0.034982
4PL | C | 359.2456 | 358.6588 | 359.7666 | 1.1078 | 0.0030836
4PL | D | 50.6425 | 46.7348 | 55.0548 | 8.32 | 0.16429
TCS | kappa | 8.2265 | 7.9763 | 8.5836 | 0.60733 | 0.073825
TCS | b | 0.003712 | 0.0030839 | 0.0068356 | 0.0037516 | 1.0107
4PL | ρ(B, C) | -0.067302
4PL | ρ(B, D) | 0.36729
4PL | ρ(C, D) | 0.88922
TCS | ρ(κ, b) | 0.35292

## 6_Model_Comparison
Metric | 4PL | TCS | Winner / Note
logL | -124799.5032 | -124824.2761 | Higher = better fit
AICc | 249607.0064 | 249653.4752 | Lower = better (penalizes complexity)
BIC | 249607.3242 | 249654.0973 | Lower = better
R² (weighted, AEB scale) | 0.99842 | 0.9977 | Higher = better
Shapiro-Wilk p (residuals) | 0.62027 | 0.88825 | Higher = more normal

## 7_LOO_Stats
Metric | Value | Note
4PL LOO failures | 0 | out of 16 wells
TCS LOO failures | 0 | out of 16 wells
LOO total logL (4PL) | -124867.0285 | Sum across wells
LOO total logL (TCS) | -124854.6952 | Sum across wells
Mean diff (4PL - TCS) | -0.77083 | Positive = TCS better
SE of diff | 1.9813
t-statistic | -0.38906 | Paired t-test LOO
TOST margin (±logL) | 78.0342 | 1% of mean per-fold ¦logL¦
Mean diff 90% CI | [-4.244, 2.702]
TOST p | 1.1102e-16 | EQUIVALENT
Smallest passing margin (logL) | 4.2441 | 0.054% of per-fold ¦logL¦

## 8_Clinical_Reference
Reference | Value (pg/mL) | Source
Manufacturer LOD | 0.0055 | Quoted by manufacturer
Analytical LLOQ | 0.01 | Quoted by manufacturer
Functional LLOQ (4× LOD) | 0.04 | Quoted by manufacturer
Plasma lower bound (normal) | 0.607 | Reference range
Plasma mean (normal) | 1.79 | Reference range
Plasma upper bound (normal) | 4.64 | Reference range
Highest calibrator | 24 | Standard curve

## 18_Clinical_BackCalc
Clinical validation: back-calculation from test_AEB (34 samples, QC excluded)
Dilution: first 26 samples = 4x, last 8 = 1x; all × 1.25 (100→125 μL)
Sample | AEB_mean | P_pos | Dilution | 4PL (pg/mL) | R2 (pg/mL) | R3 (pg/mL) | R1 (pg/mL) | Roche (pg/mL)
1 | 0.0583 | 0.056638 | 4 | 0.3432 | 0.4349 | 0.4232 | 0.4261 | 0
2 | 0.3606 | 0.30278 | 4 | 2.2323 | 2.8437 | 2.3915 | 2.7914 | 0
3 | 0.1044 | 0.099156 | 4 | 0.6305 | 0.8023 | 0.7632 | 0.7853 | 1.87
4 | 0.1259 | 0.11827 | 4 | 0.7642 | 0.9732 | 0.9161 | 0.9525 | 2.1
5 | 0.3394 | 0.28779 | 4 | 2.0989 | 2.6743 | 2.2717 | 2.6242 | 6.63
6 | 0.0932 | 0.088991 | 4 | 0.5606 | 0.7129 | 0.682 | 0.6978 | 7.5
7 | 2.3344 | 0.90313 | 4 | 15.084 | 18.5688 | 7.1925 | 18.8719 | 51.4
8 | 0.3436 | 0.29076 | 4 | 2.1251 | 2.7076 | 2.2955 | 2.657 | 7
9 |  |  | 4 | NaN | NaN | NaN | NaN | 1865
10 |  |  | 4 | NaN | NaN | NaN | NaN | 1037
11 | 16.5456 | 1 | 4 | 150.7742 | NaN | 7.9671 | 179.2076 | 558.8
12 | 0.4194 | 0.34255 | 4 | 2.6013 | 3.3117 | 2.7097 | 3.2539 | 9.93
13 | 18.5661 | 1 | 4 | 179.7675 | NaN | 7.9671 | 210.5955 | 532.8
14 | 3.514 | 0.97022 | 4 | 23.2598 | 27.9666 | 7.729 | 29.0466 | 66.82
15 | 6.0836 | 0.99772 | 4 | 42.5431 | 48.4393 | 7.9489 | 52.8337 | 123.6
16 | 9.9785 | 0.99995 | 4 | 76.3666 | 79.4699 | 7.9667 | 93.7366 | 264.7
17 | 6.8965 | 0.99899 | 4 | 49.1091 | 54.9156 | 7.959 | 60.86 | 150.9
18 | 2.1414 | 0.8825 | 4 | 13.7829 | 17.0308 | 7.0275 | 17.2482 | 45.37
19 | 11.5714 | 0.99999 | 4 | 92.1278 | 92.1604 | 7.967 | 112.3923 | 327.6
20 | 4.0688 | 0.9829 | 4 | 27.2458 | 32.387 | 7.8304 | 33.9886 | 95.15
21 | 8.8901 | 0.99986 | 4 | 66.2846 | 70.7987 | 7.966 | 81.663 | 283
22 | 7.7727 | 0.99958 | 4 | 56.4626 | 61.8963 | 7.9637 | 69.8015 | 178.2
23 | 0.7365 | 0.52123 | 4 | 4.6065 | 5.8385 | 4.1385 | 5.7673 | 14.79
24 | 1.1254 | 0.67547 | 4 | 7.0968 | 8.9363 | 5.3719 | 8.8874 | 29.54
25 | 0.3707 | 0.30974 | 4 | 2.2953 | 2.9237 | 2.4473 | 2.8704 | 6.1
26 | 0.3049 | 0.26283 | 4 | 1.883 | 2.3998 | 2.0721 | 2.3537 | 4.19
27 | 10.0381 | 0.99996 | 1 | 19.2335 | 19.9862 | 1.9917 | 23.6032 | 50
28 | 3.7313 | 0.97604 | 1 | 6.2025 | 7.4245 | 1.9439 | 7.7425 | 16.67
29 | 1.2891 | 0.72449 | 1 | 2.0392 | 2.5602 | 1.441 | 2.5536 | 5.56
30 | 0.4905 | 0.3877 | 1 | 0.7623 | 0.9696 | 0.7677 | 0.9538 | 1.85
31 | 0.1632 | 0.15055 | 1 | 0.2492 | 0.3176 | 0.2936 | 0.3109 | 0.62
32 | 0.0609 | 0.059051 | 1 | 0.0898 | 0.1138 | 0.1106 | 0.1115 | 0.21
33 | 0.0218 | 0.021519 | 1 | 0.0288 | 0.0359 | 0.0356 | 0.0355 | 0.07
34 | 0.004 | 0.003986 | 1 | 0.0011 | 0.0005 | 0.0005 | 0.001 | 0
Summary Statistics
Comparison | n | Slope | 95% CI | R² | Spearman ρ | BA bias % | BA LoA lower % | BA LoA upper %
R2 vs 4PL | 30 | 1.0459 | 1.0222-1.0696 | 0.9963 | 1 | 18.83 | 4.21 | 33.44
4PL vs Roche | 29 | 0.2937 |  | 0.9806 | 0.9788 | -102.73 | -138.55 | -66.91
R2 vs Roche | 27 | 0.2885 |  | 0.9716 | 0.9744 | -88.15 | -132.67 | -43.63
R1 vs Roche | 29 | 0.3497 |  | 0.9829 | 0.9778 | -86.27 | -126.97 | -45.58

## 9_TCS_PerPoint
TCS quantification per training well — M_hat with 3 CIs (Delta / Exact / LR)
Conc (pg/mL) | k | n | P_obs | M_hat | CV (%) | Delta_low | Delta_up | Exact_low | Exact_up | LR_low | LR_up
0.0328 | 783.2758 | 31110 | 0.025178 | 100481.8185 | 4.1833 | 92243.2798 | 108720.3572 | 92384.6719 | 109017.1504 | 92434.4798 | 108913.8034
0.0328 | 800.1181 | 30257 | 0.026444 | 106479.0505 | 4.105 | 97912.0987 | 115046.0023 | 98057.4678 | 115351.317 | 98108.8308 | 115245.0212
0.0984 | 1730.6558 | 27297 | 0.063401 | 285012.0342 | 2.5489 | 270773.3726 | 299250.6957 | 270934.1229 | 299594.3005 | 270996.2151 | 299475.3016
0.0984 | 2223.5773 | 33491 | 0.066393 | 299774.4514 | 2.2425 | 286598.8148 | 312950.0881 | 286729.7398 | 313230.2831 | 286780.8298 | 313133.3736
0.296 | 3830.497 | 23171 | 0.16531 | 816459.0553 | 1.6519 | 790024.2033 | 842893.9073 | 790212.9603 | 843322.8654 | 790303.6258 | 843174.8508
0.296 | 6384.5672 | 35706 | 0.17881 | 891655.1182 | 1.2777 | 869326.6912 | 913983.5451 | 869448.9716 | 914263.9829 | 869509.7545 | 914167.3838
0.888 | 13280.2921 | 31713 | 0.41876 | 2485996.1815 | 0.8845 | 2442899.2681 | 2529093.095 | 2443035.4879 | 2529479.9618 | 2443152.1896 | 2529346.9181
0.888 | 13471.4161 | 30509 | 0.44156 | 2670526.3972 | 0.87943 | 2624496.0528 | 2716556.7415 | 2624637.4639 | 2716969.2734 | 2624765.7298 | 2716827.4023
2.664 | 25090.2908 | 31098 | 0.80681 | 7567496.9954 | 0.70646 | 7462714.7366 | 7672279.2542 | 7462844.4696 | 7673176.4839 | 7463299.4083 | 7672867.8203
2.664 | 24884.5369 | 30145 | 0.82549 | 8036665.0687 | 0.71908 | 7923398.7632 | 8149931.3743 | 7923531.3389 | 8150940.4388 | 7924055.8409 | 8150593.2167
8 | 22722.4966 | 22988 | 0.98845 | 20563037.4336 | 1.3689 | 20011343.7893 | 21114731.0779 | 20011600.4154 | 21132372.6578 | 20022378.5906 | 21126209.4496
8 | 22692.0259 | 22866 | 0.99239 | 22488610.2341 | 1.5493 | 21805713.8074 | 23171506.6607 | 21806180.0825 | 23198524.4418 | 21822409.586 | 23189040.3525
24 | 28319.9514 | 28320 | 1 | 61227961.0596 | 34.185 | 20204372.469 | 102251549.6502 | 41128235.8322 | 127469059.1934 | 43736184.0207
24 | 28759.879 | 28760 | 1 | 57088025.6712 | 23.2287 | 31097362.7693 | 83078688.5732 | 41022053.9947 | 127469059.1934 | 43300722.4242

## 10_TCS_TestSet
TCS quantification on the 68-sample test set (CV + CI)
Index | AEB | n | k | P_obs | M_hat | CV (%) | Delta_low | Delta_up | Exact_low | Exact_up | LR_low | LR_up
0 | 0.2366 | 29071 | 6125 | 0.21069 | 1074328.4247 | 1.3012 | 1046930.0277 | 1101726.8218 | 1047080.1337 | 1102078.9585 | 1047159.8416 | 1101957.6441
1 | 0.27244 | 33328 | 7948 | 0.23848 | 1239662.5987 | 1.1407 | 1211946.3734 | 1267378.824 | 1212077.1167 | 1267692.0516 | 1212150.9702 | 1267584.2206
2 | 2.6172 | 28472 | 26393.3188 | 0.92699 | 12056598.1836 | 0.80803 | 11865656.9772 | 12247539.39 | 11865781.4656 | 12249881.9512 | 11867171.815 | 12249073.9444
3 | 2.5644 | 28456 | 26265.9644 | 0.92304 | 11813266.0676 | 0.80172 | 11627639.6657 | 11998892.4695 | 11627765.3855 | 12001123.4662 | 11629083.0647 | 12000354.0737
4 | 0.058197 | 20960 | 1185 | 0.056536 | 251323.0262 | 3.1037 | 236034.667 | 266611.3855 | 236244.2742 | 267058.1069 | 236323.552 | 266903.0333
5 | 0.058413 | 24727 | 1403 | 0.05674 | 252317.3904 | 2.8517 | 238214.8551 | 266419.9257 | 238392.4291 | 266798.27 | 238459.8537 | 266667.0768
6 | 0.35136 | 25139 | 7448 | 0.29627 | 1603777.468 | 1.1772 | 1566775.4651 | 1640779.4709 | 1566948.5578 | 1641214.2098 | 1567059.3139 | 1641064.5121
7 | 0.36993 | 27922 | 8634 | 0.30922 | 1689433.5293 | 1.0933 | 1653230.5134 | 1725636.5452 | 1653386.205 | 1726032.1275 | 1653488.9154 | 1725895.966
8 | 0.10277 | 22721 | 2219 | 0.097663 | 456934.984 | 2.2035 | 437200.6532 | 476669.3149 | 437393.6088 | 477089.7415 | 437473.7477 | 476944.3275
9 | 0.10608 | 22127 | 2227 | 0.10065 | 472212.6373 | 2.1971 | 451878.3998 | 492546.8749 | 452076.53 | 492979.3351 | 452159.3158 | 492829.7615
10 | 0.12617 | 24019 | 2847 | 0.11853 | 564878.0789 | 1.9324 | 543484.1317 | 586272.026 | 543666.4918 | 586674.2774 | 543745.788 | 586535.3116
11 | 0.12558 | 20192 | 2383 | 0.11802 | 562188.1312 | 2.1124 | 538912.1405 | 585464.1218 | 539129.1876 | 585942.8084 | 539223.2013 | 585777.3002
12 | 0.33779 | 22124 | 6342 | 0.28666 | 1541168.917 | 1.2757 | 1502633.8726 | 1579703.9613 | 1502830.7297 | 1580194.2332 | 1502953.8004 | 1580025.3353
13 | 0.34098 | 21902 | 6328 | 0.28892 | 1555849.0245 | 1.2771 | 1516904.46 | 1594793.589 | 1517103.3022 | 1595289.7573 | 1517228.2297 | 1595118.8264
14 | 0.093416 | 25262 | 2253 | 0.089185 | 413794.9659 | 2.1949 | 395993.595 | 431596.3367 | 396167.1416 | 431972.5964 | 396238.0262 | 431842.4689
15 | 0.09299 | 27985 | 2485 | 0.088798 | 411831.3475 | 2.0903 | 394958.591 | 428704.1041 | 395115.2065 | 429043.5363 | 395179.2266 | 428926.2012
16 | 2.396 | 26420 | 24013.6774 | 0.90892 | 11036303.3799 | 0.8124 | 10860575.1268 | 11212031.633 | 10860715.1219 | 11214087.6119 | 10861906.5333 | 11213378.7671
17 | 2.2728 | 23602 | 21170.4048 | 0.89698 | 10467776.7004 | 0.84644 | 10294116.5661 | 10641436.8347 | 10294277.2299 | 10643493.5032 | 10295448.7146 | 10642784.4054
18 | 0.33734 | 27587 | 7899 | 0.28633 | 1539057.8544 | 1.1431 | 1504576.1592 | 1573539.5495 | 1504733.9042 | 1573932.4056 | 1504832.7245 | 1573797.1535
19 | 0.34979 | 26023 | 7681 | 0.29516 | 1596501.723 | 1.1592 | 1560230.4794 | 1632772.9666 | 1560397.6782 | 1633192.5156 | 1560504.4415 | 1633048.061
20 | [10 empty] | 0
21 | [10 empty] | 0
22 | [10 empty] | 0
23 | [10 empty] | 0
24 | 16.7462 | 18699 | 18698.999 | 1 | 77237414.2236 | 189.0885 | -209009588.7025 | 363484417.1496 | 39335907.8293 | 127469059.1934 | 42328520.9871
25 | 16.345 | 18713 | 18712.9985 | 1 | 75386490.5439 | 158.4577 | -158742353.7467 | 309515334.8346 | 39338063.082 | 127469059.1934 | 42323295.9821
26 | 0.41737 | 23020 | 7855 | 0.34123 | 1908293.2573 | 1.1467 | 1865403.3787 | 1951183.1358 | 1865592.0943 | 1951676.897 | 1865725.7626 | 1951506.891
27 | 0.42142 | 24709 | 8497 | 0.34388 | 1926942.7066 | 1.1026 | 1885299.658 | 1968585.7552 | 1885475.3985 | 1969046.8022 | 1885600.7618 | 1968888.0934
28 | 18.3559 | 16844 | 16843.9998 | 1 | 84663471.5317 | 406.4714 | -589824418.5498 | 759151361.6131 | 38856136.3088 | 127469059.1934 | 41862625.2905
29 | 18.7763 | 16167 | 16166.9999 | 1 | 86602685.5786 | 500.4745 | -762893453.9354 | 936098825.0926 | 38667085.4136 | 127469059.1934 | 41674902.3894
30 | 3.4546 | 20085 | 19450.3166 | 0.9684 | 15919805.9388 | 1.1319 | 15566619.173 | 16272992.7046 | 15566795.3686 | 16280435.3481 | 15571361.0652 | 16277853.4854
31 | 3.5734 | 20342 | 19771.1731 | 0.97194 | 16467651.8188 | 1.156 | 16094553.965 | 16840749.6727 | 16094730.1381 | 16849005.7154 | 16099804.3754 | 16846139.6935
32 | 5.8473 | 19173 | 19117.6329 | 0.99711 | 26957803.8505 | 2.2965 | 25744410.5223 | 28171197.1786 | 25747659.003 | 28258307.0185 | 25795652.5397 | 28227118.9821
33 | 6.32 | 19937 | 19901.1147 | 0.9982 | 29138637.8288 | 2.6405 | 27630621.7119 | 30646653.9457 | 27637255.1427 | 30783245.1257 | 27708783.7781 | 30733792.3033
34 | 9.983 | 23344 | 23342.922 | 0.99995 | 46036823.4693 | 9.6511 | 37328554.7742 | 54745092.1644 | 38349497.8133 | 61944151.8715 | 39392616.6022 | 58564241.6576
35 | 9.974 | 24012 | 24010.8811 | 0.99995 | 45995424.6506 | 9.4819 | 37447581.4656 | 54543267.8356 | 38423356.3992 | 61411921.1779 | 39445592.7237 | 58200715.5898
36 | 6.9533 | 17073 | 17056.6866 | 0.99904 | 32060047.6275 | 3.5609 | 29822479.9785 | 34297615.2764 | 29845761.4442 | 34611244.1307 | 29989986.5027 | 34494450.0296
37 | 6.8398 | 17136 | 17117.6587 | 0.99893 | 31536491.7089 | 3.4139 | 29426370.1187 | 33646613.299 | 29445781.0313 | 33923445.4763 | 29576027.1132 | 33820864.9269
38 | 2.0099 | 24776 | 21455.8667 | 0.86599 | 9254889.7128 | 0.80504 | 9108862.0075 | 9400917.418 | 9109019.6984 | 9402463.7135 | 9109866.0541 | 9401931.0458
39 | 2.2728 | 21719 | 19481.5269 | 0.89698 | 10468036.5282 | 0.88238 | 10286999.1474 | 10649073.9089 | 10287174.6753 | 10651309.9004 | 10288446.5125 | 10650538.7463
40 | 11.4706 | 24918 | 24917.74 | 0.99999 | 52899703.0004 | 17.1044 | 35165637.1039 | 70633768.8969 | 40047203.1536 | 113933028.0457 | 41937393.0919 | 91598484.1818
41 | 11.6721 | 24142 | 24141.7941 | 0.99999 | 53829178.8825 | 18.887 | 33902753.7013 | 73755604.0637 | 40019396.5587 | 127469059.1934 | 42037759.2922 | 101477579.1334
42 | 3.9126 | 24784 | 24288.5909 | 0.98001 | 18032526.6914 | 1.1378 | 17630375.7905 | 18434677.5922 | 17630518.2449 | 18444131.3559 | 17636371.2708 | 18440846.3945
43 | 4.225 | 22893 | 22558.1978 | 0.98538 | 19474063.1962 | 1.2852 | 18983536.8804 | 19964589.5121 | 18983738.6557 | 19978573.8313 | 18992333.6834 | 19973699.0911
44 | 8.4397 | 23404 | 23398.9419 | 0.99978 | 38917203.8293 | 5.2702 | 34897306.2455 | 42937101.413 | 35029507.2134 | 44068081.3808 | 35405799.7192 | 43616187.8596
45 | 9.3405 | 26483 | 26480.675 | 0.99991 | 43073081.4473 | 7.0237 | 37143518.9259 | 49002643.9688 | 37528039.6664 | 51793407.3032 | 38184048.3673 | 50593566.1919
46 | 7.7491 | 23871 | 23860.7081 | 0.99957 | 35731268.7854 | 4.0236 | 32913436.4203 | 38549101.1505 | 32960272.3534 | 39064066.4493 | 33173607.5106 | 38867877.7501
47 | 7.7964 | 16792 | 16785.0946 | 0.99959 | 35949407.3878 | 4.8824 | 32509304.6068 | 39389510.1687 | 32593871.2147 | 40187731.658 | 32889051.2256 | 39876036.9488
48 | 0.72913 | 24276 | 12567 | 0.51767 | 3346508.698 | 0.91661 | 3286388.1502 | 3406629.2458 | 3286565.0744 | 3407200.4691 | 3286761.2748 | 3407003.9586
49 | 0.74396 | 31450 | 16504 | 0.52477 | 3414899.5034 | 0.80048 | 3361322.6279 | 3468476.3789 | 3361458.8895 | 3468921.497 | 3361613.6136 | 3468768.4589
50 | 1.0918 | 22859 | 15187 | 0.66438 | 5019446.4578 | 0.85527 | 4935305.263 | 5103587.6526 | 4935490.0124 | 5104373.6432 | 4935818.0914 | 5104103.2184
51 | 1.159 | 22096 | 15162 | 0.68619 | 5329421.1168 | 0.86111 | 5239474.4476 | 5419367.786 | 5239664.9261 | 5420223.431 | 5240032.4991 | 5419929.0032
52 | 0.38103 | 20070 | 6359 | 0.31684 | 1740622.045 | 1.2741 | 1697156.8383 | 1784087.2518 | 1697373.6587 | 1784641.6741 | 1697518.5839 | 1784450.6704
53 | 0.36036 | 21347 | 6459 | 0.30257 | 1645256.6589 | 1.2641 | 1604495.021 | 1686018.2967 | 1604698.9399 | 1686533.1754 | 1604831.0075 | 1686355.8068
54 | 0.30048 | 21735 | 5641 | 0.25954 | 1369022.4825 | 1.3532 | 1332712.853 | 1405332.1121 | 1332913.4598 | 1405820.6203 | 1333031.5154 | 1405652.2733
55 | 0.3094 | 28879 | 7685 | 0.26611 | 1410169.1777 | 1.1592 | 1378130.1015 | 1442208.2538 | 1378280.8906 | 1442577.5105 | 1378371.2482 | 1442450.3762
56 | 10.0901 | 26687 | 26685.8928 | 0.99996 | 46531027.3983 | 9.422 | 37938262.7219 | 55123792.0746 | 38926552.4225 | 62083891.8004 | 39954666.3084 | 58826049.894
57 | 9.986 | 31241 | 31239.5617 | 0.99995 | 46050992.1745 | 8.3529 | 38511814.986 | 53590169.3629 | 39226720.6922 | 58604690.5145 | 40113076.9444 | 56327636.689
58 | 3.6447 | 25211 | 24552.2522 | 0.97387 | 16796752.1766 | 1.056 | 16449098.6614 | 17144405.6918 | 16449228.0225 | 17151536.3991 | 16453643.6984 | 17149063.4425
59 | 3.8179 | 28192 | 27572.521 | 0.97803 | 17595859.1914 | 1.0417 | 17236591.9871 | 17955126.3957 | 17236703.3617 | 17962683.3984 | 17241403.3705 | 17960061.6164
60 | 1.3125 | 30914 | 22593.7576 | 0.73086 | 6037823.4851 | 0.7161 | 5953080.4076 | 6122566.5626 | 5953214.6173 | 6123255.0466 | 5953529.7485 | 6123018.2926
61 | 1.2657 | 31997 | 22972.7534 | 0.71797 | 5821966.2805 | 0.70678 | 5741316.72 | 5902615.8411 | 5741446.765 | 5903256.9422 | 5741735.2682 | 5903036.5078
62 | 0.49113 | 30768 | 11940 | 0.38807 | 2248552.2264 | 0.93144 | 2207502.901 | 2289601.5517 | 2207643.5802 | 2289987.6992 | 2207755.3124 | 2289854.8774
63 | 0.48994 | 29375 | 11378 | 0.38734 | 2243057.0549 | 0.95414 | 2201110.1852 | 2285003.9247 | 2201257.5703 | 2285408.1258 | 2201374.3597 | 2285269.0787
64 | 0.16466 | 19332 | 2935 | 0.15182 | 742477.9532 | 1.8906 | 714964.8833 | 769991.023 | 715191.3625 | 770501.186 | 715296.7653 | 770324.96
65 | 0.16166 | 33952 | 5068 | 0.14927 | 728622.4604 | 1.4393 | 708067.6941 | 749177.2267 | 708196.4406 | 749466.705 | 708256.4995 | 749366.9204
66 | 0.060328 | 27176 | 1591 | 0.058544 | 261152.2265 | 2.6722 | 247474.7744 | 274829.6787 | 247636.2823 | 275174.0385 | 247697.9749 | 275054.7205
67 | 0.061404 | 25337 | 1509 | 0.059557 | 266118.0684 | 2.7407 | 251823.2443 | 280412.8926 | 251996.5034 | 280782.5535 | 252062.7509 | 280654.4285
68 | 0.022457 | 28955 | 643 | 0.022207 | 86444.4112 | 4.7264 | 78436.5965 | 94452.226 | 78588.6184 | 94771.0966 | 78641.5833 | 94659.8722
69 | 0.02105 | 29381 | 612 | 0.02083 | 79951.854 | 4.9097 | 72258.1473 | 87645.5607 | 72407.9909 | 87959.7099 | 72459.985 | 87850.0826
70 | 0.004046 | 27738 | 112 | 0.0040378 | 1508.8344 | 116.8903 | -1947.9161 | 4965.5848 | 4.6132e-06 | 5302.6886 | 0 | 5182.1941
71 | 0.0039417 | 31775 | 125 | 0.0039339 | 1027.7272 | 158.2541 | -2159.9984 | 4215.4528 | 4.6132e-06 | 4509.1303 | 0 | 4404.3858

## 11_R3_Analysis
R3 linear approximation per training well — gamma-based M_r3 estimate
Conc_pgmL | P_specific | M_r3 | CV_r3_pct | R3_valid
0 | 1e-12 | 4.6133e-06 | 31365961396.482 | TRUE
0 | 1e-12 | 4.6133e-06 | 33250686184.8842 | TRUE
0.0328 | 0.021546 | 99395.4178 | 4.1379 | TRUE
0.0328 | 0.022817 | 105259.6232 | 4.0578 | TRUE
0.0984 | 0.059911 | 276386.3989 | 2.471 | TRUE
0.0984 | 0.062915 | 290242.1986 | 2.1704 | TRUE
0.296 | 0.1622 | 748290.2713 | 1.5101 | TRUE
0.296 | 0.17575 | 810778.5398 | 1.1582 | TRUE
0.888 | 0.4166 | 1921879.2019 | 0.66748 | FALSE
0.888 | 0.43947 | 2027409.5717 | 0.64931 | FALSE
2.664 | 0.80609 | 3718716.8973 | 0.27876 | FALSE
2.664 | 0.82484 | 3805218.3214 | 0.26601 | FALSE
8 | 0.98841 | 4559775.3818 | 0.071563 | FALSE
8 | 0.99236 | 4578025.0413 | 0.058122 | FALSE
24 | 1 | 4613247.4488 | 0.00078117 | FALSE
24 | 1 | 4613235.9063 | 0.0012141 | FALSE

## 12_LoB_LoD_LoQ
Per-well LoB / LoD / LoQ (R2 and R3 formulations)
Well | Conc (pg/mL) | n | R2_LoB | R2_LoD | R2_LoQ_low | R2_LoQ_high | R3_LoB | R3_LoD | R3_LoQ_low
1 | 0 | 29568 | 2701.263 | 5853.4942 | 10378.9913 | 55589663.3136 | 2693.8478 | 5387.6955 | 10352.9886
2 | 0 | 33529 | 2561.2543 | 5349.1449 | 9609.0185 | 56284188.9165 | 2529.7274 | 5059.4548 | 9586.861
3 | 0.0328 | 31110 | 2608.4558 | 5668.5467 | 10059.4403 | 55870713.9551 | 2626.2376 | 5252.4752 | 10035.0727
4 | 0.0328 | 30257 | 2707.4827 | 5821.095 | 10232.7796 | 55717046.5986 | 2662.9995 | 5325.9991 | 10207.5321
5 | 0.0984 | 27297 | 2829.3038 | 5948.2941 | 10905.139 | 55147274.2254 | 2803.6679 | 5607.3357 | 10876.322
6 | 0.0984 | 33491 | 2583.6743 | 5522.5985 | 9615.6655 | 56277931.7925 | 2531.1621 | 5062.3243 | 9593.4763
7 | 0.296 | 23171 | 3196.4364 | 6836.4354 | 12082.9228 | 54237862.871 | 3043.0678 | 6086.1357 | 12047.2513
8 | 0.296 | 35706 | 2524.2967 | 5359.6072 | 9248.349 | 56631089.9316 | 2451.3957 | 4902.7914 | 9227.8818
9 | 0.888 | 31713 | 2670.1198 | 5702.9732 | 9941.7076 | 55976773.9909 | 2601.1498 | 5202.2996 | 9917.9284
10 | 0.888 | 30509 | 2694.92 | 5793.6882 | 10180.7117 | 55762899.3113 | 2651.9787 | 5303.9575 | 10155.7302
11 | 2.664 | 31098 | 2616.0989 | 5677.3732 | 10061.8226 | 55868582.0991 | 2626.7443 | 5253.4885 | 10037.4431
12 | 2.664 | 30145 | 2781.4435 | 5906.649 | 10256.1589 | 55696542.4785 | 2667.942 | 5335.8839 | 10230.7913
13 | 8 | 22988 | 3157.2468 | 6591.7697 | 12143.5269 | 54193782.3145 | 3055.1563 | 6110.3125 | 12107.4819
14 | 8 | 22866 | 3063.2256 | 6718.7535 | 12184.395 | 54164195.7701 | 3063.2957 | 6126.5914 | 12148.0969
15 | 24 | 28320 | 2760.2232 | 5991.4024 | 10659.2029 | 55351027.9059 | 2752.5638 | 5505.1277 | 10631.7203
16 | 24 | 28760 | 2777.0613 | 5982.5811 | 10558.041 | 55436365.1483 | 2731.4269 | 5462.8538 | 10531.0977

## 13_LOO_PerFold
LOO per-fold logL: positive Diff = 4PL better, negative = TCS better
Fold | 4PL_logL | TCS_logL | Diff(4PL-TCS)
1 | -588.6824 | -593.2918 | 4.6093
2 | -820.5875 | -817.1996 | -3.3879
3 | -3658.4051 | -3658.2192 | -0.1859
4 | -3700.9755 | -3700.2936 | -0.6819
5 | -6448.4359 | -6448.2228 | -0.2131
6 | -8180.796 | -8181.5184 | 0.7224
7 | -10400.9086 | -10394.1582 | -6.7504
8 | -16769.0267 | -16772.4434 | 3.4168
9 | -21588.0014 | -21571.4568 | -16.5446
10 | -20942.9045 | -20950.9438 | 8.0393
11 | -15269.7034 | -15269.4476 | -0.2558
12 | -13996.5545 | -13982.7858 | -13.7687
13 | -1470.6165 | -1488.1464 | 17.5299
14 | -1029.0191 | -1023.9963 | -5.0228
15 | -0.7753 | -0.74 | -0.0354
16 | -1.636 | -1.8316 | 0.1956

## 14_Reference_CV
Manufacturer concentration -> TCS predicted CV (typical n=mean(n_obs)=29408)
Conc (pg/mL) | M (molecules) | P_pos | CV (%) | Note
0 | 0 | 0.003712 |  | Blank (background only)
0.0055 | 15923.5577 | 0.0071449 | 14.3849 | Manufacturer LOD
0.01 | 28951.9231 | 0.0099449 | 9.3472 | Analytical LLOQ
0.04 | 115807.6923 | 0.028411 | 3.9871 | Functional LLOQ
0.607 | 1757381.7308 | 0.31932 | 1.0524 | Plasma lower
1.79 | 5182394.2308 | 0.67603 | 0.75264 | Plasma mean
4.64 | 13433692.3077 | 0.94583 | 0.83992 | Plasma upper
24 | 69484615.3846 | 1 | 72.6055 | Highest calibrator

## 15_TestSet_BackCalc
Test set 68 samples — back-calculated concentration (4PL vs TCS) — KEY for clinical validation
Test sample # | AEB (raw) | n (beads) | k (positive) | P_obs | Conc_4PL (pg/mL) | Conc_TCS (pg/mL) | Diff_4PL-TCS | ¦log10 ratio¦
1 | 0.2366 | 29071 | 6125 | 0.21069 | 0.36384 | 0.37107 | -0.0072331 | 0.0085491
2 | 0.27244 | 33328 | 7948 | 0.23848 | 0.41988 | 0.42818 | -0.0083003 | 0.0085015
3 | 2.6172 | 28472 | 26393.3188 | 0.92699 | 4.2521 | 4.1644 | 0.087704 | 0.0090516
4 | 2.5644 | 28456 | 26265.9644 | 0.92304 | 4.1619 | 4.0803 | 0.081602 | 0.0085998
5 | 0.058197 | 20960 | 1185 | 0.056536 | 0.085626 | 0.086807 | -0.0011809 | 0.0059487
6 | 0.058413 | 24727 | 1403 | 0.05674 | 0.085962 | 0.08715 | -0.0011887 | 0.0059642
7 | 0.35136 | 25139 | 7448 | 0.29627 | 0.54351 | 0.55395 | -0.010433 | 0.0082573
8 | 0.36993 | 27922 | 8634 | 0.30922 | 0.57264 | 0.58353 | -0.010889 | 0.0081808
9 | 0.10277 | 22721 | 2219 | 0.097663 | 0.15505 | 0.15783 | -0.0027777 | 0.0077116
10 | 0.10608 | 22127 | 2227 | 0.10065 | 0.16021 | 0.1631 | -0.0028952 | 0.0077783
11 | 0.12617 | 24019 | 2847 | 0.11853 | 0.19151 | 0.19511 | -0.0036017 | 0.008092
12 | 0.12558 | 20192 | 2383 | 0.11802 | 0.1906 | 0.19418 | -0.0035814 | 0.0080847
13 | 0.33779 | 22124 | 6342 | 0.28666 | 0.52223 | 0.53232 | -0.010088 | 0.0083094
14 | 0.34098 | 21902 | 6328 | 0.28892 | 0.52722 | 0.53739 | -0.01017 | 0.0082975
15 | 0.093416 | 25262 | 2253 | 0.089185 | 0.14048 | 0.14292 | -0.0024448 | 0.007493
16 | 0.09299 | 27985 | 2485 | 0.088798 | 0.13982 | 0.14225 | -0.0024296 | 0.0074818
17 | 2.396 | 26420 | 24013.6774 | 0.90892 | 3.8753 | 3.8119 | 0.063406 | 0.0071644
18 | 2.2728 | 23602 | 21170.4048 | 0.89698 | 3.6669 | 3.6156 | 0.051322 | 0.0061214
19 | 0.33734 | 27587 | 7899 | 0.28633 | 0.52151 | 0.53159 | -0.010076 | 0.0083111
20 | 0.34979 | 26023 | 7681 | 0.29516 | 0.54104 | 0.55143 | -0.010393 | 0.0082636
21
22
23
24
25 | 16.7462 | 18699 | 18698.999 | 1 | 38.3747
26 | 16.345 | 18713 | 18712.9985 | 1 | 37.0203
27 | 0.41737 | 23020 | 7855 | 0.34123 | 0.64715 | 0.65912 | -0.011974 | 0.007962
28 | 0.42142 | 24709 | 8497 | 0.34388 | 0.65351 | 0.66557 | -0.012061 | 0.007942
29 | 18.3559 | 16844 | 16843.9998 | 1 | 44.1458
30 | 18.7763 | 16167 | 16166.9999 | 1 | 45.7484
31 | 3.4546 | 20085 | 19450.3166 | 0.9684 | 5.7097 | 5.4987 | 0.21096 | 0.016351
32 | 3.5734 | 20342 | 19771.1731 | 0.97194 | 5.9205 | 5.6879 | 0.23256 | 0.017403
33 | 5.8473 | 19173 | 19117.6329 | 0.99711 | 10.1695 | 9.3112 | 0.8583 | 0.038294
34 | 6.32 | 19937 | 19901.1147 | 0.9982 | 11.1069 | 10.0645 | 1.0424 | 0.042802
35 | 9.983 | 23344 | 23342.922 | 0.99995 | 19.1023 | 15.9011 | 3.2012 | 0.079658
36 | 9.974 | 24012 | 24010.8811 | 0.99995 | 19.081 | 15.8868 | 3.1942 | 0.079564
37 | 6.9533 | 17073 | 17056.6866 | 0.99904 | 12.3941 | 11.0735 | 1.3206 | 0.048929
38 | 6.8398 | 17136 | 17117.6587 | 0.99893 | 12.1607 | 10.8927 | 1.268 | 0.047824
39 | 2.0099 | 24776 | 21455.8667 | 0.86599 | 3.2256 | 3.1966 | 0.028972 | 0.0039185
40 | 2.2728 | 21719 | 19481.5269 | 0.89698 | 3.667 | 3.6157 | 0.051327 | 0.0061218
41 | 11.4706 | 24918 | 24917.74 | 0.99999 | 22.7733 | 18.2716 | 4.5018 | 0.095651
42 | 11.6721 | 24142 | 24141.7941 | 0.99999 | 23.2919 | 18.5926 | 4.6993 | 0.097865
43 | 3.9126 | 24784 | 24288.5909 | 0.98001 | 6.5285 | 6.2284 | 0.30004 | 0.020433
44 | 4.225 | 22893 | 22558.1978 | 0.98538 | 7.0963 | 6.7263 | 0.36993 | 0.023251
45 | 8.4397 | 23404 | 23398.9419 | 0.99978 | 15.5658 | 13.442 | 2.1238 | 0.063707
46 | 9.3405 | 26483 | 26480.675 | 0.99991 | 17.5983 | 14.8775 | 2.7208 | 0.072941
47 | 7.7491 | 23871 | 23860.7081 | 0.99957 | 14.0651 | 12.3416 | 1.7235 | 0.056771
48 | 7.7964 | 16792 | 16785.0946 | 0.99959 | 14.1663 | 12.4169 | 1.7494 | 0.057242
49 | 0.72913 | 24276 | 12567 | 0.51767 | 1.1398 | 1.1559 | -0.016039 | 0.0060683
50 | 0.74396 | 31450 | 16504 | 0.52477 | 1.1634 | 1.1795 | -0.016095 | 0.0059668
51 | 1.0918 | 22859 | 15187 | 0.66438 | 1.7201 | 1.7337 | -0.013666 | 0.0034369
52 | 1.159 | 22096 | 15162 | 0.68619 | 1.8284 | 1.8408 | -0.012354 | 0.0029244
53 | 0.38103 | 20070 | 6359 | 0.31684 | 0.59006 | 0.60121 | -0.011153 | 0.0081324
54 | 0.36036 | 21347 | 6459 | 0.30257 | 0.55762 | 0.56827 | -0.010656 | 0.008221
55 | 0.30048 | 21735 | 5641 | 0.25954 | 0.46377 | 0.47286 | -0.009093 | 0.0084327
56 | 0.3094 | 28879 | 7685 | 0.26611 | 0.47774 | 0.48707 | -0.0093371 | 0.0084062
57 | 10.0901 | 26687 | 26685.8928 | 0.99996 | 19.3577 | 16.0718 | 3.2859 | 0.080789
58 | 9.986 | 31241 | 31239.5617 | 0.99995 | 19.1096 | 15.906 | 3.2036 | 0.07969
59 | 3.6447 | 25211 | 24552.2522 | 0.97387 | 6.0476 | 5.8016 | 0.24603 | 0.018038
60 | 3.8179 | 28192 | 27572.521 | 0.97803 | 6.358 | 6.0776 | 0.28034 | 0.019584
61 | 1.3125 | 30914 | 22593.7576 | 0.73086 | 2.0772 | 2.0855 | -0.0083058 | 0.0017331
62 | 1.2657 | 31997 | 22972.7534 | 0.71797 | 2.0012 | 2.0109 | -0.0096947 | 0.0020988
63 | 0.49113 | 30768 | 11940 | 0.38807 | 0.76323 | 0.77665 | -0.013423 | 0.0075714
64 | 0.48994 | 29375 | 11378 | 0.38734 | 0.76135 | 0.77475 | -0.013402 | 0.0075781
65 | 0.16466 | 19332 | 2935 | 0.15182 | 0.25153 | 0.25645 | -0.004921 | 0.0084146
66 | 0.16166 | 33952 | 5068 | 0.14927 | 0.24685 | 0.25167 | -0.00482 | 0.0083984
67 | 0.060328 | 27176 | 1591 | 0.058544 | 0.088945 | 0.090202 | -0.0012575 | 0.0060969
68 | 0.061404 | 25337 | 1509 | 0.059557 | 0.090621 | 0.091917 | -0.0012961 | 0.0061677
69 | 0.022457 | 28955 | 643 | 0.022207 | 0.029939 | 0.029858 | 8.0797e-05 | 0.0011736
70 | 0.02105 | 29381 | 612 | 0.02083 | 0.027744 | 0.027615 | 0.00012847 | 0.0020156
71 | 0.004046 | 27738 | 112 | 0.0040378 | 0.0011749 | 0.00052115 | 0.00065378 | 0.35305
72 | 0.0039417 | 31775 | 125 | 0.0039339 | 0.0010111 | 0.00035498 | 0.00065608 | 0.45457
Summary | 4PL | TCS
Min (pg/mL) | 0.0010111 | 0.00035498
Median (pg/mL) | 1.4417 | 0.96627
Max (pg/mL) | 45.7484 | 18.5926
# NaN (above ULOQ) | 4 | 8
# =0 (below LLOQ) | 0 | 0

## 16_Train_RealConc
Training-set real concentrations in pg/mL (TCS scale and 4PL scale)
Concentration (input, pg/mL) | conc_train_real (TCS scale, pg/mL) | conc_train_4pl_real (4PL scale, pg/mL) | Ratio (4PL / TCS)
0 | 0 | 0
0 | 0 | 0
0.0328 | 0.0328 | 0.0328 | 1
0.0328 | 0.0328 | 0.0328 | 1
0.0984 | 0.0984 | 0.0984 | 1
0.0984 | 0.0984 | 0.0984 | 1
0.296 | 0.296 | 0.296 | 1
0.296 | 0.296 | 0.296 | 1
0.888 | 0.888 | 0.888 | 1
0.888 | 0.888 | 0.888 | 1
2.664 | 2.664 | 2.664 | 1
2.664 | 2.664 | 2.664 | 1
8 | 8 | 8 | 1
8 | 8 | 8 | 1
24 | 24 | 24 | 1
24 | 24 | 24 | 1

## 17_LoB_LoD_LoQ_Dist
Per-well LoB / LoD / LoQ in pg/mL (converted from molecules, used for plotting distributions)
Well | R2_LoB (pg/mL) | R2_LoD (pg/mL) | R2_LoQ_low (pg/mL) | R2_LoQ_high (pg/mL) | R3_LoB (pg/mL) | R3_LoD (pg/mL) | R3_LoQ_low (pg/mL)
1 | 0.00093302 | 0.0020218 | 0.0035849 | 19.2007 | 0.00093046 | 0.0018609 | 0.0035759
2 | 0.00088466 | 0.0018476 | 0.003319 | 19.4406 | 0.00087377 | 0.0017475 | 0.0033113
3 | 0.00090096 | 0.0019579 | 0.0034745 | 19.2978 | 0.0009071 | 0.0018142 | 0.0034661
4 | 0.00093517 | 0.0020106 | 0.0035344 | 19.2447 | 0.0009198 | 0.0018396 | 0.0035257
5 | 0.00097724 | 0.0020545 | 0.0037666 | 19.0479 | 0.00096839 | 0.0019368 | 0.0037567
6 | 0.0008924 | 0.0019075 | 0.0033213 | 19.4384 | 0.00087426 | 0.0017485 | 0.0033136
7 | 0.001104 | 0.0023613 | 0.0041734 | 18.7338 | 0.0010511 | 0.0021022 | 0.0041611
8 | 0.00087189 | 0.0018512 | 0.0031944 | 19.5604 | 0.00084671 | 0.0016934 | 0.0031873
9 | 0.00092226 | 0.0019698 | 0.0034339 | 19.3344 | 0.00089844 | 0.0017969 | 0.0034257
10 | 0.00093083 | 0.0020011 | 0.0035164 | 19.2605 | 0.00091599 | 0.001832 | 0.0035078
11 | 0.0009036 | 0.001961 | 0.0034754 | 19.297 | 0.00090728 | 0.0018146 | 0.0034669
12 | 0.00096071 | 0.0020402 | 0.0035425 | 19.2376 | 0.00092151 | 0.001843 | 0.0035337
13 | 0.0010905 | 0.0022768 | 0.0041944 | 18.7185 | 0.0010553 | 0.0021105 | 0.0041819
14 | 0.001058 | 0.0023207 | 0.0042085 | 18.7083 | 0.0010581 | 0.0021161 | 0.004196
15 | 0.00095338 | 0.0020694 | 0.0036817 | 19.1183 | 0.00095074 | 0.0019015 | 0.0036722
16 | 0.0009592 | 0.0020664 | 0.0036467 | 19.1477 | 0.00094344 | 0.0018869 | 0.0036374
Median (pg/mL) | 0.00093409 | 0.0020162 | 0.0035384 | 19.2411 | 0.00092065 | 0.0018413 | 0.0035297
Min (pg/mL) | 0.00087189 | 0.0018476 | 0.0031944 | 18.7083 | 0.00084671 | 0.0016934 | 0.0031873
Max (pg/mL) | 0.001104 | 0.0023613 | 0.0042085 | 19.5604 | 0.0010581 | 0.0021161 | 0.004196