## README
SI Table 6: droplet digital immunoassay workbook
Raw droplet-bead counts, TCS/R2 fit results and counting-statistics
performance bounds of the home-built platform (main-text Fig. 3c-e).
Sheets:
Standards_raw - positive/negative bead counts per standard
Fit_parameters - R2 binomial-likelihood fit (lambda_max, C50, b),
kappa, slope, R2, Fisher covariance, bootstrap CI
LoB_LoD_LoQ - point estimates and parametric-bootstrap 95% CI
Clinical_12 - twelve clinical plasma samples: raw counts,
droplet-platform back-calculation, Roche Cobas 411
and Simoa HD-X reference values
Bootstrap_samples - raw bootstrap draws for LoB / LoD / LoQ (fg/mL)
Model: lambda_tot(C) = b + lambda_max*C/(C50 + C); positives binomial
with p = 1 - exp(-lambda_tot); fitted jointly in log space (SI S2f).
kappa = V_d / (slope in molecule units) - 1, V_d = 160 uL / 2e6 beads.

## Standards_raw
nominal_IL6_fg_per_mL | positive_beads | negative_beads | total_beads | lambda_observed | lambda_CI95_lower | lambda_CI95_upper
1000000 | 26675 | 66273 | 92948 | 0.33826 | 0.33419 | 0.34236
100000 | 5655 | 90581 | 96236 | 0.060559 | 0.05899 | 0.062159
10000 | 776 | 114680 | 115456 | 0.0067439 | 0.0062776 | 0.0072355
1000 | 125 | 139732 | 139857 | 0.00089417 | 0.0007443 | 0.0010654
100 | 56 | 193471 | 193527 | 0.00028941 | 0.00021861 | 0.00037582
10 | 35 | 133020 | 133055 | 0.00026308 | 0.00018325 | 0.00036589
1 | 18 | 161295 | 161313 | 0.00011159 | 6.6135e-05 | 0.00017636
0 | 21 | 154409 | 154430 | 0.00013599 | 8.4182e-05 | 0.00020788

## Fit_parameters
quantity | estimate | CI95_lower | CI95_upper | note
lambda_max | 0.68975 | 0.66206 | 0.71814 | -
C50_fg_per_mL | 1040331.3976 | 973084.9468 | 1111481.0557 | -
background_b | 0.0001792 | 0.00014863 | 0.00021605 | measured blank lambda = 1.360e-04
linear_slope_per_fg_mL | 6.6301e-07 | 6.4368e-07 | 6.8271e-07 | -
kappa | 3475.7802 | 3375.4444 | 3580.1893 | platform binding-regime parameter
R_squared_fit | 1 |  |  | all seven standards + blank
R_squared_vs_Cobas | 0.96973 |  |  | n = 12
R_squared_vs_Simoa | 0.93885 |  |  | n = 10 in range
degeneracy_major_slope | 0.60216 |  |  | d(ln lambda_max)/d(ln C50) along the 95% ellipse major axis
degeneracy_axis_ratio | 8.2743 |  |  | major/minor axis ratio of the log-space Fisher 95% ellipse
bootstrap_corr_ln_lmax_ln_C50 | 0.96332 |  |  | Pearson correlation of the log-space bootstrap samples
log-space Fisher covariance (R2 likelihood)
| ln_lambda_max | ln_C50 | ln_b
ln_lambda_max | 0.00042663 | 0.00067111 | 0.00012668
ln_C50 | 0.00067111 | 0.001137 | 0.00025234
ln_b | 0.00012668 | 0.00025234 | 0.0091048

## LoB_LoD_LoQ
bound | point_estimate_fg_per_mL | bootstrap_median | bootstrap_CI95_lower | bootstrap_CI95_upper
LoB | 73.6302 | 73.3615 | 56.698 | 88.8847
LoD | 177.7591 | 177.1103 | 136.881 | 214.5866
LoQ | 491.5822 | 491.1603 | 459.3461 | 524.8515

## Clinical_12
total_beads | positive_beads | lambda_observed | lambda_CI95_lower | lambda_CI95_upper | droplet_platform_pg_per_mL | Roche_Cobas411_pg_per_mL | Simoa_HDX_pg_per_mL
130564 | 44046 | 0.41151 | 0.40765 | 0.4154 | 1536.997 | 1865
136445 | 32237 | 0.26953 | 0.26659 | 0.2725 | 666.5605 | 1037
115343 | 13018 | 0.11976 | 0.11771 | 0.12183 | 218.1807 | 558.8 | 161.13
165529 | 880 | 0.0053305 | 0.004984 | 0.0056946 | 7.828 | 9.93 | 3.27
118828 | 17596 | 0.16026 | 0.1579 | 0.16265 | 314.4252 | 532.8 | 184.92
172737 | 2577 | 0.015031 | 0.014456 | 0.015623 | 22.8937 | 66.82 | 29.43
142637 | 3898 | 0.027708 | 0.026845 | 0.028592 | 43.248 | 123.6 | 52.67
197965 | 10288 | 0.053368 | 0.052341 | 0.054409 | 86.9267 | 264.7 | 90.35
170980 | 5435 | 0.032304 | 0.03145 | 0.033174 | 50.8194 | 150.9 | 60.28
195295 | 1869 | 0.0096162 | 0.0091851 | 0.010062 | 14.4311 | 45.37 | 17.56
203242 | 13412 | 0.068268 | 0.067118 | 0.069434 | 113.9462 | 327.6 | 106.65
171355 | 2880 | 0.01695 | 0.016337 | 0.017581 | 25.9256 | 95.15 | 34.34

## Bootstrap_samples
lambda_max | C50_fg_per_mL | LoB_fg_per_mL | LoD_fg_per_mL | LoQ_fg_per_mL
0.71687 | 1105143.6241 | 80.4543 | 194.2338 | 512.3124
0.7087 | 1108863.7815 | 74.5406 | 179.957 | 506.5101
0.68643 | 1033241.6076 | 81.7638 | 197.3954 | 460.6244
0.68483 | 1033771.8219 | 77.1213 | 186.1873 | 501.2932
0.68053 | 1002820.4937 | 60.7964 | 146.7755 | 475.3175
0.66829 | 998527.3529 | 87.1826 | 210.4775 | 475.769
0.69101 | 1042352.3253 | 78.7238 | 190.056 | 500.583
0.7382 | 1151243.4993 | 83.0679 | 200.5436 | 493.2542
0.67921 | 1025344.3843 | 64.325 | 155.2943 | 483.8545
0.67985 | 1014953.4133 | 71.1231 | 171.7064 | 509.3505
0.66066 | 977488.2869 | 68.7025 | 165.8625 | 507.5535
0.68913 | 1059847.0869 | 86.6952 | 209.3007 | 510.4867
0.70946 | 1075517.4311 | 77.4498 | 186.9803 | 499.2081
0.6978 | 1053212.0848 | 80.394 | 194.0882 | 500.2764
0.69261 | 1055237.1487 | 72.5834 | 175.2319 | 473.8696
0.69024 | 1041499.7735 | 68.195 | 164.6372 | 495.7913
0.69543 | 1032416.2838 | 72.4728 | 174.9647 | 478.1981
0.71235 | 1093688.0653 | 61.195 | 147.7377 | 499.1865
0.70228 | 1069638.0254 | 81.1268 | 195.8575 | 485.4011
0.67061 | 992166.1379 | 73.9247 | 178.47 | 468.5312
0.71075 | 1085524.901 | 79.7069 | 192.4295 | 493.0933
0.66212 | 985863.0891 | 69.1377 | 166.9131 | 498.5398
0.68581 | 1015798.9018 | 86.4248 | 208.6479 | 509.0838
0.68441 | 1030619.773 | 83.3564 | 201.2402 | 493.9444
0.69099 | 1037989.1051 | 78.3963 | 189.2655 | 492.7198
0.69664 | 1051448.9418 | 66.2923 | 160.0438 | 475.7432
0.70202 | 1084796.2381 | 73.6168 | 177.7266 | 505.9833
0.67832 | 1009211.2972 | 59.3018 | 143.1672 | 507.1026
0.68376 | 1029063.4316 | 64.1291 | 154.8213 | 467.7636
0.65546 | 963562.5381 | 75.1047 | 181.3188 | 470.1792
0.67808 | 1011625.4485 | 77.8598 | 187.9701 | 491.1742
0.69194 | 1059220.6981 | 89.3212 | 215.6404 | 510.21
0.68299 | 1015694.5504 | 69.0533 | 166.7093 | 489.5984
0.68111 | 1026674.5155 | 69.9925 | 168.9768 | 505.2503
0.72235 | 1113787.128 | 73.4564 | 177.3394 | 465.8377
0.69007 | 1022887.1849 | 65.1051 | 157.1775 | 466.6299
0.71565 | 1103264.5713 | 65.6896 | 158.5888 | 493.9068
0.68142 | 1024622.8107 | 73.4042 | 177.2133 | 457.158
0.66777 | 986613.864 | 64.8938 | 156.6674 | 485.3933
0.69867 | 1050086.0898 | 76.7858 | 185.3774 | 488.9536
0.69314 | 1055702.3396 | 72.5598 | 175.1748 | 497.5257
0.67858 | 1014983.5398 | 79.6703 | 192.3412 | 502.843
0.67862 | 1024560.1342 | 77.1336 | 186.2171 | 468.2104
0.68488 | 1041179.6559 | 70.5911 | 170.422 | 502.8097
0.70018 | 1065876.6129 | 81.0844 | 195.755 | 478.3611
0.69425 | 1059344.3697 | 81.2752 | 196.2156 | 464.9398
0.70498 | 1075590.2847 | 72.685 | 175.4771 | 485.6506
0.71449 | 1109085.9899 | 70.1561 | 169.3718 | 493.2892
0.67006 | 999837.8472 | 76.2332 | 184.0432 | 493.4678
0.67618 | 1014366.7243 | 63.9221 | 154.3217 | 463.3331
0.67835 | 1017666.0788 | 65.8927 | 159.079 | 525.171
0.68214 | 1015216.831 | 52.5811 | 126.9419 | 475.6056
0.69235 | 1047539.2862 | 80.5905 | 194.5627 | 493.3
0.69889 | 1048117.6059 | 76.6186 | 184.9736 | 494.2704
0.70655 | 1085061.1903 | 78.4585 | 189.4155 | 462.2748
0.67718 | 1009141.8417 | 77.7713 | 187.7566 | 489.4671
0.65463 | 945827.8939 | 68.8325 | 166.1762 | 479.3288
0.67998 | 1028124.0122 | 75.5484 | 182.39 | 473.5544
0.67322 | 1005317.2047 | 63.6297 | 153.6156 | 502.5195
0.67481 | 1003688.2944 | 63.3778 | 153.0075 | 493.1032
0.66344 | 986843.2733 | 75.994 | 183.4658 | 495.5747
0.70084 | 1071422.6539 | 74.6307 | 180.1744 | 466.548
0.70731 | 1099989.371 | 77.7059 | 187.5986 | 453.8822
0.68287 | 1035929.7676 | 79.1713 | 191.1364 | 476.5085
0.69651 | 1039431.4755 | 74.5667 | 180.0199 | 468.3687
0.65353 | 952336.9731 | 71.138 | 171.7424 | 476.868
0.70059 | 1078942.3453 | 76.9505 | 185.7749 | 504.974
0.66243 | 971005.2175 | 68.0641 | 164.3213 | 491.8791
0.69047 | 1044417.7135 | 58.0974 | 140.2595 | 518.3665
0.70742 | 1078112.2876 | 72.6043 | 175.2822 | 501.1107
0.66566 | 979501.3557 | 64.6304 | 156.0317 | 488.5757
0.70283 | 1068868.4557 | 72.4517 | 174.9138 | 513.2914
0.68036 | 998624.491 | 79.7293 | 192.4837 | 496.4521
0.67936 | 1037694.5691 | 69.0339 | 166.6625 | 469.0918
0.66041 | 973007.1733 | 58.7246 | 141.7736 | 486.4352
0.66804 | 998466.0471 | 67.5499 | 163.0799 | 517.9831
0.68891 | 1047111.2873 | 68.695 | 165.8444 | 471.2994
0.66864 | 976573.7545 | 74.6184 | 180.1448 | 458.4931
0.68074 | 1021253.5568 | 73.2365 | 176.8086 | 491.5871
0.68827 | 1036765.3601 | 80.235 | 193.7045 | 498.4554
0.68564 | 1033055.018 | 76.9767 | 185.8383 | 465.3467
0.67299 | 1003090.534 | 71.0081 | 171.4286 | 486.1117
0.6946 | 1044424.2077 | 80.0904 | 193.3552 | 486.9489
0.69564 | 1057668.6311 | 66.7804 | 161.2222 | 509.8877
0.69699 | 1067799.4863 | 56.533 | 136.4828 | 487.913
0.68059 | 1028267.1659 | 62.3331 | 150.4854 | 491.0095
0.70109 | 1060763.8108 | 78.9619 | 190.631 | 471.4323
0.6737 | 1001652.3472 | 70.8315 | 171.0023 | 486.3816
0.69062 | 1049233.6729 | 64.7369 | 156.2887 | 480.8691
0.68405 | 1014322.46 | 72.3871 | 174.7579 | 517.5134
0.67179 | 992000.4465 | 62.9206 | 151.9038 | 509.6558
0.67003 | 979319.534 | 74.673 | 180.2765 | 479.7066
0.68584 | 1028818.0982 | 71.4646 | 172.5307 | 495.5167
0.71598 | 1113578.5767 | 72.2203 | 174.3553 | 479.1037
0.718 | 1109077.7319 | 69.8125 | 168.5422 | 510.8534
0.67595 | 1001105.124 | 74.002 | 178.6567 | 510.7634
0.7002 | 1074538.6154 | 71.2583 | 172.0327 | 461.3946
0.69438 | 1048058.4981 | 62.2713 | 150.3362 | 460.4364
0.67536 | 993822.7029 | 62.7032 | 151.3789 | 493.8626
0.69375 | 1058779.5058 | 87.5543 | 211.3749 | 481.8472
0.70209 | 1089060.8086 | 85.865 | 207.2964 | 507.9868
0.69032 | 1050191.3995 | 77.7224 | 187.6386 | 490.0404
0.67809 | 1009874.4277 | 67.3098 | 162.5002 | 460.9135
0.69659 | 1057115.1777 | 88.5486 | 213.7752 | 470.6498
0.68259 | 1013243.9293 | 77.4685 | 187.0256 | 477.2149
0.70251 | 1087235.5568 | 80.7691 | 194.994 | 497.1177
0.69908 | 1067286.6762 | 72.733 | 175.593 | 479.545
0.69302 | 1047025.4179 | 68.2821 | 164.8476 | 493.9615
0.67028 | 991264.2491 | 61.0142 | 147.3014 | 478.0746
0.69551 | 1052196.0958 | 83.7428 | 202.1729 | 519.1851
0.69871 | 1056362.0061 | 72.0262 | 173.8866 | 480.4483
0.69206 | 1062268.5153 | 67.4173 | 162.7597 | 497.4142
0.68083 | 1014876.156 | 69.2165 | 167.1035 | 489.409
0.68055 | 1002644.422 | 73.615 | 177.7223 | 462.7515
0.66727 | 996450.4816 | 65.5895 | 158.3471 | 474.4088
0.6565 | 967350.4945 | 80.0398 | 193.2333 | 466.2758
0.68009 | 1028390.392 | 78.9164 | 190.521 | 516.8787
0.69967 | 1061430.943 | 79.1725 | 191.1393 | 507.6624
0.68681 | 1046080.3941 | 72.5612 | 175.1783 | 495.7079
0.70455 | 1068193.0347 | 75.755 | 182.8889 | 474.0443
0.68445 | 1033902.5398 | 75.4769 | 182.2173 | 509.523
0.69773 | 1066972.5377 | 78.1258 | 188.6124 | 485.6732
0.66763 | 996463.4717 | 61.5782 | 148.663 | 504.3398
0.68508 | 1027815.9332 | 71.4745 | 172.5546 | 497.9075
0.66588 | 981382.2832 | 56.6064 | 136.6598 | 477.8914
0.67137 | 1012788.4739 | 73.6428 | 177.7894 | 492.9037
0.68093 | 1009049.0073 | 68.8093 | 166.1203 | 493.2678
0.70813 | 1083807.2406 | 65.2165 | 157.4465 | 489.0459
0.67868 | 1016908.3268 | 61.8185 | 149.2432 | 506.191
0.65843 | 966937.3336 | 73.3778 | 177.1497 | 473.4092
0.69754 | 1048937.3004 | 66.0485 | 159.4551 | 478.9697
0.67347 | 999825.5692 | 83.6872 | 202.0387 | 484.111
0.71104 | 1094443.8319 | 76.9084 | 185.6733 | 509.0903
0.70501 | 1074639.1556 | 70.7787 | 170.8749 | 468.5138
0.6922 | 1037813.0563 | 74.914 | 180.8585 | 470.9884
0.6987 | 1055608.3479 | 55.7511 | 134.595 | 477.2978
0.70271 | 1063106.3442 | 86.7908 | 209.5316 | 472.6787
0.69906 | 1055046.9453 | 73.6768 | 177.8716 | 505.2928
0.68104 | 1020511.0575 | 79.8144 | 192.6891 | 482.2017
0.66812 | 1003985.8052 | 61.9979 | 149.6763 | 522.1242
0.67911 | 1018021.2177 | 73.1799 | 176.6719 | 476.8426
0.67819 | 1015908.4392 | 73.1263 | 176.5426 | 480.643
0.68913 | 1038502.0112 | 84.9495 | 205.0861 | 486.5393
0.70177 | 1078910.8066 | 76.8189 | 185.4573 | 460.9515
0.68613 | 1045128.086 | 74.3596 | 179.5201 | 504.8639
0.67881 | 1007256.8612 | 67.0642 | 161.9072 | 517.7079
0.67959 | 991382.2161 | 67.7383 | 163.5347 | 516.5819
0.66692 | 971755.2272 | 71.1303 | 171.7238 | 492.0067
0.68168 | 1024223.169 | 71.58 | 172.8095 | 530.0413
0.67513 | 1000943.9053 | 54.7098 | 132.0812 | 499.2828
0.72133 | 1117275.9701 | 82.5018 | 199.1769 | 489.929
0.69861 | 1042784.4988 | 82.6256 | 199.4759 | 490.5248
0.69213 | 1036013.641 | 63.7813 | 153.9816 | 504.7236
0.68142 | 1017255.0355 | 72.8767 | 175.9398 | 486.8583
0.66358 | 973321.6702 | 60.5151 | 146.0964 | 475.9265
0.68435 | 1024847.1401 | 76.5085 | 184.7078 | 495.8188
0.71339 | 1098153.2602 | 69.5716 | 167.9606 | 496.415
0.67764 | 1015798.3517 | 76.5839 | 184.89 | 500.9616
0.69797 | 1046903.4689 | 78.2781 | 188.9801 | 490.6042
0.67964 | 1019072.0017 | 78.2524 | 188.9181 | 482.0251
0.67882 | 1010559.2998 | 61.4197 | 148.2802 | 509.9071
0.71213 | 1102403.0303 | 85.6907 | 206.8756 | 495.8107
0.68598 | 1044201.9497 | 68.7967 | 166.0898 | 507.7268
0.68947 | 1039210.9496 | 57.891 | 139.7613 | 508.5723
0.6856 | 1024337.2119 | 72.9362 | 176.0836 | 496.3509
0.69909 | 1062894.7451 | 70.598 | 170.4386 | 494.6049
0.70338 | 1072435.798 | 81.2124 | 196.0641 | 498.1796
0.66037 | 977127.9721 | 78.8135 | 190.2726 | 467.4793
0.68115 | 1033527.2675 | 70.4561 | 170.0961 | 474.3943
0.6943 | 1076135.5622 | 88.9187 | 214.6687 | 464.3812
0.69056 | 1061762.8642 | 69.4898 | 167.7633 | 467.5774
0.66711 | 968513.374 | 70.8728 | 171.1021 | 519.6253
0.67033 | 989560.0079 | 66.7192 | 161.0743 | 512.4602
0.69624 | 1061756.4096 | 85.9645 | 207.5367 | 516.8681
0.70253 | 1077522.938 | 67.3661 | 162.6362 | 486.4059
0.70251 | 1074238.7604 | 60.9492 | 147.1444 | 497.0188
0.67029 | 995585.9127 | 70.761 | 170.8321 | 490.7384
0.69326 | 1054392.1571 | 93.0769 | 224.7076 | 500.2917
0.7121 | 1088485.3278 | 84.613 | 204.2738 | 504.8244
0.69356 | 1057529.3583 | 79.5759 | 192.1133 | 492.6615
0.68261 | 1027575.1776 | 83.3282 | 201.1722 | 511.2359
0.69234 | 1039283.3096 | 83.0935 | 200.6054 | 472.4395
0.67117 | 987859.3765 | 71.8515 | 173.4648 | 496.9432
0.68423 | 1025491.8698 | 69.5935 | 168.0135 | 509.0271
0.68361 | 1019307.5736 | 77.8166 | 187.866 | 481.3272
0.67143 | 1007538.6338 | 76.6645 | 185.0845 | 516.8309
0.6756 | 1006607.9548 | 69.184 | 167.025 | 476.7451
0.68791 | 1062983.8836 | 63.7519 | 153.9107 | 464.7174
0.70122 | 1077932.7909 | 80.2251 | 193.6806 | 505.2135
0.68529 | 1028595.6949 | 67.8372 | 163.7735 | 491.3341
0.68859 | 1032582.9842 | 84.5311 | 204.0761 | 498.5375
0.68293 | 1026292.2623 | 67.9193 | 163.9716 | 508.2655
0.6505 | 953237.4938 | 62.4408 | 150.7455 | 496.0952
0.68098 | 1023879.136 | 80.0852 | 193.3427 | 502.7691
0.66715 | 975601.8576 | 62.3112 | 150.4325 | 481.8714
0.69769 | 1041869.8787 | 65.5894 | 158.3467 | 494.0486
0.67964 | 1023085.6945 | 75.2161 | 181.5878 | 491.0228
0.6997 | 1079539.284 | 83.8077 | 202.3296 | 450.0048
0.70084 | 1064896.4469 | 66.7381 | 161.12 | 506.6165
0.68363 | 1040207.2713 | 70.6536 | 170.5728 | 481.5836
0.70101 | 1071558.2756 | 79.775 | 192.5938 | 476.758
0.67744 | 1014983.4466 | 87.4229 | 211.0575 | 499.5316
0.67924 | 1021265.9803 | 73.3993 | 177.2016 | 534.8489
0.69277 | 1064309.0714 | 65.463 | 158.0417 | 487.0952
0.69199 | 1043604.4429 | 55.652 | 134.3558 | 466.9868
0.6729 | 992701.2035 | 66.6752 | 160.9682 | 509.9457
0.6873 | 1027366.7897 | 63.6933 | 153.7693 | 474.1208
0.64932 | 961237.3256 | 75.6317 | 182.591 | 463.8094
0.68983 | 1045524.9818 | 64.5818 | 155.9143 | 478.9386
0.69065 | 1039262.4208 | 73.4581 | 177.3436 | 471.4048
0.68819 | 1036051.2925 | 76.9136 | 185.6859 | 483.7757
0.71056 | 1095624.0992 | 77.0441 | 186.0008 | 453.761
0.68946 | 1042608.5112 | 66.4192 | 160.3502 | 495.756
0.69184 | 1042127.5734 | 80.2331 | 193.6997 | 482.0745
0.67617 | 1011485.7801 | 76.4253 | 184.507 | 460.7728
0.66917 | 993819.2986 | 77.5074 | 187.1195 | 511.8099
0.69078 | 1048850.1694 | 64.6982 | 156.1952 | 497.7248
0.71186 | 1117727.2893 | 66.9048 | 161.5226 | 460.8601
0.67986 | 1036301.5069 | 85.9249 | 207.441 | 483.0922
0.70519 | 1068270.0615 | 72.169 | 174.2314 | 506.6357
0.68365 | 1028522.5796 | 67.9952 | 164.1548 | 502.5775
0.68452 | 1029032.1643 | 73.3865 | 177.1707 | 497.7848
0.70281 | 1058685.4006 | 76.9595 | 185.7966 | 478.3452
0.68939 | 1027242.1091 | 85.4838 | 206.3762 | 526.1301
0.71035 | 1088627.1416 | 69.2633 | 167.2165 | 471.5653
0.67825 | 1030373.0846 | 85.6356 | 206.7427 | 463.8324
0.67227 | 1010165.2359 | 57.7125 | 139.3303 | 491.7148
0.67966 | 1018625.4154 | 73.1636 | 176.6327 | 467.0919
0.65915 | 953162.5634 | 72.2533 | 174.4348 | 508.9438
0.69649 | 1054025.0217 | 62.4363 | 150.7344 | 491.6618
0.68008 | 1007118.051 | 63.1014 | 152.3403 | 490.0805
0.67757 | 1015400.7169 | 71.3943 | 172.361 | 504.8272
0.67629 | 1014232.5007 | 87.5069 | 211.2605 | 502.7294
0.66872 | 987759.6518 | 73.8042 | 178.1792 | 494.9569
0.68443 | 1020106.5142 | 84.0179 | 202.8371 | 479.7397
0.69833 | 1079292.161 | 82.3226 | 198.7444 | 477.1757
0.69936 | 1070816.0577 | 72.9445 | 176.1036 | 478.75
0.67912 | 1024859.6939 | 80.3821 | 194.0595 | 489.2653
0.71072 | 1102397.0654 | 57.2374 | 138.1834 | 478.2735
0.69224 | 1040834.4062 | 62.0332 | 149.7614 | 496.8903
0.69034 | 1051809.4384 | 81.1543 | 195.9238 | 491.1652
0.67996 | 1014741.7024 | 79.4898 | 191.9053 | 491.9058
0.70891 | 1078135.3751 | 68.7346 | 165.94 | 456.4019
0.68284 | 1011148.8305 | 72.2881 | 174.5189 | 491.7351
0.6839 | 1051349.7349 | 76.8118 | 185.4402 | 482.7844
0.6813 | 1026238.9404 | 81.8218 | 197.5352 | 492.4413
0.70825 | 1074890.5775 | 75.8322 | 183.0751 | 502.8067
0.6795 | 1012157.527 | 70.9639 | 171.3219 | 504.2038
0.67315 | 1000420.0159 | 69.009 | 166.6024 | 508.2907
0.64305 | 921991.1827 | 68.3061 | 164.9054 | 502.7164
0.68654 | 1024811.708 | 76.2616 | 184.1119 | 492.5523
0.67263 | 1002793.468 | 80.9824 | 195.5088 | 517.9901
0.7025 | 1070333.1165 | 72.5855 | 175.2368 | 501.3859
0.66582 | 993370.8438 | 67.4293 | 162.7888 | 487.7351
0.70941 | 1086376.281 | 69.2119 | 167.0924 | 489.9636
0.6636 | 975072.9368 | 76.6834 | 185.1302 | 467.8039
0.68266 | 1033308.4705 | 80.6236 | 194.6426 | 469.9211
0.70959 | 1082120.5861 | 62.9171 | 151.8953 | 470.7898
0.68921 | 1051539.7394 | 60.8121 | 146.8134 | 485.3245
0.67204 | 1003894.7574 | 74.6399 | 180.1966 | 472.241
0.68685 | 1044841.9573 | 64.8193 | 156.4876 | 504.1987
0.68858 | 1043905.661 | 80.7505 | 194.949 | 483.4094
0.67084 | 1009950.6052 | 69.9071 | 168.7707 | 485.0543
0.68987 | 1038136.9852 | 64.1214 | 154.8027 | 512.9535
0.66653 | 996231.4131 | 84.2542 | 203.4076 | 508.0602
0.69882 | 1053365.7409 | 66.2063 | 159.8361 | 499.723
0.67579 | 991442.7499 | 66.3061 | 160.0772 | 481.6864
0.71284 | 1097126.4334 | 65.5811 | 158.3268 | 465.3218
0.69129 | 1034067.4273 | 63.7392 | 153.88 | 500.7161
0.68591 | 1033271.5627 | 55.589 | 134.2037 | 479.5072
0.69076 | 1062550.2907 | 61.3116 | 148.0194 | 527.0154
0.68064 | 1028239.4842 | 66.3528 | 160.1899 | 520.0178
0.71463 | 1088638.239 | 60.7183 | 146.5869 | 475.1477
0.67822 | 1029019.9214 | 75.8109 | 183.0237 | 496.302
0.70359 | 1076650.7487 | 78.1784 | 188.7395 | 504.2364
0.70928 | 1082548.1286 | 72.7117 | 175.5416 | 486.0931
0.68843 | 1032351.4804 | 69.6314 | 168.1051 | 484.9339
0.71206 | 1083498.2578 | 74.2825 | 179.3338 | 459.5603
0.71521 | 1109748.596 | 84.2844 | 203.4804 | 476.2858
0.67981 | 1021112.2898 | 81.5912 | 196.9787 | 489.0375
0.6997 | 1075276.212 | 61.2532 | 147.8782 | 460.5174
0.67992 | 1015912.9959 | 65.6265 | 158.4363 | 511.193
0.69377 | 1044172.678 | 68.0227 | 164.2214 | 503.1148
0.68257 | 1037729.2989 | 68.7125 | 165.8866 | 505.8883
0.67474 | 1000645.5477 | 74.1008 | 178.8951 | 483.8995
0.68532 | 1018905.869 | 70.8297 | 170.998 | 482.5844
0.68192 | 1013037.3994 | 59.2118 | 142.95 | 489.9802
0.67201 | 1023872.7854 | 70.7467 | 170.7977 | 515.236
0.68951 | 1068080.0894 | 75.6196 | 182.5618 | 507.4358
0.68501 | 1021430.6165 | 61.5195 | 148.5212 | 500.006
0.69424 | 1074740.4983 | 82.458 | 199.0711 | 495.995
0.68447 | 1021808.5199 | 59.5024 | 143.6516 | 454.8172
0.69138 | 1048579.2471 | 85.494 | 206.4008 | 514.0784
0.6889 | 1041517.9095 | 62.3753 | 150.5872 | 485.098
0.66847 | 977390.3844 | 79.4226 | 191.7431 | 488.2568
0.681 | 1017838.7786 | 82.7341 | 199.7379 | 514.5522
0.69372 | 1060718.6801 | 90.6929 | 218.9521 | 497.103
0.71733 | 1086429.1086 | 80.6717 | 194.7588 | 461.9781
0.687 | 1032156.1062 | 78.4081 | 189.2938 | 491.4765
0.71385 | 1084208.2098 | 77.5955 | 187.3321 | 525.5136
0.66205 | 976035.8106 | 76.9391 | 185.7475 | 497.2028
0.68335 | 1045245.1298 | 65.1769 | 157.3509 | 497.8883
0.67078 | 1009542.1533 | 73.4718 | 177.3767 | 512.04
0.70993 | 1079224.1862 | 74.2112 | 179.1617 | 503.1259
0.68003 | 1019998.4682 | 81.4754 | 196.6991 | 503.5366
0.67698 | 993541.9366 | 69.9176 | 168.796 | 501.2419
0.67726 | 1018466.4965 | 71.6423 | 172.9597 | 507.086
0.67791 | 1007084.8031 | 74.2282 | 179.2028 | 486.6996
0.68469 | 1020404.9131 | 67.3552 | 162.6099 | 501.2966
0.6968 | 1047952.9898 | 64.0844 | 154.7133 | 468.4362
0.70111 | 1066587.1981 | 74.2648 | 179.2911 | 506.2981
0.69879 | 1055688.0449 | 71.9725 | 173.757 | 529.1667
0.70252 | 1092037.7032 | 70.2548 | 169.6102 | 531.4692
0.69584 | 1058847.6307 | 70.6581 | 170.5838 | 502.6459
0.71082 | 1107402.9349 | 72.3404 | 174.6451 | 494.4844
0.68866 | 1033376.1028 | 69.6777 | 168.2168 | 509.7466
0.6934 | 1057916.3563 | 70.8447 | 171.0343 | 463.974
0.73471 | 1146640.9541 | 79.7337 | 192.4941 | 464.5047
0.69178 | 1044172.3038 | 68.2183 | 164.6936 | 506.1059
0.65652 | 955479.1727 | 83.4919 | 201.5672 | 509.1051
0.67109 | 1010728.4309 | 83.3693 | 201.2714 | 510.5902
0.68231 | 1015455.162 | 72.653 | 175.3999 | 478.631
0.70578 | 1088240.2266 | 71.5966 | 172.8494 | 488.8991
0.66567 | 985656.0802 | 72.2839 | 174.5087 | 452.1307
0.66568 | 1004411.0374 | 81.9602 | 197.8695 | 498.5847
0.68119 | 1008721.8771 | 75.6545 | 182.6461 | 487.9973
0.68521 | 1042035.153 | 81.0029 | 195.5582 | 505.1445
0.689 | 1052945.1411 | 74.6036 | 180.1091 | 510.6028
0.68974 | 1043529.5409 | 75.5951 | 182.5027 | 512.1432
0.69421 | 1053086.5669 | 68.5595 | 165.5173 | 490.8792
0.69129 | 1028866.4749 | 77.6735 | 187.5203 | 476.8824
0.69944 | 1052714.5882 | 66.1064 | 159.595 | 481.2336
0.69771 | 1052622.3331 | 77.0779 | 186.0825 | 489.9347
0.67295 | 1008048.0314 | 81.3683 | 196.4403 | 517.6639
0.68549 | 1017749.369 | 63.2643 | 152.7335 | 480.3949
0.69491 | 1040334.6379 | 74.8031 | 180.5906 | 471.8728
0.68999 | 1047155.5021 | 62.6139 | 151.1632 | 523.9934
0.67473 | 1007358.3783 | 77.9164 | 188.1067 | 496.0827
0.67168 | 997854.0993 | 61.2923 | 147.9726 | 501.3747
0.69056 | 1050779.6093 | 85.7755 | 207.0804 | 474.9292
0.67864 | 1007801.2689 | 63.2776 | 152.7655 | 466.1757
0.68529 | 1026416.4412 | 85.9255 | 207.4425 | 473.7956
0.7171 | 1114233.6929 | 82.7633 | 199.8084 | 493.2376
0.70712 | 1087496.5506 | 76.8439 | 185.5176 | 492.7025
0.71361 | 1093530.3941 | 81.6221 | 197.0533 | 498.2328
0.69244 | 1058722.2308 | 81.4404 | 196.6146 | 472.9911
0.70029 | 1056383.0168 | 86.5399 | 208.9259 | 479.6776
0.68528 | 1027116.2296 | 67.7404 | 163.5398 | 512.7706
0.68159 | 1010530.7219 | 74.0805 | 178.8461 | 500.4
0.67995 | 1039242.7953 | 63.0583 | 152.2363 | 506.4164
0.69257 | 1049216.2424 | 75.6968 | 182.7483 | 483.3578
0.71582 | 1097805.9499 | 67.3604 | 162.6224 | 484.6737
0.68427 | 1023544.0572 | 74.7403 | 180.4391 | 487.5249
0.68004 | 1020115.1748 | 74.9533 | 180.9532 | 512.6157
0.70549 | 1090188.3688 | 73.619 | 177.732 | 491.3604
0.70953 | 1082701.7494 | 74.4927 | 179.8412 | 519.1422
0.68219 | 1036300.163 | 79.2783 | 191.3948 | 502.3406
0.70349 | 1076442.7174 | 81.5029 | 196.7653 | 516.4479
0.68574 | 1036553.3396 | 73.7919 | 178.1493 | 507.2756
0.67451 | 998731.3873 | 73.9837 | 178.6126 | 497.913
0.67211 | 1008804.5941 | 84.6093 | 204.2649 | 488.5819
0.66392 | 980955.6595 | 62.9583 | 151.9947 | 466.0997
0.68888 | 1033636.3957 | 63.9355 | 154.3541 | 473.3873
0.71123 | 1091907.0569 | 63.3396 | 152.9153 | 488.0554
0.68848 | 1036130.7454 | 76.8866 | 185.6207 | 493.5298
0.68549 | 1016005.3197 | 70.6113 | 170.4708 | 478.0865
0.6867 | 1051840.4005 | 79.9383 | 192.9882 | 470.4602
0.67429 | 1007593.4371 | 76.3433 | 184.3089 | 486.3654
0.68533 | 1025834.07 | 69.5051 | 167.8002 | 475.9247
0.71336 | 1104782.7643 | 85.7278 | 206.9653 | 487.2013
0.72577 | 1128720.6914 | 81.1638 | 195.9469 | 449.8121
0.69105 | 1039321.2808 | 67.9728 | 164.1008 | 494.5191
0.6552 | 971396.3001 | 68.8434 | 166.2028 | 492.2732
0.67801 | 1002543.6524 | 68.6597 | 165.7593 | 499.5264
0.66176 | 976711.1724 | 62.8905 | 151.8311 | 514.9743
0.70294 | 1073902.0762 | 65.0977 | 157.1597 | 486.0353
0.67813 | 1005853.0979 | 83.6126 | 201.8586 | 479.3499
0.6869 | 1036228.497 | 71.8685 | 173.5058 | 508.1024
0.70607 | 1078855.498 | 70.9501 | 171.2888 | 475.5981
0.70174 | 1079126.9954 | 80.2549 | 193.7525 | 465.3298
0.70171 | 1073357.9758 | 74.672 | 180.2741 | 500.9087
0.71272 | 1087103.8608 | 85.982 | 207.5789 | 459.8652
0.70729 | 1081747.8006 | 79.8187 | 192.6995 | 524.3182
0.68621 | 1034963.0095 | 73.6274 | 177.7523 | 485.3376
0.71544 | 1103766.1065 | 71.6377 | 172.9486 | 473.6564
0.69879 | 1068321.7551 | 92.1317 | 222.4255 | 478.9047
0.69372 | 1049529.8233 | 78.9556 | 190.6156 | 487.3614
0.69534 | 1071236.4071 | 88.3818 | 213.3725 | 463.6281
0.70906 | 1085161.2969 | 65.212 | 157.4356 | 498.9068
0.68895 | 1035936.143 | 81.6778 | 197.1876 | 496.9949
0.70568 | 1062683.1152 | 75.2445 | 181.6562 | 485.2438
0.70221 | 1062311.467 | 64.4621 | 155.6252 | 502.2028
0.66058 | 988673.263 | 74.7836 | 180.5436 | 517.7981
0.69441 | 1038533.2282 | 76.4074 | 184.4638 | 498.0545
0.68223 | 1012393.3849 | 52.4283 | 126.573 | 506.0968
0.69924 | 1058094.0459 | 72.0896 | 174.0397 | 485.8868
0.69778 | 1057072.7113 | 80.6914 | 194.8062 | 497.4474
0.71036 | 1085450.5221 | 69.0598 | 166.7251 | 469.5698
0.67574 | 1005169.0675 | 70.8661 | 171.0859 | 461.5398
0.69427 | 1030311.8681 | 70.6999 | 170.6846 | 506.846
0.67852 | 1006004.8555 | 68.8453 | 166.2073 | 498.7811
0.6738 | 1002119.3013 | 83.8379 | 202.4027 | 519.9822
0.70242 | 1055838.4636 | 78.4472 | 189.3882 | 462.4292
0.65788 | 968858.8648 | 60.76 | 146.6877 | 481.2239
0.6728 | 1015077.6062 | 95.0887 | 229.5644 | 489.1705
0.68286 | 1035943.9959 | 82.4067 | 198.9475 | 444.2644
0.69438 | 1059121.0216 | 68.9359 | 166.4259 | 528.4952
0.70004 | 1061302.2775 | 80.7529 | 194.9548 | 479.4972
0.69195 | 1033731.6171 | 61.6362 | 148.8029 | 509.7511
0.68603 | 1028084.167 | 97.1121 | 234.4493 | 479.1905
0.71941 | 1111001.9148 | 78.8984 | 190.4775 | 456.2154
0.69218 | 1054673.2909 | 77.8449 | 187.9342 | 486.6892
0.68175 | 1034057.6912 | 62.5778 | 151.0761 | 497.8972
0.68406 | 1020907.4209 | 82.6129 | 199.4452 | 484.3091
0.69154 | 1041487.666 | 57.8444 | 139.6488 | 498.4975
0.6828 | 1023547.8362 | 67.7504 | 163.564 | 496.5303
0.69856 | 1062579.3045 | 64.8151 | 156.4774 | 470.1128
0.68613 | 1021800.5264 | 77.7195 | 187.6315 | 503.2535
0.68974 | 1025329.7779 | 65.2918 | 157.6283 | 512.0866
0.67303 | 1003655.1132 | 82.5474 | 199.2871 | 488.1431
0.67039 | 982173.3604 | 74.8493 | 180.7023 | 456.1966
0.67153 | 1003717.745 | 57.4073 | 138.5935 | 489.3838
0.69336 | 1045762.2828 | 50.8071 | 122.6593 | 470.9607
0.70894 | 1075450.5438 | 60.4644 | 145.974 | 484.6392
0.7039 | 1063675.4947 | 75.5045 | 182.284 | 464.1689
0.6719 | 982355.3446 | 77.8757 | 188.0085 | 482.8036
0.68645 | 1036239.4386 | 68.2253 | 164.7104 | 511.302
0.68353 | 1029019.2893 | 81.7753 | 197.423 | 472.9086
0.67788 | 1006641.791 | 59.1889 | 142.8946 | 498.5739
0.68935 | 1036527.7646 | 80.0909 | 193.3565 | 476.9781
0.66608 | 1000874.2387 | 84.7043 | 204.4942 | 506.0645
0.71248 | 1077583.5212 | 78.9314 | 190.5572 | 500.5791
0.67888 | 1023285.7174 | 77.0075 | 185.9126 | 471.0131
0.66096 | 975827.7297 | 80.1965 | 193.6115 | 499.3168
0.66601 | 996666.7676 | 57.4766 | 138.7608 | 462.2345
0.66719 | 992462.9938 | 65.335 | 157.7326 | 497.43
0.70791 | 1086151.2047 | 65.3772 | 157.8345 | 531.7012
0.68585 | 1029635.1598 | 73.2876 | 176.9319 | 496.8785
0.6859 | 1045668.7181 | 79.5628 | 192.0815 | 496.9147
0.69813 | 1069239.6102 | 76.5274 | 184.7534 | 501.1324
0.68773 | 1027160.6757 | 81.129 | 195.8627 | 480.9918
0.70984 | 1074206.482 | 64.4828 | 155.6752 | 514.8396
0.68317 | 1023322.0351 | 78.1728 | 188.7259 | 487.9989
0.70675 | 1064727.6628 | 66.1695 | 159.7473 | 486.6279
0.68421 | 1024865.8778 | 67.6975 | 163.4363 | 475.3677
0.70253 | 1063092.2208 | 77.3097 | 186.642 | 493.4948
0.69176 | 1049919.0527 | 90.0237 | 217.3364 | 477.697
0.69773 | 1048885.2944 | 75.113 | 181.3389 | 494.9676
0.68763 | 1045037.7953 | 60.5753 | 146.2416 | 476.8674
0.68899 | 1039097.5045 | 75.3561 | 181.9258 | 475.5532
0.69683 | 1049500.5162 | 57.8473 | 139.6557 | 485.4273
0.70752 | 1068771.2651 | 83.6179 | 201.8716 | 464.8939
0.67218 | 1000946.1668 | 76.0777 | 183.6678 | 508.9423
0.68777 | 1038200.221 | 66.3013 | 160.0654 | 489.715
0.68707 | 1026672.4265 | 61.65 | 148.8363 | 482.3185
0.69564 | 1050553.8433 | 66.3306 | 160.1363 | 518.5262
0.65283 | 958462.646 | 87.083 | 210.237 | 526.1437
0.6836 | 1037938.5009 | 72.3349 | 174.6318 | 517.544
0.67643 | 994484.2041 | 78.3095 | 189.0558 | 466.4416
0.69291 | 1048115.794 | 75.5802 | 182.4667 | 481.2272
0.68188 | 1018595.7759 | 74.6401 | 180.1972 | 461.6213
0.69242 | 1042606.5145 | 64.1604 | 154.8969 | 489.7973
0.67922 | 1000730.6196 | 84.5239 | 204.0588 | 482.4054
0.70097 | 1050639.6729 | 67.7404 | 163.5397 | 483.3623
0.70337 | 1077822.8057 | 76.5666 | 184.848 | 486.5703
0.68594 | 1027848.397 | 76.5551 | 184.8203 | 481.2323
0.67052 | 1005065.6962 | 76.5793 | 184.8788 | 484.5041
0.69001 | 1043395.6123 | 85.2405 | 205.7887 | 499.3714
0.69432 | 1057907.5961 | 64.9235 | 156.7392 | 511.8564
0.70149 | 1087728.5933 | 75.696 | 182.7464 | 505.4747
0.6773 | 998267.9752 | 70.2175 | 169.52 | 483.206
0.69966 | 1069408.3468 | 74.616 | 180.139 | 479.3998
0.69295 | 1053636.8317 | 64.7896 | 156.4158 | 486.2077
0.69049 | 1056918.9317 | 65.2229 | 157.462 | 495.6275
0.68288 | 1046220.5596 | 69.243 | 167.1673 | 504.6782
0.68513 | 1016096.4618 | 68.8656 | 166.2562 | 488.3914
0.65826 | 972444.7529 | 70.379 | 169.91 | 490.9854
0.68266 | 1028742.6297 | 75.2972 | 181.7836 | 522.2682
0.71176 | 1078004.4549 | 75.6773 | 182.701 | 508.4511
0.68647 | 1045314.6053 | 72.5441 | 175.1369 | 482.7745
0.6795 | 1030881.7607 | 72.276 | 174.4898 | 471.4458
0.67148 | 995167.8972 | 65.0948 | 157.1527 | 472.4151
0.66925 | 1003461.2477 | 74.9181 | 180.8683 | 513.5019
0.69147 | 1042594.875 | 70.0135 | 169.0275 | 516.2952
0.69456 | 1029334.8034 | 56.9206 | 137.4185 | 487.5008
0.68383 | 1032384.7863 | 71.923 | 173.6375 | 489.4978
0.68677 | 1036161.3652 | 73.6531 | 177.8143 | 477.0423
0.69598 | 1036773.8082 | 65.429 | 157.9595 | 478.0252
0.66791 | 982873.2088 | 73.5289 | 177.5145 | 482.1593
0.69458 | 1065707.5936 | 76.6636 | 185.0824 | 490.548
0.67546 | 1007319.9242 | 76.1901 | 183.9391 | 515.9458
0.68228 | 1010897.0124 | 75.6966 | 182.7476 | 505.4647
0.68725 | 1046476.4985 | 77.7938 | 187.8108 | 475.8653
0.70211 | 1061697.0953 | 70.2153 | 169.5147 | 481.1441
0.68836 | 1037844.1107 | 80.3071 | 193.8785 | 491.6474
0.68458 | 1014492.0222 | 72.3429 | 174.6511 | 500.37
0.72358 | 1136020.4409 | 74.7961 | 180.5738 | 491.5615
0.70679 | 1077310.6617 | 76.16 | 183.8665 | 509.0895
0.71793 | 1122397.0155 | 72.5938 | 175.2569 | 501.4213
0.68401 | 1028926.8769 | 73.4341 | 177.2855 | 522.8205
0.69128 | 1024589.1518 | 77.3511 | 186.7421 | 470.6862
0.70494 | 1074188.5447 | 84.3497 | 203.6383 | 468.4265
0.70106 | 1066151.6027 | 85.7271 | 206.9635 | 502.1423
0.68351 | 1034297.7407 | 73.8706 | 178.3395 | 522.1898
0.71076 | 1093406.415 | 75.0986 | 181.3042 | 488.9928
0.68244 | 1013437.4146 | 79.0996 | 190.9632 | 501.0769
0.69113 | 1048296.3161 | 70.4304 | 170.0341 | 491.6762
0.6877 | 1041656.7115 | 80.6796 | 194.7777 | 489.437
0.6893 | 1036232.2234 | 73.3878 | 177.1738 | 503.4513
0.70561 | 1082835.7494 | 65.3902 | 157.8659 | 467.6836
0.69113 | 1034323.1218 | 61.744 | 149.0632 | 513.3889
0.67816 | 1023335.7056 | 81.9679 | 197.888 | 538.072
0.7011 | 1061791.7756 | 60.364 | 145.7317 | 508.379
0.7053 | 1063036.7474 | 68.1195 | 164.4549 | 508.6402
0.66352 | 971444.2192 | 67.9835 | 164.1267 | 520.4426
0.70506 | 1067854.9989 | 66.5222 | 160.5989 | 480.3677
0.69715 | 1052202.3222 | 70.0823 | 169.1937 | 488.2723
0.67336 | 988189.9506 | 76.5894 | 184.9031 | 480.6992
0.70998 | 1097057.685 | 57.0197 | 137.6577 | 482.3634
0.69187 | 1043519.4532 | 80.3368 | 193.9502 | 487.5283
0.68159 | 1021276.1359 | 81.3907 | 196.4946 | 482.0201
0.69436 | 1038392.1753 | 69.4412 | 167.6458 | 499.8071
0.71548 | 1098476.5036 | 65.4201 | 157.9381 | 463.2433
0.70429 | 1079681.5455 | 76.5987 | 184.9256 | 465.646
0.71663 | 1103996.6627 | 65.6427 | 158.4755 | 514.3352
0.69228 | 1046302.8817 | 70.1801 | 169.4297 | 486.8664
0.6851 | 1025646.0118 | 85.8845 | 207.3436 | 481.1578
0.69561 | 1067194.6376 | 95.3021 | 230.0795 | 474.5817
0.69024 | 1055247.7012 | 79.7865 | 192.6216 | 500.4172
0.69435 | 1051727.1825 | 70.3337 | 169.8006 | 488.9183
0.69104 | 1045788.7506 | 77.3167 | 186.659 | 513.7491
0.69536 | 1055373.1356 | 77.5398 | 187.1976 | 490.9451
0.66585 | 990667.8406 | 70.8812 | 171.1224 | 493.5366
0.67885 | 1000750.0369 | 75.3153 | 181.8273 | 492.526
0.7058 | 1083027.6564 | 54.2128 | 130.8813 | 447.5456
0.6805 | 1012424.8458 | 65.3459 | 157.759 | 503.5365
0.69561 | 1068665.8247 | 74.9983 | 181.0619 | 498.1154
0.67731 | 1019245.0521 | 66.0955 | 159.5687 | 515.2481
0.69741 | 1054548.6388 | 66.414 | 160.3376 | 493.8321
0.67734 | 1007008.8988 | 79.1891 | 191.1793 | 484.4633
0.70093 | 1059553.7496 | 91.0969 | 219.9274 | 478.5789
0.69445 | 1060644.0015 | 74.5593 | 180.0021 | 515.2944
0.70934 | 1096200.3578 | 83.9443 | 202.6595 | 534.8068
0.70031 | 1064859.1049 | 91.6341 | 221.2244 | 499.8998
0.69241 | 1044896.0276 | 64.3021 | 155.2389 | 497.6925
0.69502 | 1049284.0375 | 71.9236 | 173.639 | 488.7111
0.69404 | 1063948.3607 | 73.0319 | 176.3147 | 475.0647
0.68975 | 1039134.9958 | 62.1558 | 150.0574 | 486.0807
0.69997 | 1062071.2347 | 79.186 | 191.1719 | 480.7018
0.69659 | 1065000.7563 | 72.8367 | 175.8433 | 514.52
0.70678 | 1061184.4939 | 71.5297 | 172.6879 | 459.4079
0.68971 | 1037583.6236 | 64.1018 | 154.7555 | 511.3389
0.6954 | 1063399.4361 | 67.1656 | 162.1522 | 465.529
0.72435 | 1144911.4452 | 87.494 | 211.2292 | 479.6699
0.6769 | 1002535.8338 | 59.0331 | 142.5186 | 523.4105
0.69465 | 1057527.8608 | 68.8057 | 166.1116 | 475.1768
0.68413 | 1020326.5672 | 67.4061 | 162.7326 | 496.0637
0.70661 | 1104161.0921 | 76.2831 | 184.1637 | 514.1159
0.67609 | 1007798.3456 | 74.4813 | 179.8137 | 482.4284
0.68556 | 1035773.5424 | 73.7551 | 178.0605 | 482.6766
0.68526 | 1009555.8718 | 60.7823 | 146.7414 | 491.9615
0.69595 | 1051508.9147 | 68.2859 | 164.8567 | 495.3063
0.68984 | 1047970.0169 | 82.5198 | 199.2204 | 512.1016
0.71914 | 1101242.3485 | 76.5148 | 184.723 | 497.0024
0.6989 | 1064861.3762 | 77.8415 | 187.9259 | 500.7896
0.67374 | 1015668.4682 | 77.0179 | 185.9376 | 539.9071
0.69142 | 1052535.6827 | 79.4454 | 191.7982 | 483.6003
0.6784 | 1017126.7931 | 59.7594 | 144.2719 | 514.1198
0.69121 | 1040163.6458 | 69.8766 | 168.697 | 483.8456
0.70804 | 1075300.2013 | 74.139 | 178.9873 | 464.7978
0.68507 | 1028121.6681 | 69.6867 | 168.2385 | 488.7463
0.6937 | 1056688.7731 | 82.7436 | 199.7608 | 511.9926
0.69513 | 1045031.3823 | 73.3903 | 177.1798 | 493.134
0.69117 | 1051807.7002 | 79.4192 | 191.735 | 512.0476
0.70359 | 1077316.2266 | 83.1724 | 200.796 | 482.9627
0.69266 | 1040333.4152 | 73.3207 | 177.0118 | 478.4411
0.6983 | 1046612.0085 | 82.9653 | 200.2959 | 473.7572
0.69204 | 1033046.168 | 74.5869 | 180.0687 | 490.8177
0.67518 | 1003252.6867 | 80.7139 | 194.8606 | 474.8296
0.6945 | 1063547.744 | 71.1091 | 171.6726 | 480.8675
0.70645 | 1075097.296 | 74.2912 | 179.3548 | 471.0392
0.71198 | 1109892.467 | 76.1007 | 183.7234 | 497.3337
0.66861 | 972833.2958 | 60.0299 | 144.9251 | 472.6774
0.70104 | 1069450.7218 | 72.6771 | 175.4581 | 517.3571
0.69708 | 1058450.2622 | 84.0514 | 202.9179 | 511.905
0.69025 | 1045496.2875 | 75.6824 | 182.7136 | 497.6314
0.70275 | 1075008.7258 | 67.1884 | 162.207 | 490.0473
0.7067 | 1079035.3387 | 70.8992 | 171.1657 | 494.8483
0.71561 | 1101813.3676 | 67.6259 | 163.2633 | 483.862
0.69512 | 1063016.1224 | 58.7359 | 141.8009 | 471.2347
0.69023 | 1042719.3866 | 85.1581 | 205.5899 | 484.9272
0.69606 | 1035706.1252 | 70.8874 | 171.1373 | 466.7196
0.71914 | 1111730.7558 | 71.7832 | 173.3 | 479.92
0.68652 | 1030038.6452 | 74.9681 | 180.9889 | 488.1028
0.70261 | 1070376.0692 | 53.8232 | 129.9408 | 482.7499
0.69213 | 1053845.6752 | 68.8148 | 166.1337 | 486.8031
0.68125 | 1013965.8045 | 76.0408 | 183.5788 | 516.595
0.68967 | 1042784.2436 | 75.5493 | 182.3921 | 502.3346
0.67926 | 1029736.4848 | 89.9182 | 217.0817 | 484.8881
0.68687 | 1044299.5957 | 75.9675 | 183.4019 | 464.8833
0.70797 | 1080034.5337 | 74.4721 | 179.7915 | 499.3778
0.71979 | 1118220.8389 | 81.0765 | 195.7359 | 498.356
0.68787 | 1011574.8588 | 81.4043 | 196.5273 | 470.5571
0.68544 | 1035294.4777 | 78.8263 | 190.3034 | 488.8928
0.68547 | 1018375.2748 | 86.6878 | 209.2828 | 485.3342
0.67901 | 1013653.9019 | 77.9095 | 188.0901 | 508.535
0.68189 | 1019542.7837 | 81.2175 | 196.0765 | 483.2903
0.68137 | 1029877.2247 | 64.4054 | 155.4884 | 477.4803
0.6918 | 1048222.6938 | 79.0761 | 190.9066 | 483.5967
0.67287 | 997186.6155 | 70.6027 | 170.4499 | 499.1753
0.70371 | 1082249.5234 | 75.0771 | 181.2521 | 446.8575
0.66649 | 994453.1003 | 76.229 | 184.0331 | 522.8298
0.68634 | 1009804.2529 | 85.8491 | 207.258 | 488.534
0.70035 | 1061297.2816 | 89.884 | 216.9993 | 488.7904
0.68441 | 1043063.8092 | 68.8799 | 166.2907 | 505.674
0.69861 | 1084768.3047 | 81.0357 | 195.6374 | 469.7426
0.69728 | 1051818.207 | 78.7243 | 190.0573 | 479.3495
0.6755 | 1013466.5794 | 69.6664 | 168.1895 | 472.4567
0.68286 | 1015923.3703 | 63.3938 | 153.0461 | 489.109
0.68564 | 1020330.0459 | 57.1566 | 137.9883 | 478.3973
0.70124 | 1068353.898 | 79.5098 | 191.9536 | 509.2725
0.6983 | 1072177.7245 | 74.9545 | 180.9562 | 500.0817
0.69849 | 1060185.5349 | 91.4688 | 220.8252 | 486.0721
0.6992 | 1074312.0676 | 73.1997 | 176.7196 | 494.8627
0.69588 | 1066878.0502 | 90.9372 | 219.5419 | 499.7169
0.67621 | 1002515.1807 | 70.6301 | 170.5161 | 467.0338
0.69066 | 1038089.6738 | 80.059 | 193.2794 | 482.8316
0.6827 | 1020926.2887 | 61.6968 | 148.9492 | 470.7752
0.69547 | 1055160.0904 | 74.0649 | 178.8085 | 490.7036
0.71242 | 1096984.0999 | 82.0165 | 198.0054 | 491.9393
0.68035 | 1018193.8267 | 84.3628 | 203.6698 | 458.1426
0.67762 | 1006632.9172 | 79.1271 | 191.0297 | 489.5222
0.69936 | 1053640.1337 | 57.865 | 139.6984 | 487.3018
0.6899 | 1059711.0512 | 71.3243 | 172.1922 | 472.2084
0.68754 | 1029259.3975 | 59.6683 | 144.052 | 492.4574
0.70853 | 1071933.6763 | 62.4179 | 150.6901 | 516.1817
0.69928 | 1071783.5549 | 78.3043 | 189.0432 | 486.2243
0.69845 | 1052907.7073 | 77.0167 | 185.9347 | 472.7229
0.69985 | 1072137.0634 | 79.9506 | 193.0178 | 471.1107
0.69265 | 1041578.3192 | 81.6843 | 197.2033 | 522.4034
0.68787 | 1056867.6457 | 80.1838 | 193.5808 | 482.5639
0.66891 | 1002119.2196 | 79.7974 | 192.6479 | 482.9367
0.7151 | 1103066.4259 | 61.4829 | 148.4328 | 513.2297
0.69727 | 1064553.0663 | 69.0022 | 166.5861 | 499.9505
0.6742 | 1004707.2417 | 74.4602 | 179.7628 | 491.8108
0.70393 | 1069577.0266 | 82.5354 | 199.2581 | 474.1571
0.67914 | 1020641.6019 | 50.6246 | 122.2185 | 499.7875
0.68742 | 1035740.5197 | 73.5536 | 177.5742 | 480.5725
0.68829 | 1039263.4767 | 77.1406 | 186.2338 | 485.7201
0.67808 | 1017148.7491 | 57.6141 | 139.0927 | 499.2199
0.68955 | 1039381.7571 | 71.8099 | 173.3644 | 470.3577
0.69316 | 1041399.387 | 83.1643 | 200.7765 | 508.5388
0.69936 | 1069873.8249 | 78.1555 | 188.6841 | 493.6034
0.69746 | 1067698.7955 | 69.1867 | 167.0315 | 491.6597
0.68817 | 1044221.8298 | 82.4245 | 198.9904 | 471.6968
0.67858 | 1020386.016 | 69.8237 | 168.5694 | 479.4907
0.69528 | 1058190.7358 | 82.6732 | 199.5907 | 487.3509
0.68381 | 1010986.7055 | 66.8195 | 161.3166 | 462.9367
0.66027 | 987929.5007 | 74.7618 | 180.4908 | 501.6273
0.67443 | 1009700.4779 | 71.3238 | 172.191 | 504.7755
0.68247 | 1019127.999 | 67.4905 | 162.9365 | 485.5792
0.68858 | 1046660.6634 | 84.1404 | 203.1328 | 446.2905
0.70651 | 1072157.7408 | 66.6531 | 160.9149 | 482.232
0.69347 | 1056020.8565 | 62.8273 | 151.6785 | 470.1
0.70259 | 1070482.6746 | 70.7481 | 170.8009 | 472.5372
0.69409 | 1041487.8949 | 78.3095 | 189.0558 | 494.5998
0.72771 | 1126457.8095 | 82.451 | 199.0542 | 484.0884
0.68476 | 1032471.4005 | 96.3779 | 232.677 | 505.7685
0.68119 | 1024705.7119 | 67.9876 | 164.1366 | 486.0352
0.70094 | 1060295.9146 | 72.0647 | 173.9795 | 493.703
0.68512 | 1026759.3702 | 63.8589 | 154.1689 | 471.196
0.67614 | 1010016.8731 | 74.6392 | 180.195 | 490.402
0.67332 | 995235.5363 | 93.1587 | 224.9051 | 539.2385
0.66944 | 1008654.462 | 76.9776 | 185.8403 | 490.2959
0.6876 | 1039722.611 | 75.5543 | 182.4041 | 484.0709
0.70032 | 1075753.9976 | 73.1805 | 176.6734 | 496.6957
0.67551 | 1000358.2071 | 81.9743 | 197.9035 | 475.014
0.68083 | 1020724.6473 | 73.1886 | 176.6928 | 508.7495
0.68205 | 1028791.0101 | 64.2725 | 155.1675 | 497.5253
0.68292 | 1021986.4106 | 69.4884 | 167.7598 | 494.0803
0.67515 | 1014338.0441 | 59.8829 | 144.5702 | 510.2551
0.69754 | 1059898.1861 | 66.7391 | 161.1225 | 482.3097
0.71321 | 1105668.3793 | 57.2074 | 138.1109 | 486.8914
0.67358 | 997777.7721 | 70.5708 | 170.373 | 471.7175
0.6885 | 1035115.3183 | 75.1209 | 181.358 | 471.2762
0.69447 | 1054448.8822 | 70.503 | 170.2093 | 461.5109
0.70146 | 1067519.889 | 72.5023 | 175.036 | 495.7154
0.70082 | 1074012.6131 | 87.9182 | 212.2533 | 517.1233
0.68762 | 1030906.1003 | 78.2432 | 188.8957 | 502.167
0.68985 | 1038473.7732 | 73.488 | 177.4157 | 495.2751
0.67516 | 999989.9598 | 66.9398 | 161.607 | 491.5973
0.68289 | 1016592.4834 | 82.4048 | 198.9427 | 496.3443
0.67561 | 1014803.1957 | 67.8861 | 163.8915 | 509.3498
0.67074 | 1006689.1746 | 69.6913 | 168.2498 | 497.4456
0.68223 | 1033275.6883 | 83.8375 | 202.4017 | 504.797
0.67048 | 982328.8183 | 71.5229 | 172.6716 | 508.9521
0.70248 | 1057773.1884 | 78.5835 | 189.7174 | 475.3001
0.67923 | 1000160.2141 | 73.5744 | 177.6244 | 516.873
0.71826 | 1101541.7576 | 78.3523 | 189.1592 | 490.932
0.71367 | 1103697.6582 | 69.8955 | 168.7426 | 503.3173
0.70992 | 1074522.4174 | 68.4069 | 165.1489 | 484.401
0.69028 | 1041841.087 | 70.0829 | 169.1951 | 518.237
0.71844 | 1105916.7658 | 65.5912 | 158.3512 | 481.1453
0.69179 | 1042035.643 | 81.821 | 197.5333 | 507.6555
0.69104 | 1054963.684 | 68.9967 | 166.5727 | 519.128
0.66125 | 967134.9255 | 73.0803 | 176.4314 | 510.7653
0.70914 | 1086999.3026 | 81.6469 | 197.113 | 511.0642
0.68719 | 1033421.048 | 80.1018 | 193.3829 | 478.026
0.67517 | 1024715.4376 | 72.3048 | 174.5592 | 501.689
0.71557 | 1100661.8556 | 75.089 | 181.2809 | 522.147
0.66997 | 986523.5391 | 91.4687 | 220.825 | 491.8769
0.67707 | 1014638.3241 | 76.5614 | 184.8356 | 488.0522
0.66747 | 1002628.4976 | 80.0103 | 193.162 | 480.8644
0.68999 | 1024356.8637 | 82.1796 | 198.3991 | 466.8251
0.69341 | 1070841.5251 | 82.2573 | 198.5866 | 500.461
0.70835 | 1092332.9041 | 67.731 | 163.5172 | 511.0917
0.72978 | 1128395.7083 | 77.2589 | 186.5196 | 459.6349
0.67737 | 1009847.1811 | 61.508 | 148.4934 | 482.1845
0.69733 | 1051487.8835 | 77.0367 | 185.9829 | 472.4998
0.71618 | 1114472.8545 | 89.2728 | 215.5235 | 519.3832
0.69802 | 1046380.4535 | 73.1808 | 176.6741 | 487.1688
0.6877 | 1040755.2224 | 72.0989 | 174.062 | 461.1727
0.69223 | 1047673.5074 | 58.1301 | 140.3384 | 497.0606
0.6912 | 1037202.5331 | 86.0867 | 207.8317 | 511.6951
0.68693 | 1040305.0153 | 62.4809 | 150.8423 | 490.5403
0.67328 | 1006701.1789 | 57.429 | 138.6458 | 503.9549
0.69743 | 1065879.6757 | 69.0718 | 166.754 | 512.4124
0.68525 | 1022395.4541 | 59.4687 | 143.5702 | 488.6618
0.69434 | 1058463.1604 | 72.624 | 175.3299 | 477.1806
0.67574 | 988861.4093 | 71.4384 | 172.4675 | 511.9237
0.66406 | 987313.3404 | 69.0372 | 166.6704 | 491.8301
0.68764 | 1032779.4594 | 59.8636 | 144.5235 | 472.3611
0.6861 | 1025152.1047 | 63.6677 | 153.7074 | 494.1305
0.69639 | 1042795.6399 | 71.338 | 172.2252 | 492.6174
0.70776 | 1109396.5382 | 62.4764 | 150.8314 | 503.3543
0.6883 | 1038644.5808 | 77.0944 | 186.1223 | 519.2726
0.6925 | 1039299.4901 | 65.9178 | 159.1397 | 477.4097
0.69379 | 1051364.8818 | 70.3659 | 169.8784 | 492.9406
0.69462 | 1045307.8071 | 71.6928 | 173.0816 | 486.9253
0.68256 | 1024162.0776 | 71.4836 | 172.5767 | 497.6575
0.68232 | 1028448.2746 | 48.1683 | 116.2887 | 498.9894
0.69261 | 1045882.2643 | 64.3447 | 155.3418 | 494.2464
0.67192 | 987452.3379 | 79.8282 | 192.7223 | 495.8748
0.7148 | 1109501.1666 | 49.6031 | 119.7525 | 513.6432
0.70166 | 1059787.057 | 73.7337 | 178.0089 | 497.9751
0.67789 | 1009458.2791 | 72.6947 | 175.5005 | 460.3443
0.70208 | 1090645.3531 | 72.1331 | 174.1448 | 496.2497
0.71287 | 1098005.8414 | 73.3794 | 177.1535 | 508.924
0.71586 | 1109311.3626 | 80.8723 | 195.243 | 512.303
0.69644 | 1065139.1003 | 83.0773 | 200.5664 | 479.1162
0.68621 | 1044386.4828 | 68.7859 | 166.0637 | 501.2741
0.70033 | 1051338.1701 | 69.7075 | 168.2888 | 481.0922
0.68714 | 1031416.7634 | 67.8397 | 163.7796 | 484.8093
0.71871 | 1103116.3944 | 74.9277 | 180.8915 | 512.5054
0.69571 | 1070511.2702 | 78.6131 | 189.7889 | 472.927
0.6809 | 1015007.8514 | 67.3727 | 162.6522 | 466.2925
0.68889 | 1018888.2713 | 68.6778 | 165.8029 | 493.6085
0.69655 | 1061325.8465 | 85.8918 | 207.3612 | 511.8902
0.68895 | 1033196.6411 | 59.7743 | 144.3078 | 478.8068
0.68933 | 1042534.9055 | 72.0508 | 173.946 | 474.8392
0.68747 | 1042854.093 | 77.5001 | 187.1019 | 484.9206
0.68389 | 1030529.1336 | 64.2086 | 155.0133 | 523.8685
0.66398 | 984064.9151 | 59.0723 | 142.6132 | 494.7938
0.66767 | 979557.0306 | 62.5148 | 150.9241 | 478.9175
0.71276 | 1097187.3594 | 86.7744 | 209.4919 | 480.0179
0.68881 | 1052005.3102 | 53.9594 | 130.2695 | 482.6884
0.70594 | 1080420.6272 | 67.2215 | 162.2872 | 506.034
0.68383 | 1026502.9697 | 57.6548 | 139.191 | 471.7197
0.69114 | 1047268.7279 | 62.5165 | 150.9281 | 498.5771
0.67064 | 1007031.8283 | 75.0295 | 181.1373 | 504.552
0.69158 | 1037111.2655 | 67.7767 | 163.6275 | 496.6866
0.67158 | 985401.9849 | 73.3153 | 176.9988 | 486.2767
0.66156 | 982760.9754 | 74.226 | 179.1974 | 485.7584
0.66739 | 991166.5293 | 80.6722 | 194.76 | 478.3517
0.72294 | 1114092.9628 | 51.9119 | 125.3265 | 494.4115
0.68736 | 1027948.2431 | 81.2355 | 196.1199 | 463.2797
0.69244 | 1052998.9435 | 75.9844 | 183.4425 | 485.9724
0.69556 | 1043178.3292 | 63.9057 | 154.282 | 485.4334
0.67607 | 1011075.0928 | 78.0488 | 188.4264 | 502.2061
0.69123 | 1050907.4502 | 82.5845 | 199.3766 | 499.4305
0.6847 | 1040395.8803 | 68.6738 | 165.7933 | 552.0275
0.68395 | 1038413.8183 | 85.5856 | 206.622 | 520.8828
0.71508 | 1109240.4277 | 73.9003 | 178.411 | 515.1662
0.69776 | 1052130.5162 | 75.3428 | 181.8935 | 462.0444
0.67132 | 1002757.958 | 88.598 | 213.8946 | 482.8018
0.71626 | 1114905.5283 | 72.2775 | 174.4932 | 492.6699
0.6835 | 1017947.1503 | 79.3277 | 191.5139 | 487.4987
0.70606 | 1065813.3992 | 70.0932 | 169.2199 | 493.4673
0.69498 | 1051750.9512 | 80.6086 | 194.6064 | 497.5
0.67796 | 1013485.1799 | 74.6947 | 180.329 | 527.1189
0.70403 | 1078713.7906 | 67.2976 | 162.4708 | 514.4835
0.69594 | 1046859.8549 | 71.6626 | 173.0087 | 481.2706
0.7105 | 1099136.35 | 71.8329 | 173.42 | 490.7683
0.68199 | 1026227.3901 | 62.0818 | 149.8787 | 486.3909
0.69951 | 1066571.6001 | 58.5629 | 141.3832 | 493.5355
0.71208 | 1086953.8162 | 70.8797 | 171.1188 | 473.1188
0.68854 | 1039819.3337 | 55.7276 | 134.5383 | 533.188
0.71595 | 1096068.04 | 67.2417 | 162.3359 | 500.7215
0.69217 | 1055667.9376 | 79.5958 | 192.1613 | 479.0357
0.7008 | 1069539.9223 | 77.9714 | 188.2396 | 494.1064
0.68702 | 1031150.8144 | 78.3293 | 189.1036 | 502.6601
0.68977 | 1027608.1921 | 69.1766 | 167.0072 | 495.3337
0.6949 | 1067270.5029 | 85.0166 | 205.2483 | 488.0902
0.66676 | 991036.4753 | 74.2676 | 179.2978 | 517.9153
0.67957 | 998065.3579 | 62.5808 | 151.0835 | 493.5995
0.70402 | 1073241.1668 | 64.9579 | 156.8223 | 476.3401
0.65765 | 958422.4178 | 74.4545 | 179.7491 | 503.2353
0.67837 | 1015533.1293 | 65.7517 | 158.7386 | 496.438
0.68485 | 1027886.8008 | 73.269 | 176.887 | 456.5782
0.71451 | 1076038.4595 | 92.1625 | 222.4999 | 533.6632
0.67376 | 994934.0745 | 49.7438 | 120.0922 | 485.9618
0.71505 | 1108938.8004 | 57.2285 | 138.1618 | 512.2495
0.68179 | 1027711.5532 | 92.2483 | 222.7071 | 488.0161
0.67686 | 999767.85 | 73.8035 | 178.1774 | 486.0604
0.70299 | 1070563.9161 | 64.8907 | 156.6599 | 480.6667
0.68419 | 1006783.7263 | 81.4547 | 196.6489 | 494.2531
0.68536 | 1038221.3363 | 82.287 | 198.6584 | 515.1493
0.67547 | 1011377.6241 | 73.0943 | 176.4652 | 512.0982
0.68264 | 997070.4971 | 72.9815 | 176.193 | 486.7051
0.66668 | 986182.2593 | 66.8556 | 161.4036 | 483.8321
0.67457 | 991812.8735 | 58.6035 | 141.4813 | 494.5721
0.70951 | 1108373.745 | 78.0553 | 188.4421 | 485.9149
0.7037 | 1072441.7817 | 68.8784 | 166.2872 | 472.2875
0.68028 | 1023975.1657 | 76.9014 | 185.6563 | 474.687
0.67423 | 1003825.4598 | 79.3026 | 191.4534 | 509.3603
0.70261 | 1079746.111 | 56.7086 | 136.9067 | 491.2642
0.68476 | 1010855.4911 | 73.7611 | 178.0751 | 488.3074
0.67702 | 985842.0744 | 65.8114 | 158.8828 | 521.5999
0.67719 | 1003204.3526 | 80.4704 | 194.2727 | 486.1416
0.67855 | 1017207.6063 | 67.752 | 163.5677 | 489.191
0.689 | 1042781.6398 | 75.6224 | 182.5687 | 474.8202
0.68895 | 1045185.1242 | 74.0589 | 178.7941 | 505.6574
0.66349 | 963449.1763 | 69.1781 | 167.0108 | 495.5206
0.73699 | 1156897.1205 | 80.1979 | 193.6148 | 495.4899
0.69761 | 1056427.9398 | 79.0316 | 190.7992 | 453.6679
0.70032 | 1075549.1425 | 76.7379 | 185.2616 | 476.6614
0.70801 | 1091003.2018 | 80.4188 | 194.1481 | 502.8642
0.68001 | 1019020.4448 | 88.8846 | 214.5863 | 510.065
0.68034 | 1025807.0422 | 71.8313 | 173.4162 | 506.4111
0.68997 | 1037578.7077 | 73.4116 | 177.2313 | 523.1199
0.68132 | 1024958.09 | 53.1497 | 128.3146 | 501.7728
0.68216 | 1005442.4362 | 78.5072 | 189.5332 | 505.9451
0.69532 | 1044046.6149 | 73.3009 | 176.9641 | 472.3784
0.67994 | 1022706.2344 | 71.6565 | 172.9941 | 495.8142
0.69565 | 1048104.3312 | 62.1607 | 150.0691 | 482.1036
0.68183 | 1019321.194 | 82.7544 | 199.7868 | 497.4992
0.66529 | 995913.8912 | 71.3157 | 172.1714 | 483.7811
0.67527 | 1006475.0065 | 76.1476 | 183.8366 | 493.0584
0.66406 | 978957.6716 | 62.8163 | 151.6519 | 492.7107
0.69839 | 1055486.77 | 86.7024 | 209.3181 | 520.4699
0.69966 | 1065221.723 | 66.8706 | 161.4398 | 507.5488
0.68725 | 1020713.4222 | 65.2337 | 157.4881 | 490.0828
0.69544 | 1056707.5104 | 82.5374 | 199.2628 | 492.1605
0.68089 | 1017990.1804 | 82.7606 | 199.8017 | 487.6047
0.6898 | 1028688.8992 | 65.5001 | 158.1312 | 504.53
0.70277 | 1072566.4997 | 72.7086 | 175.5342 | 504.1397
0.68201 | 1023200.2719 | 78.2967 | 189.0251 | 535.7329
0.70238 | 1069837.6247 | 82.7377 | 199.7464 | 485.8894
0.66747 | 990076.7695 | 68.8771 | 166.284 | 489.4559
0.69491 | 1043410.5242 | 73.2992 | 176.9599 | 501.6139
0.69976 | 1065609.1502 | 79.4731 | 191.865 | 465.2789
0.68433 | 1032149.1969 | 78.7141 | 190.0325 | 489.3265
0.68538 | 1025915.0484 | 79.7298 | 192.4847 | 495.6742
0.69115 | 1063683.7175 | 76.8985 | 185.6495 | 464.2002
0.70123 | 1056638.1128 | 62.1676 | 150.0858 | 461.0338
0.71404 | 1098324.7817 | 85.1452 | 205.5586 | 513.4832
0.67833 | 1006538.6723 | 80.6026 | 194.592 | 489.6166
0.66943 | 989883.1797 | 78.7619 | 190.148 | 517.2861
0.6765 | 1000369.6702 | 61.0089 | 147.2886 | 525.2342
0.69993 | 1066863.8978 | 66.9477 | 161.626 | 490.8791
0.68354 | 1023136.258 | 78.1171 | 188.5914 | 476.3736
0.68403 | 1031633.8484 | 55.6531 | 134.3584 | 464.6465
0.71086 | 1090326.7966 | 90.9769 | 219.6377 | 494.032
0.69836 | 1057420.3941 | 70.308 | 169.7385 | 498.4403
0.68549 | 1024925.8543 | 78.0305 | 188.3823 | 488.6295
0.71902 | 1112258.4487 | 85.6286 | 206.7257 | 485.5997
0.70335 | 1079557.0166 | 73.1221 | 176.5323 | 509.6329
0.68461 | 1007145.7274 | 68.3106 | 164.9164 | 482.3579
0.69964 | 1068699.1349 | 76.3233 | 184.2608 | 486.4603
0.70745 | 1092901.7396 | 77.1897 | 186.3523 | 506.5185
0.67569 | 1003860.9778 | 70.779 | 170.8757 | 483.4653
0.70725 | 1061742.5027 | 75.0104 | 181.0912 | 486.8073
0.67297 | 1017210.6749 | 85.2057 | 205.7047 | 479.0649
0.66564 | 978898.9384 | 70.0607 | 169.1415 | 531.903
0.70278 | 1081486.8696 | 67.5901 | 163.177 | 482.7814
0.69018 | 1046002.9818 | 70.3736 | 169.8969 | 499.308
0.68527 | 1030027.1786 | 71.6084 | 172.8779 | 491.888
0.67261 | 1007586.5174 | 63.8318 | 154.1035 | 497.6783
0.69971 | 1056390.9803 | 64.3318 | 155.3108 | 496.5063
0.72657 | 1130589.3945 | 79.4979 | 191.925 | 482.2241
0.70399 | 1080403.4875 | 71.2621 | 172.042 | 476.1555
0.6951 | 1051563.521 | 86.7891 | 209.5274 | 473.7457
0.6774 | 1013292.1713 | 67.6058 | 163.2148 | 486.0237
0.68665 | 1048741.8146 | 63.0131 | 152.1272 | 486.9097
0.68066 | 1013791.2788 | 77.7304 | 187.6578 | 501.2478
0.69273 | 1055436.4322 | 82.7612 | 199.8033 | 501.1662
0.69006 | 1030427.1833 | 67.4884 | 162.9314 | 475.3651
0.67391 | 1012605.3386 | 65.9967 | 159.33 | 516.886
0.66923 | 1003590.6673 | 65.8664 | 159.0154 | 504.8884
0.65681 | 974515.1532 | 72.4311 | 174.8641 | 486.5748
0.68558 | 1027956.1389 | 71.4317 | 172.4514 | 490.0246
0.71575 | 1106383.605 | 69.8619 | 168.6616 | 471.8487
0.68241 | 1022434.9662 | 59.7183 | 144.1727 | 515.3286
0.71747 | 1107760.0371 | 75.3733 | 181.9672 | 471.3721
0.68364 | 1015443.4816 | 85.2129 | 205.7222 | 495.1821
0.68623 | 1035643.9437 | 77.103 | 186.143 | 479.9741
0.66963 | 1006151.7004 | 73.3505 | 177.0839 | 499.8736
0.69452 | 1054108.7342 | 62.6179 | 151.1731 | 510.8188
0.68494 | 1034476.7266 | 68.2598 | 164.7938 | 506.1794
0.67834 | 1025267.6994 | 68.3102 | 164.9155 | 489.6787
0.72186 | 1126331.0964 | 81.4301 | 196.5896 | 517.9191
0.68362 | 1022828.2926 | 81.2723 | 196.2087 | 496.0931
0.72228 | 1122468.4578 | 86.024 | 207.6803 | 502.2061
0.69499 | 1069875.9695 | 91.3094 | 220.4405 | 487.6937
0.68674 | 1025285.4614 | 88.5545 | 213.7896 | 518.1925
0.7173 | 1116296.7231 | 89.28 | 215.541 | 470.4822
0.68069 | 1015721.5288 | 61.5642 | 148.6291 | 485.069
0.68233 | 1018554.9479 | 63.6073 | 153.5615 | 469.468
0.67228 | 1006096.2089 | 69.4907 | 167.7653 | 455.3191
0.71708 | 1095534.973 | 74.5819 | 180.0567 | 504.2907
0.65945 | 973855.286 | 60.928 | 147.0932 | 462.8099
0.71214 | 1081285.4738 | 82.4773 | 199.1177 | 465.2227
0.69005 | 1030203.1358 | 74.597 | 180.093 | 472.978
0.69432 | 1051349.6162 | 73.9201 | 178.4589 | 490.989
0.68785 | 1046381.476 | 66.8154 | 161.3067 | 483.3285
0.66588 | 976933.2519 | 66.3075 | 160.0805 | 482.7609
0.70461 | 1081680.5067 | 65.4137 | 157.9227 | 495.1856
0.66842 | 983817.0232 | 79.9502 | 193.0169 | 503.4814
0.68028 | 1008931.9379 | 74.1055 | 178.9065 | 476.5907
0.69271 | 1036635.2308 | 78.0992 | 188.5481 | 479.0459
0.68992 | 1039656.1184 | 68.1061 | 164.4227 | 496.4691
0.6764 | 1023716.3537 | 70.2773 | 169.6645 | 524.8425
0.70087 | 1048880.9558 | 71.2959 | 172.1234 | 504.7297
0.69173 | 1048194.1027 | 79.0819 | 190.9206 | 505.7408
0.68737 | 1032893.3949 | 78.422 | 189.3274 | 495.0442
0.6689 | 989891.2544 | 64.9989 | 156.9211 | 474.4154
0.6934 | 1042019.0912 | 84.712 | 204.5128 | 523.2364
0.71363 | 1104163.6721 | 65.9295 | 159.168 | 506.0328
0.68929 | 1029108.6615 | 82.6447 | 199.5218 | 480.4053
0.69077 | 1041518.2596 | 62.2063 | 150.1794 | 491.1679
0.70078 | 1060351.1073 | 68.3854 | 165.0971 | 496.3148
0.68474 | 1025687.8572 | 81.3669 | 196.437 | 488.0159
0.7005 | 1058979.7158 | 66.3991 | 160.3017 | 478.4949
0.68897 | 1047256.0484 | 72.4152 | 174.8258 | 476.5996
0.68753 | 1032527.3288 | 71.5465 | 172.7286 | 509.4464
0.68166 | 1033037.8933 | 68.4927 | 165.3561 | 489.4224
0.68221 | 1023965.8613 | 76.683 | 185.1292 | 496.0821
0.68132 | 1017835.5774 | 81.149 | 195.9109 | 487.3013
0.68855 | 1032674.7464 | 81.4681 | 196.6814 | 504.6395
0.71076 | 1096940.1019 | 71.6638 | 173.0117 | 457.7555
0.68572 | 1027548.6882 | 73.1523 | 176.6052 | 503.1364
0.67741 | 1002942.1454 | 66.9143 | 161.5454 | 501.5074
0.71697 | 1102307.8415 | 69.4864 | 167.755 | 479.5006
0.71063 | 1096572.6087 | 65.752 | 158.7395 | 490.9543
0.69753 | 1069554.5137 | 86.4351 | 208.6728 | 512.2579
0.67896 | 1004874.0784 | 81.9258 | 197.7863 | 455.1117
0.69961 | 1070391.572 | 79.8479 | 192.7698 | 505.4986
0.68175 | 1001662.1565 | 71.725 | 173.1595 | 475.1677
0.69198 | 1024995.241 | 72.311 | 174.5742 | 493.1836
0.69436 | 1062540.7699 | 74.7027 | 180.3483 | 486.0573
0.64959 | 939633.0491 | 70.6142 | 170.4776 | 482.8884
0.70887 | 1083529.6662 | 72.8206 | 175.8045 | 484.6424
0.6907 | 1036321.0015 | 63.9322 | 154.3461 | 479.8863
0.70401 | 1072076.1193 | 64.8882 | 156.6541 | 493.0377
0.68471 | 1026157.4059 | 85.9763 | 207.5652 | 453.8721
0.70409 | 1069881.8923 | 74.1795 | 179.0851 | 488.5097
0.69838 | 1052335.087 | 86.444 | 208.6943 | 493.7256
0.69787 | 1067048.4459 | 67.1575 | 162.1325 | 482.5092
0.69793 | 1043931.1547 | 73.0189 | 176.2832 | 490.0396
0.67443 | 1006108.0724 | 65.5221 | 158.1844 | 486.476
0.71039 | 1088892.1184 | 69.2759 | 167.2468 | 466.2349
0.69897 | 1070188.561 | 58.8063 | 141.971 | 493.8585
0.69516 | 1053228.4109 | 64.5585 | 155.8579 | 511.5395
0.69657 | 1043185.9605 | 82.8991 | 200.1362 | 470.9721
0.66859 | 993914.5005 | 69.0283 | 166.6491 | 466.3447
0.69972 | 1067312.6203 | 76.2157 | 184.001 | 525.4486
0.67612 | 1006939.4471 | 70.9508 | 171.2904 | 513.6247
0.66046 | 994342.0416 | 71.7243 | 173.1577 | 508.6853
0.69315 | 1043264.9238 | 73.4753 | 177.3851 | 492.2644
0.71007 | 1069719.1661 | 66.1689 | 159.7459 | 503.7684
0.6735 | 996703.609 | 89.1835 | 215.308 | 490.764
0.67467 | 1012816.5262 | 79.9613 | 193.0437 | 501.8316
0.70768 | 1076392.2516 | 79.3789 | 191.6376 | 508.3699
0.66599 | 980877.9946 | 70.1656 | 169.3946 | 489.2981
0.70201 | 1060973.8365 | 64.3992 | 155.4733 | 478.4453
0.6875 | 1049157.1322 | 86.0244 | 207.6813 | 484.9709
0.67974 | 1017076.5863 | 69.4787 | 167.7364 | 494.9153
0.66413 | 989586.9823 | 59.3906 | 143.3815 | 478.8028
0.67741 | 1035720.2198 | 70.9954 | 171.398 | 487.633
0.69012 | 1051847.0649 | 81.183 | 195.9931 | 493.5594
0.69271 | 1056504.4443 | 76.2072 | 183.9805 | 520.815
0.69193 | 1050243.3889 | 79.2144 | 191.2406 | 507.323
0.65462 | 962219.2233 | 78.2935 | 189.0172 | 498.724
0.68527 | 1021841.8644 | 65.4943 | 158.1172 | 489.1715
0.68994 | 1031219.4014 | 84.254 | 203.4073 | 515.9785
0.6751 | 1011889.2357 | 71.4074 | 172.3928 | 482.8317
0.68759 | 1050905.3585 | 76.3673 | 184.367 | 478.6565
0.66839 | 988753.2461 | 47.2743 | 114.1302 | 489.6547
0.68352 | 1019263.7438 | 91.2578 | 220.3158 | 482.2798
0.6966 | 1047131.5109 | 67.9378 | 164.0163 | 512.0255
0.69676 | 1056932.2708 | 60.4623 | 145.969 | 505.2634
0.71166 | 1101629.3961 | 73.7461 | 178.0389 | 516.3536
0.67564 | 1005158.3135 | 63.3925 | 153.043 | 475.1916
0.72922 | 1138407.362 | 66.5207 | 160.5951 | 493.4882
0.71125 | 1088616.7194 | 76.4771 | 184.6321 | 475.9271
0.67642 | 984867.0406 | 69.3647 | 167.4611 | 474.8961
0.69888 | 1071462.4649 | 86.4221 | 208.6415 | 476.245
0.70184 | 1063154.4474 | 70.3392 | 169.8139 | 479.2186
0.66092 | 972499.4049 | 76.7921 | 185.3926 | 476.5801
0.67546 | 1009372.4205 | 71.1918 | 171.8722 | 489.0829
0.68397 | 1037895.3846 | 77.5262 | 187.1647 | 502.3143
0.6682 | 971747.5959 | 72.6652 | 175.4293 | 456.4761
0.69036 | 1041457.6652 | 73.6443 | 177.793 | 488.7008
0.693 | 1057996.8225 | 87.5842 | 211.4469 | 507.7367
0.68082 | 1036254.2815 | 70.6762 | 170.6274 | 482.3571
0.67082 | 1005893.9768 | 65.8615 | 159.0037 | 455.0142
0.70902 | 1081009.3164 | 77.8935 | 188.0516 | 482.943
0.69617 | 1069419.3785 | 83.4425 | 201.448 | 474.8452
0.70608 | 1072262.1182 | 58.3276 | 140.8153 | 499.4413
0.69774 | 1087617.2019 | 74.2612 | 179.2823 | 490.5154
0.69025 | 1046676.1394 | 83.9386 | 202.6457 | 485.9762
0.70312 | 1083554.4867 | 80.4261 | 194.1657 | 493.4491
0.71782 | 1101632.1575 | 93.9199 | 226.7428 | 508.0042
0.69942 | 1066808.6349 | 74.4603 | 179.7631 | 479.1043
0.69658 | 1072445.0927 | 78.6559 | 189.8922 | 478.8384
0.68467 | 1021148.3009 | 69.2542 | 167.1944 | 536.8841
0.67703 | 1005038.8584 | 85.1623 | 205.6001 | 486.6978
0.69551 | 1063524.7173 | 71.0043 | 171.4196 | 487.0395
0.72271 | 1131821.1247 | 81.731 | 197.3162 | 540.2869
0.70181 | 1063719.9997 | 72.2075 | 174.3244 | 471.0677
0.66436 | 972763.4722 | 73.1608 | 176.6259 | 509.8171
0.67575 | 1008107.3779 | 77.8559 | 187.9607 | 499.4059
0.68511 | 1026388.0215 | 73.1345 | 176.5622 | 493.6915
0.68969 | 1054914.3366 | 84.6676 | 204.4056 | 491.6978
0.67825 | 1007049.068 | 70.736 | 170.7717 | 472.6992
0.70619 | 1076034.3276 | 77.8462 | 187.9374 | 512.3342
0.67294 | 1009460.2413 | 69.6547 | 168.1614 | 482.9935
0.68653 | 1017063.1026 | 70.577 | 170.3879 | 494.1493
0.67571 | 1017531.5888 | 71.7409 | 173.1978 | 517.1058
0.68182 | 1039300.6292 | 64.9516 | 156.8071 | 485.4006
0.70648 | 1089414.6384 | 63.6204 | 153.5931 | 493.4482
0.70112 | 1079438.311 | 78.6566 | 189.8938 | 499.9826
0.67627 | 1023611.6036 | 77.3297 | 186.6905 | 497.0104
0.68392 | 1021983.1169 | 76.3426 | 184.3073 | 497.8461
0.69735 | 1057007.7522 | 55.9333 | 135.0349 | 501.9255
0.69311 | 1061849.648 | 71.1374 | 171.7408 | 485.0844
0.67626 | 993000.5097 | 64.4942 | 155.7027 | 486.9808
0.68904 | 1039149.0764 | 78.7064 | 190.0141 | 513.7704
0.6995 | 1069175.1063 | 83.027 | 200.4449 | 481.8648
0.66722 | 992619.9419 | 67.2376 | 162.3259 | 475.1708
0.67315 | 992506.2563 | 75.3275 | 181.8568 | 493.751
0.69138 | 1037109.6125 | 63.9178 | 154.3112 | 507.7699
0.68549 | 1032820.5312 | 68.0956 | 164.3973 | 494.2473
0.67287 | 1010529.0875 | 63.9936 | 154.4943 | 493.0561
0.68257 | 1014365.4933 | 67.1649 | 162.1504 | 475.1636
0.7074 | 1090118.6633 | 78.7301 | 190.0713 | 498.5391
0.66917 | 989075.1879 | 81.8177 | 197.5255 | 491.4335
0.70541 | 1100768.9855 | 68.5388 | 165.4673 | 466.1339
0.68247 | 1023777.4468 | 73.2312 | 176.7958 | 481.7507
0.70621 | 1082051.396 | 78.2791 | 188.9824 | 472.494
0.67883 | 1017827.0967 | 86.0177 | 207.6652 | 471.604
0.69947 | 1064244.0363 | 76.0236 | 183.5372 | 471.6922
0.6937 | 1066443.2508 | 69.4807 | 167.7413 | 489.5234
0.6788 | 1008194.1743 | 74.2124 | 179.1646 | 488.5765
0.70546 | 1082762.5623 | 71.2682 | 172.0566 | 477.0276
0.69785 | 1056634.9748 | 89.8094 | 216.8191 | 485.5512
0.70869 | 1078144.9743 | 72.4763 | 174.9733 | 516.091
0.68822 | 1033979.9908 | 76.7561 | 185.3057 | 490.5157
0.65416 | 957944.1819 | 54.0379 | 130.459 | 484.2312
0.71044 | 1077359.5439 | 72.2452 | 174.4154 | 514.317
0.68531 | 1025496.8334 | 81.2838 | 196.2365 | 507.3057
0.67063 | 995461.0789 | 74.1685 | 179.0586 | 505.686
0.72098 | 1109287.3949 | 73.2988 | 176.9589 | 479.5884
0.69799 | 1065379.84 | 86.0418 | 207.7232 | 466.4118
0.70417 | 1078227.1063 | 83.1744 | 200.8007 | 487.7636
0.70865 | 1088162.0659 | 71.302 | 172.1383 | 492.9849
0.68054 | 1031147.3362 | 73.9672 | 178.5726 | 496.2862
0.68861 | 1029714.4189 | 72.9994 | 176.2361 | 494.2334
0.6697 | 988941.7925 | 73.7846 | 178.1318 | 471.6176
0.68072 | 1012096.2537 | 69.0386 | 166.674 | 479.4278
0.69072 | 1035128.2078 | 67.7306 | 163.5161 | 487.7496
0.68302 | 1031143.9755 | 83.5675 | 201.7498 | 487.5781
0.69825 | 1056362.2593 | 85.281 | 205.8864 | 519.0487
0.68708 | 1025848.4854 | 81.1024 | 195.7984 | 505.7981
0.71026 | 1095765.5493 | 91.5083 | 220.9206 | 490.6963
0.69161 | 1043776.9986 | 78.7622 | 190.1487 | 487.1816
0.68068 | 1023913.9897 | 78.505 | 189.5278 | 496.8862
0.69516 | 1043109.5545 | 78.3099 | 189.0568 | 495.9383
0.69706 | 1064438.1667 | 74.5454 | 179.9686 | 481.2565
0.68938 | 1034177.5501 | 71.4682 | 172.5394 | 466.5919
0.69903 | 1069756.3182 | 69.1645 | 166.978 | 507.7721
0.69073 | 1050048.0201 | 77.6658 | 187.5018 | 505.1906
0.67843 | 1009243.4332 | 67.2333 | 162.3154 | 486.007
0.67623 | 1007972.379 | 67.3677 | 162.64 | 505.6395
0.69215 | 1025497.1306 | 74.0309 | 178.7265 | 475.994
0.70607 | 1072452.9247 | 72.3618 | 174.6968 | 469.5721
0.69389 | 1051834.4809 | 66.5795 | 160.737 | 500.5442
0.68678 | 1035637.8302 | 80.3208 | 193.9116 | 488.3348
0.69742 | 1066376.6259 | 83.0568 | 200.5169 | 457.2614
0.66589 | 978070.9526 | 64.5138 | 155.7501 | 468.2472
0.70103 | 1062029.3911 | 53.5238 | 129.2179 | 466.7259
0.69003 | 1036011.4308 | 69.7167 | 168.311 | 512.4081
0.71859 | 1098742.2147 | 81.4435 | 196.6221 | 493.2957
0.67213 | 1008401.8783 | 81.4962 | 196.7491 | 494.7794
0.66714 | 986381.3804 | 63.0003 | 152.0963 | 474.7043
0.70498 | 1074431.9634 | 79.5381 | 192.022 | 558.4188
0.71533 | 1089753.2907 | 64.914 | 156.7164 | 510.1425
0.68469 | 1032800.322 | 71.8623 | 173.4909 | 517.1887
0.66315 | 973823.1987 | 93.8651 | 226.6104 | 473.4295
0.70933 | 1088614.5578 | 67.4078 | 162.7369 | 502.2602
0.69509 | 1048998.9543 | 64.306 | 155.2484 | 487.0716
0.68835 | 1025321.6636 | 70.9623 | 171.3182 | 481.5699
0.6861 | 1037420.9211 | 75.5514 | 182.3971 | 501.1402
0.66935 | 992579.8558 | 83.5917 | 201.8083 | 511.7119
0.67655 | 1003377.7171 | 87.968 | 212.3735 | 461.2711
0.69022 | 1054505.1813 | 60.8949 | 147.0132 | 488.5556
0.71034 | 1088368.6991 | 65.2874 | 157.6176 | 461.3319
0.67474 | 1001072.7678 | 79.0256 | 190.7847 | 501.2522
0.6945 | 1049690.0318 | 68.3105 | 164.9162 | 476.1154
0.68508 | 1038998.8308 | 83.9513 | 202.6764 | 490.8412
0.67692 | 1008520.9083 | 76.1158 | 183.7599 | 482.6157
0.67372 | 1003644.6882 | 63.4773 | 153.2478 | 495.5137
0.68069 | 1008162.7302 | 78.8899 | 190.4571 | 489.5923
0.66529 | 981777.6228 | 64.8165 | 156.481 | 478.8833
0.72432 | 1145714.565 | 75.3569 | 181.9276 | 511.5907
0.69078 | 1039148.211 | 73.4359 | 177.2898 | 488.1331
0.69701 | 1062882.4671 | 74.4421 | 179.7191 | 491.2367
0.70392 | 1070370.1045 | 77.6862 | 187.551 | 523.8435
0.69834 | 1059098.1225 | 77.4817 | 187.0574 | 527.3304
0.69457 | 1060725.7767 | 93.46 | 225.6324 | 469.0332
0.68488 | 1031882.9609 | 97.6352 | 235.7121 | 499.9762
0.68153 | 1018953.8798 | 72.9869 | 176.2059 | 471.945
0.699 | 1081415.8395 | 63.8289 | 154.0967 | 485.4258
0.67249 | 1012368.9135 | 78.5648 | 189.6721 | 518.5067
0.67671 | 1003626.8832 | 87.9687 | 212.3751 | 496.1953
0.72696 | 1122429.1797 | 71.6948 | 173.0866 | 498.7235
0.68119 | 1018044.4808 | 71.199 | 171.8895 | 496.8884
0.70273 | 1048700.0386 | 71.0955 | 171.6396 | 513.8984
0.69297 | 1038038.7983 | 59.7059 | 144.1429 | 496.7007
0.69752 | 1068206.8627 | 78.2396 | 188.887 | 517.2266
0.70695 | 1074510.1506 | 70.5765 | 170.3866 | 479.1025
0.68473 | 1036068.4444 | 80.5955 | 194.5748 | 483.9897
0.69126 | 1058310.2913 | 81.5477 | 196.8735 | 469.6713
0.71204 | 1082186.0017 | 80.9543 | 195.4409 | 508.9182
0.71638 | 1091183.4927 | 72.5658 | 175.1892 | 471.4916
0.68779 | 1026164.2484 | 69.2787 | 167.2535 | 533.9385
0.70863 | 1091873.0357 | 80.4127 | 194.1334 | 492.3205
0.67544 | 1014505.4837 | 80.003 | 193.1443 | 494.5282
0.68016 | 1007886.7908 | 72.3392 | 174.6422 | 495.5641
0.68452 | 1031970.6915 | 83.4522 | 201.4714 | 450.6018
0.69594 | 1057505.8811 | 64.7478 | 156.3149 | 497.8519
0.70926 | 1083391.1352 | 82.9732 | 200.3151 | 486.8978
0.68039 | 1031202.4306 | 62.5298 | 150.9602 | 470.9962
0.67204 | 1002436.0035 | 63.5591 | 153.4452 | 485.6334
0.68884 | 1037355.6401 | 73.5157 | 177.4826 | 484.4631
0.71268 | 1090000.5667 | 60.9606 | 147.1719 | 527.5155
0.69942 | 1070308.1413 | 60.9946 | 147.2539 | 496.846
0.67465 | 1009608.6853 | 45.0885 | 108.8532 | 493.4071
0.67322 | 999942.2744 | 70.7607 | 170.8315 | 460.267
0.70376 | 1075164.5567 | 79.7309 | 192.4874 | 480.8622
0.69808 | 1067720.6957 | 58.7461 | 141.8257 | 456.4065
0.69506 | 1052717.0804 | 77.3786 | 186.8085 | 486.2706
0.68218 | 1014502.6634 | 79.2129 | 191.2368 | 482.1874
0.68013 | 1003829.213 | 64.8265 | 156.5051 | 483.8512
0.68573 | 1029432.1294 | 63.9677 | 154.4318 | 504.7938
0.68209 | 1008386.4559 | 80.3056 | 193.8748 | 497.3389
0.7093 | 1097304.3664 | 82.4021 | 198.9362 | 507.2963
0.69548 | 1053864.6851 | 73.9726 | 178.5856 | 489.3422
0.69643 | 1066051.0557 | 81.5337 | 196.8398 | 501.1063
0.68491 | 1030522.0169 | 62.0763 | 149.8654 | 493.9827
0.70484 | 1083397.4425 | 73.2276 | 176.787 | 480.2701
0.70563 | 1073899.4023 | 79.4254 | 191.7498 | 502.5291
0.70953 | 1088797.382 | 69.3537 | 167.4345 | 493.9692
0.67441 | 1004304.411 | 77.7163 | 187.6237 | 526.7659
0.69564 | 1056186.9334 | 74.1194 | 178.9401 | 473.8881
0.68812 | 1036880.2095 | 78.6395 | 189.8527 | 475.0713
0.69018 | 1031516.5205 | 71.2014 | 171.8953 | 481.1855
0.69319 | 1032166.9432 | 69.1413 | 166.9218 | 493.5648
0.68481 | 1017635.4377 | 65.2682 | 157.5714 | 493.3899
0.68022 | 993969.4758 | 64.1811 | 154.9469 | 522.1844
0.67121 | 979429.7995 | 76.1537 | 183.8512 | 510.3021
0.70585 | 1083965.2889 | 65.4361 | 157.9768 | 484.3777
0.68194 | 1018201.587 | 61.6009 | 148.7177 | 486.048
0.67959 | 1007599.5359 | 65.1217 | 157.2177 | 484.1186
0.69621 | 1053630.5858 | 66.4712 | 160.4756 | 485.6746
0.70281 | 1054336.9398 | 79.9062 | 192.9106 | 479.6833
0.69041 | 1040901.7521 | 71.8254 | 173.4019 | 487.3209
0.67839 | 1000131.968 | 88.8454 | 214.4918 | 491.3532
0.69533 | 1066080.1864 | 65.3309 | 157.7226 | 512.7517
0.68413 | 1031506.5488 | 77.0305 | 185.968 | 474.3164
0.68217 | 1025736.2621 | 71.6342 | 172.9401 | 514.8021
0.66453 | 978749.2356 | 70.1677 | 169.3998 | 461.3818
0.70219 | 1078008.0408 | 83.3918 | 201.3256 | 465.3447
0.69802 | 1060058.1552 | 88.6137 | 213.9324 | 524.4615
0.68688 | 1025396.7624 | 81.0898 | 195.7681 | 494.1389
0.69556 | 1059572.483 | 66.9077 | 161.5295 | 503.287
0.6947 | 1062232.6477 | 67.1593 | 162.1369 | 469.5685
0.70649 | 1072811.3525 | 77.5791 | 187.2926 | 492.0202
0.66777 | 972089.0526 | 65.7923 | 158.8367 | 480.3554
0.68495 | 1045995.7247 | 79.698 | 192.4081 | 465.9663
0.68638 | 1044770.3055 | 76.0561 | 183.6157 | 499.6207
0.68071 | 1017733.9179 | 72.9868 | 176.2057 | 495.4493
0.69756 | 1039506.4344 | 79.3757 | 191.6299 | 498.1251
0.69356 | 1052154.1245 | 77.504 | 187.1112 | 498.4574
0.67967 | 1038506.3399 | 58.686 | 141.6805 | 487.9025
0.68408 | 1032138.4012 | 60.1382 | 145.1864 | 505.1265
0.67219 | 1001685.0812 | 74.4592 | 179.7605 | 492.4887
0.69118 | 1043139.0653 | 73.6761 | 177.8698 | 490.5549
0.69341 | 1053817.4777 | 68.6866 | 165.8242 | 466.623
0.69037 | 1047909.0907 | 72.3131 | 174.5793 | 470.5298
0.70357 | 1076629.1756 | 79.8611 | 192.8016 | 475.7253
0.66901 | 986879.2922 | 78.5721 | 189.6897 | 495.7012
0.66404 | 978199.6168 | 73.6059 | 177.7005 | 517.4588
0.69021 | 1042607.8733 | 71.9645 | 173.7376 | 469.0198
0.69972 | 1046200.806 | 67.5751 | 163.1408 | 501.6152
0.67155 | 1015422.7677 | 73.8147 | 178.2044 | 508.348
0.68157 | 1025014.7651 | 69.8328 | 168.5914 | 495.1803
0.68905 | 1049709.7501 | 76.1192 | 183.7681 | 503.2496
0.68955 | 1042387.8192 | 77.232 | 186.4544 | 491.3413
0.71349 | 1095140.9458 | 61.1785 | 147.6979 | 482.3764
0.68211 | 1029841.9623 | 45.4888 | 109.8196 | 535.8993
0.68117 | 1004240.5968 | 64.754 | 156.3299 | 465.3063
0.70901 | 1094905.1443 | 85.4822 | 206.3723 | 497.8164
0.69772 | 1069129.2088 | 78.2857 | 188.9985 | 457.3749
0.71318 | 1098693.934 | 67.6643 | 163.356 | 497.7069
0.68888 | 1042755.4235 | 83.7904 | 202.288 | 496.4836
0.70995 | 1071869.5026 | 71.9266 | 173.6462 | 494.3636
0.68589 | 1035689.8433 | 73.7139 | 177.9612 | 499.6653
0.6768 | 1004791.8032 | 83.6888 | 202.0426 | 495.0775
0.67766 | 1014749.2376 | 59.6853 | 144.093 | 470.5175
0.67167 | 994943.185 | 65.0616 | 157.0725 | 512.6995
0.71453 | 1102468.1378 | 77.0938 | 186.121 | 479.6556
0.68706 | 1038557.7003 | 66.3922 | 160.2848 | 480.3342
0.71239 | 1094525.3056 | 78.494 | 189.5013 | 486.8553
0.67801 | 1010140.7156 | 65.4375 | 157.9801 | 492.5669
0.68867 | 1055998.8198 | 63.2632 | 152.731 | 480.0113
0.70133 | 1081709.2356 | 67.7439 | 163.5482 | 496.0179
0.67625 | 1004653.6838 | 70.7765 | 170.8696 | 479.3901
0.69031 | 1041389.8487 | 66.2603 | 159.9665 | 488.4981
0.68887 | 1032481.5302 | 63.8647 | 154.183 | 476.3777
0.69676 | 1054502.2241 | 82.2097 | 198.4719 | 486.6999
0.69515 | 1050904.7603 | 70.1973 | 169.4712 | 468.2067
0.72148 | 1122029.159 | 72.2133 | 174.3384 | 499.9976
0.70246 | 1054123.3124 | 59.8117 | 144.3982 | 498.8605
0.69573 | 1051670.0095 | 70.1903 | 169.4543 | 476.17
0.67364 | 995343.0816 | 66.7795 | 161.22 | 500.887
0.67667 | 1020676.3264 | 75.3684 | 181.9555 | 460.2062
0.69871 | 1062415.8838 | 66.7852 | 161.2338 | 503.2891
0.6654 | 982124.3284 | 81.7033 | 197.2491 | 494.7644
0.68568 | 1026313.343 | 84.3746 | 203.6983 | 503.2229
0.68033 | 1016470.4618 | 61.6424 | 148.8178 | 454.8773
0.71685 | 1105485.6914 | 63.6246 | 153.6033 | 520.674
0.70615 | 1082859.4078 | 65.3419 | 157.7493 | 483.5467
0.70131 | 1074164.829 | 76.5313 | 184.7628 | 514.019
0.69263 | 1035598.3882 | 63.7103 | 153.8103 | 493.9533
0.69491 | 1059315.0563 | 82.8051 | 199.9092 | 492.2749
0.67637 | 998723.672 | 77.0608 | 186.0412 | 469.7511
0.69968 | 1084611.6455 | 82.5689 | 199.339 | 485.1634
0.68963 | 1040906.2054 | 73.6837 | 177.8882 | 487.983
0.69954 | 1067086.7732 | 94.7565 | 228.7625 | 499.3419
0.66551 | 982634.5621 | 58.851 | 142.0789 | 474.394
0.67495 | 1003802.0103 | 70.8523 | 171.0525 | 483.0811
0.70618 | 1075048.2848 | 70.6885 | 170.6571 | 508.6257
0.70491 | 1080760.3295 | 87.9565 | 212.3457 | 508.689
0.67242 | 998773.9108 | 74.2169 | 179.1755 | 499.0281
0.69629 | 1049126.1736 | 80.256 | 193.755 | 501.8478
0.68825 | 1024859.0806 | 65.4031 | 157.897 | 494.1974
0.70319 | 1083160.3832 | 91.3644 | 220.5732 | 512.6716
0.69448 | 1049329.8508 | 73.7613 | 178.0756 | 494.5512
0.69007 | 1048123.7653 | 74.1466 | 179.0057 | 466.352
0.68952 | 1032677.215 | 87.3889 | 210.9756 | 493.5717
0.70025 | 1064858.3487 | 70.6118 | 170.472 | 498.5867
0.68564 | 1029721.6728 | 71.549 | 172.7345 | 499.9283
0.69136 | 1044814.422 | 75.5116 | 182.301 | 517.869
0.69244 | 1039612.029 | 63.9739 | 154.4467 | 481.491
0.67578 | 1001128.6152 | 68.7897 | 166.0731 | 488.4993
0.69373 | 1030610.0505 | 65.251 | 157.53 | 511.2048
0.69586 | 1045070.7139 | 61.9618 | 149.5891 | 479.2662
0.69666 | 1081547.3463 | 84.3294 | 203.5892 | 453.873
0.68286 | 1034737.9949 | 75.7141 | 182.7899 | 475.8787
0.6752 | 1013268.5867 | 69.6837 | 168.2313 | 502.6237
0.68803 | 1051445.4828 | 72.8048 | 175.7663 | 511.4029
0.67791 | 1011228.3713 | 71.0652 | 171.5666 | 490.2342
0.68646 | 1030815.3098 | 75.0313 | 181.1416 | 475.6346
0.68914 | 1038394.4915 | 84.9392 | 205.0615 | 504.3604
0.67961 | 1031451.0782 | 85.5538 | 206.5451 | 461.2084
0.70351 | 1086397.7297 | 73.5688 | 177.6109 | 498.8976
0.69133 | 1056573.1238 | 70.9661 | 171.3273 | 487.6052
0.69326 | 1053126.6679 | 72.371 | 174.7192 | 507.7778
0.69685 | 1041483.1124 | 74.6772 | 180.2866 | 473.8521
0.66843 | 978973.1614 | 69.7733 | 168.4476 | 470.3799
0.69153 | 1034732.0398 | 76.445 | 184.5547 | 470.0898
0.70927 | 1087631.0885 | 74.8592 | 180.7262 | 470.2118
0.70141 | 1058273.5534 | 70.0589 | 169.1372 | 471.9595
0.69441 | 1037800.9473 | 69.3961 | 167.5371 | 494.5264
0.68835 | 1035627.9043 | 69.8604 | 168.658 | 509.1765
0.72682 | 1129389.0052 | 64.1087 | 154.7721 | 493.4565
0.68619 | 1021409.3145 | 65.3786 | 157.8379 | 478.2593
0.68143 | 1023661.9674 | 78.3982 | 189.27 | 454.2493
0.70414 | 1076522.4891 | 72.8348 | 175.8388 | 511.8644
0.68074 | 1009695.9721 | 65.1468 | 157.2782 | 483.3283
0.6894 | 1050510.3063 | 76.1392 | 183.8162 | 494.7154
0.70052 | 1071978.3364 | 79.8616 | 192.8031 | 472.2357
0.70545 | 1078357.776 | 69.0866 | 166.7899 | 501.3281
0.69025 | 1018805.1869 | 78.6187 | 189.8024 | 499.526
0.69603 | 1060308.337 | 85.8725 | 207.3146 | 505.2034
0.68806 | 1031140.5399 | 74.8806 | 180.7778 | 487.2695
0.70355 | 1081366.2907 | 78.5252 | 189.5765 | 484.1365
0.69273 | 1053077.4705 | 60.5921 | 146.2824 | 477.1887
0.6857 | 1025989.0763 | 85.8385 | 207.2324 | 516.2002
0.67656 | 1002015.8677 | 70.5581 | 170.3424 | 516.5464
0.66971 | 1002741.1102 | 84.4024 | 203.7654 | 473.2455
0.67573 | 1003788.4698 | 74.2244 | 179.1935 | 499.7747
0.65467 | 955281.6289 | 49.1541 | 118.6686 | 510.5157
0.68645 | 1031875.9158 | 57.7356 | 139.3861 | 493.2584
0.702 | 1084049.7612 | 78.8934 | 190.4654 | 474.9983
0.69518 | 1062941.8204 | 76.3987 | 184.4428 | 502.4489
0.67209 | 1002021.3982 | 57.2632 | 138.2455 | 496.8524
0.70045 | 1063936.743 | 79.2707 | 191.3763 | 489.8878
0.72852 | 1138015.7603 | 84.8528 | 204.8529 | 513.0696
0.67928 | 1012034.1612 | 72.7311 | 175.5884 | 503.5062
0.69888 | 1060409.7528 | 72.2848 | 174.5108 | 503.6985
0.67313 | 998533.7331 | 49.9707 | 120.6398 | 488.6288
0.66582 | 975953.2919 | 78.075 | 188.4898 | 512.4866
0.68793 | 1031663.8801 | 69.6358 | 168.1156 | 505.2642
0.66089 | 970651.995 | 79.78 | 192.606 | 529.0161
0.66178 | 967482.8197 | 67.8839 | 163.8862 | 494.2845
0.66841 | 988520.9735 | 73.8953 | 178.3989 | 481.1846
0.68153 | 1008452.3108 | 66.8756 | 161.4519 | 483.4617
0.68788 | 1025525.5998 | 63.526 | 153.3652 | 476.1409
0.69041 | 1027549.4117 | 77.6727 | 187.5184 | 502.529
0.68799 | 1046430.5317 | 74.2508 | 179.2573 | 489.8737
0.68557 | 1018269.5131 | 72.5078 | 175.0493 | 495.9385
0.71533 | 1110265.9957 | 90.5641 | 218.6411 | 510.7611
0.67063 | 984130.5351 | 85.626 | 206.7194 | 475.8869
0.66108 | 969462.2646 | 79.6589 | 192.3136 | 483.3392
0.72222 | 1106930.9865 | 67.3182 | 162.5204 | 486.8003
0.67872 | 1014060.4379 | 74.6538 | 180.2301 | 484.8717
0.67262 | 986846.0909 | 79.6964 | 192.4042 | 482.5988
0.68807 | 1036211.0088 | 69.928 | 168.8211 | 503.3187
0.69282 | 1044368.6789 | 62.1917 | 150.1441 | 505.5496
0.69434 | 1053549.1286 | 80.8209 | 195.1188 | 472.7969
0.68483 | 1030435.4415 | 73.4534 | 177.3323 | 490.6027
0.6764 | 996048.8419 | 73.579 | 177.6355 | 473.5362
0.71257 | 1108443.1743 | 92.2666 | 222.7513 | 473.0078
0.70414 | 1072874.9926 | 76.1322 | 183.7994 | 532.0663
0.70015 | 1060323.5947 | 72.1481 | 174.181 | 504.7431
0.71207 | 1091101.6072 | 71.1514 | 171.7746 | 520.9159
0.68075 | 1017849.6723 | 71.2321 | 171.9695 | 520.3267
0.68411 | 1021587.9595 | 63.6304 | 153.6174 | 489.6829
0.67865 | 1007963.3135 | 79.1112 | 190.9914 | 506.9604
0.71891 | 1117648.0078 | 84.4477 | 203.8747 | 490.8836
0.68066 | 1012439.6441 | 67.2261 | 162.2981 | 495.3502
0.67553 | 997131.2489 | 68.5403 | 165.471 | 490.3076
0.67563 | 1021480.0562 | 82.1253 | 198.2681 | 470.4705
0.68987 | 1038471.2843 | 55.5483 | 134.1053 | 520.7938
0.68666 | 1048204.1277 | 74.5204 | 179.9081 | 482.3833
0.67821 | 1031734.4828 | 72.4742 | 174.9683 | 506.7143
0.68768 | 1028959.1349 | 76.4437 | 184.5513 | 496.514
0.69777 | 1061108.2048 | 66.7929 | 161.2522 | 455.3675
0.72025 | 1112965.2802 | 78.9462 | 190.593 | 480.7261
0.69542 | 1056914.8006 | 70.5715 | 170.3748 | 486.7192
0.663 | 974694.4778 | 73.4566 | 177.3399 | 480.7219
0.7198 | 1107342.3865 | 59.0873 | 142.6494 | 509.7082
0.68583 | 1031429.7996 | 71.6471 | 172.9715 | 490.5552
0.66483 | 987338.3628 | 72.499 | 175.0281 | 443.9777
0.67747 | 1007757.3925 | 65.3351 | 157.7329 | 484.7171
0.67702 | 1000382.433 | 62.9621 | 152.0039 | 479.2024
0.67713 | 1006303.4761 | 65.274 | 157.5853 | 481.0925
0.70856 | 1093833.2479 | 71.6823 | 173.0564 | 495.0583
0.67161 | 991579.459 | 60.9134 | 147.058 | 462.6306
0.67277 | 983074.5466 | 83.8287 | 202.3804 | 470.6601
0.70709 | 1082668.892 | 72.9456 | 176.1063 | 485.1431
0.69683 | 1051665.9905 | 80.3877 | 194.0731 | 487.2063
0.68936 | 1034325.9158 | 87.5484 | 211.3604 | 501.063
0.69399 | 1053684.7638 | 62.6407 | 151.2281 | 479.4238
0.69943 | 1063674.6495 | 77.6948 | 187.5719 | 497.5094
0.70225 | 1060031.3753 | 81.9939 | 197.9508 | 515.8795
0.67446 | 1018198.8916 | 75.4313 | 182.1072 | 502.6889
0.67413 | 1004559.2519 | 70.9925 | 171.391 | 488.511
0.69958 | 1062274.2528 | 72.3396 | 174.6433 | 508.5126
0.68143 | 998408.3337 | 56.2746 | 135.8589 | 474.821
0.69277 | 1055012.6319 | 74.3435 | 179.4811 | 476.9418
0.70462 | 1082049.2581 | 93.9792 | 226.8859 | 496.1553
0.68856 | 1023575.055 | 85.2812 | 205.8869 | 506.9649
0.69026 | 1035726.3782 | 57.6309 | 139.1332 | 475.4354
0.6821 | 1027557.8032 | 76.9643 | 185.8083 | 502.2499
0.68399 | 1026649.2973 | 69.696 | 168.2609 | 508.7982
0.68901 | 1027312.3038 | 82.5339 | 199.2543 | 492.5134
0.69241 | 1039379.7502 | 59.8311 | 144.445 | 480.0737
0.67361 | 1006541.4664 | 84.2313 | 203.3524 | 476.0117
0.67559 | 1021069.5759 | 80.5034 | 194.3525 | 491.4691
0.69592 | 1050528.1103 | 64.3227 | 155.2887 | 477.8647
0.69924 | 1048656.4536 | 59.7756 | 144.3111 | 485.2083
0.70957 | 1096692.2507 | 82.325 | 198.7501 | 529.3229
0.71262 | 1087441.8536 | 65.0229 | 156.9791 | 512.1858
0.69233 | 1042843.6963 | 73.5328 | 177.524 | 510.0142
0.67732 | 1009951.3795 | 82.5394 | 199.2678 | 500.3271
0.69823 | 1071190.2568 | 76.6561 | 185.0642 | 485.7896
0.68355 | 1011643.4661 | 70.5074 | 170.2199 | 507.3593
0.68823 | 1029738.9015 | 71.2807 | 172.0868 | 530.0919
0.69488 | 1071489.7594 | 61.4605 | 148.3788 | 458.4424
0.68862 | 1035441.1421 | 87.7369 | 211.8156 | 473.8487
0.68612 | 1039585.4529 | 70.3557 | 169.8537 | 493.4257
0.70113 | 1063787.3513 | 66.6409 | 160.8853 | 470.6564
0.68484 | 1013838.5452 | 66.9074 | 161.5288 | 460.7042
0.68047 | 1002292.952 | 85.9455 | 207.4907 | 490.673
0.69771 | 1044868.9917 | 61.786 | 149.1646 | 453.4756
0.70095 | 1081379.7812 | 83.8007 | 202.3128 | 476.838
0.67216 | 1007257.5204 | 76.5597 | 184.8314 | 503.5931
0.71333 | 1089476.5519 | 81.3522 | 196.4015 | 472.9437
0.68303 | 1030310.1665 | 89.4725 | 216.0056 | 466.4512
0.70266 | 1067170.351 | 82.499 | 199.1702 | 500.8559
0.70099 | 1085071.4519 | 87.2567 | 210.6564 | 491.4849
0.69632 | 1057163.0932 | 64.6916 | 156.1793 | 501.736
0.68549 | 1040220.0441 | 80.8282 | 195.1365 | 481.4524
0.69985 | 1077270.7051 | 75.144 | 181.4138 | 477.9235
0.70449 | 1089148.9946 | 75.4715 | 182.2044 | 510.1503
0.71443 | 1101827.4585 | 85.3706 | 206.1028 | 493.1445
0.66552 | 986670.8864 | 86.5061 | 208.8442 | 479.308
0.68681 | 1023068.7303 | 76.1028 | 183.7283 | 500.1994
0.69627 | 1069874.7326 | 75.0115 | 181.0938 | 486.987
0.68071 | 1026169.643 | 71.818 | 173.3839 | 489.1669
0.6895 | 1037037.4526 | 71.653 | 172.9857 | 483.9582
0.6784 | 1014494.7593 | 69.4383 | 167.6389 | 500.1811
0.69611 | 1042680.1159 | 71.359 | 172.2758 | 487.4464
0.67534 | 1008572.2065 | 61.6148 | 148.7513 | 501.2197
0.69493 | 1064829.3195 | 63.2181 | 152.622 | 494.5105
0.68561 | 1032501.3923 | 78.5931 | 189.7405 | 508.5716
0.67567 | 1011515.7419 | 71.3206 | 172.1832 | 508.6073
0.68618 | 1029631.818 | 84.5854 | 204.2072 | 492.7671
0.66457 | 971446.923 | 73.0394 | 176.3328 | 487.3059
0.67964 | 1014348.4869 | 57.3238 | 138.3919 | 489.1084
0.68001 | 1027640.871 | 64.394 | 155.4609 | 488.4398
0.69242 | 1052839.5992 | 66.7848 | 161.2328 | 513.864
0.68784 | 1029464.6055 | 87.3299 | 210.833 | 504.507
0.67045 | 998969.7235 | 65.4434 | 157.9944 | 449.0696
0.69273 | 1058501.8421 | 74.5937 | 180.0851 | 475.5144
0.70154 | 1069755.0699 | 85.9577 | 207.5201 | 485.9311
0.69365 | 1057400.1737 | 90.4191 | 218.291 | 502.2692
0.69433 | 1059214.6834 | 68.9463 | 166.451 | 477.1442
0.68253 | 1018803.3921 | 59.4956 | 143.6352 | 503.8993
0.66667 | 976462.1401 | 62.4108 | 150.673 | 486.2964
0.70725 | 1093103.3021 | 61.604 | 148.7253 | 473.9385
0.71608 | 1122408.228 | 68.8452 | 166.2071 | 482.2736
0.69233 | 1033162.7658 | 74.5645 | 180.0147 | 484.9759
0.70781 | 1069763.3248 | 83.6617 | 201.9772 | 493.9788
0.68812 | 1023691.3858 | 72.6237 | 175.3292 | 543.9605
0.68361 | 1022149.3966 | 74.7101 | 180.3662 | 531.6167
0.69498 | 1058415.2609 | 66.891 | 161.4892 | 479.5784
0.67793 | 1021604.5975 | 80.2674 | 193.7828 | 490.164
0.69874 | 1062665.9093 | 75.9902 | 183.4565 | 473.3905
0.69821 | 1063222.8429 | 82.7166 | 199.6956 | 482.6248
0.70191 | 1066813.0713 | 87.1926 | 210.5016 | 518.9395
0.69479 | 1055852.5932 | 70.5642 | 170.3572 | 516.279
0.68829 | 1041517.828 | 62.4307 | 150.7211 | 500.2257
0.69287 | 1042601.2061 | 66.0918 | 159.5596 | 489.2026
0.68965 | 1031342.7034 | 71.245 | 172.0007 | 485.7954
0.68263 | 1019894.7879 | 85.7121 | 206.9273 | 469.4339
0.68628 | 1033756.9486 | 69.9446 | 168.8612 | 452.3028
0.6667 | 986701.3689 | 77.2377 | 186.4683 | 545.0713
0.69066 | 1036015.75 | 73.2272 | 176.7861 | 475.7183
0.67937 | 1011964.2295 | 65.4242 | 157.9481 | 464.3535
0.71195 | 1094573.5021 | 71.3892 | 172.3489 | 497.6694
0.68937 | 1031584.6383 | 76.4514 | 184.5701 | 482.2824
0.6916 | 1044291.6895 | 77.1432 | 186.2402 | 486.2333
0.68707 | 1045443.6031 | 72.4897 | 175.0057 | 491.4397
0.70248 | 1092671.4855 | 70.2994 | 169.7177 | 504.5817
0.68858 | 1042960.8545 | 79.0475 | 190.8375 | 489.5621
0.69281 | 1031501.6822 | 65.3945 | 157.8762 | 479.1398
0.70617 | 1076580.2654 | 85.9393 | 207.4759 | 529.9658
0.67982 | 1002735.4661 | 64.7852 | 156.4054 | 506.9784
0.70133 | 1071138.2032 | 63.0119 | 152.1241 | 481.9558
0.68194 | 1025028.4568 | 73.3771 | 177.148 | 480.4129
0.70821 | 1071137.3576 | 73.8337 | 178.2503 | 517.3878
0.67829 | 1025230.9529 | 73.7865 | 178.1363 | 486.0246
0.69783 | 1049538.4007 | 71.6521 | 172.9836 | 465.8201
0.69466 | 1056499.3539 | 94.4749 | 228.0826 | 500.6205
0.68906 | 1052137.8663 | 74.5402 | 179.9559 | 481.494
0.67644 | 1008922.6443 | 72.8119 | 175.7836 | 475.9397
0.71961 | 1109123.6991 | 73.4274 | 177.2694 | 494.6053
0.69956 | 1057748.2502 | 60.2661 | 145.4954 | 469.8112
0.6865 | 1034211.5643 | 75.2746 | 181.7289 | 477.7539
0.7112 | 1102751.9703 | 70.078 | 169.1832 | 479.6739
0.68852 | 1040677.9571 | 75.5226 | 182.3277 | 486.9677
0.68154 | 1012693.0278 | 59.2248 | 142.9813 | 486.5252
0.67277 | 996859.3505 | 87.8871 | 212.1782 | 474.0054
0.67137 | 986574.3622 | 70.0072 | 169.0124 | 492.9381
0.67241 | 989018.3471 | 81.4185 | 196.5616 | 507.1932
0.65759 | 952650.1403 | 70.7219 | 170.7378 | 499.7853
0.67614 | 1003123.4874 | 59.1339 | 142.7617 | 508.5061
0.67994 | 1015512.5217 | 79.5523 | 192.0562 | 467.5951
0.71322 | 1071857.3247 | 76.7793 | 185.3616 | 474.9446
0.6554 | 956366.835 | 76.1539 | 183.8517 | 468.0034
0.73121 | 1149437.7321 | 74.8893 | 180.7988 | 471.0508
0.68216 | 1014704.9618 | 74.3241 | 179.4343 | 492.2481
0.69977 | 1065360.4625 | 82.699 | 199.6529 | 491.9158
0.67288 | 999340.1066 | 74.2085 | 179.1553 | 492.1417
0.66654 | 993122.4306 | 69.1858 | 167.0294 | 503.7556
0.69739 | 1044344.775 | 76.507 | 184.7041 | 531.8161
0.69144 | 1049643.9023 | 77.5567 | 187.2384 | 501.8484
0.68714 | 1039851.3776 | 86.8166 | 209.5938 | 477.4422
0.67633 | 1011871.0772 | 69.4712 | 167.7182 | 468.8542
0.69214 | 1051345.9578 | 72.3653 | 174.7052 | 490.1628
0.66322 | 961473.7491 | 77.218 | 186.4208 | 465.6221
0.67791 | 1028020.6869 | 74.0296 | 178.7234 | 478.5756
0.70348 | 1084904.2777 | 80.4846 | 194.307 | 494.7497
0.67814 | 997765.9368 | 70.0954 | 169.2254 | 478.9475
0.68062 | 1015029.8864 | 65.5026 | 158.1373 | 522.3301
0.6915 | 1041619.6191 | 78.6119 | 189.786 | 486.8663
0.67196 | 995130.1546 | 63.1037 | 152.3457 | 498.8576
0.69751 | 1055514.5504 | 66.4653 | 160.4614 | 493.4436
0.69758 | 1049553.5683 | 75.1774 | 181.4944 | 479.7727
0.69604 | 1057286.27 | 79.2743 | 191.3851 | 492.2438
0.68078 | 1029307.8111 | 83.6936 | 202.0542 | 496.3531
0.68771 | 1030736.6821 | 67.7386 | 163.5355 | 481.8578
0.70317 | 1077089.0075 | 63.1962 | 152.5692 | 485.9901
0.69162 | 1044230.4574 | 86.6164 | 209.1105 | 531.6344
0.68464 | 1032498.3615 | 64.2603 | 155.1381 | 501.2273
0.68646 | 1030506.6683 | 65.9349 | 159.181 | 472.7718
0.70278 | 1067148.9503 | 68.6281 | 165.6829 | 489.0112
0.67917 | 1010024.9473 | 72.5987 | 175.2689 | 497.8656
0.70547 | 1091601.2983 | 85.652 | 206.7823 | 505.7939
0.70293 | 1073923.1309 | 79.7318 | 192.4897 | 494.5819
0.68291 | 1023296.2522 | 76.5541 | 184.8178 | 531.1567
0.72409 | 1131830.82 | 68.6545 | 165.7467 | 508.8112
0.66932 | 996859.4013 | 69.1572 | 166.9601 | 493.1459
0.71386 | 1123194.7581 | 76.8093 | 185.4341 | 489.3861
0.71163 | 1084901.5101 | 64.9611 | 156.83 | 502.2711
0.67766 | 1014456.0425 | 67.6575 | 163.3396 | 478.0797
0.69022 | 1036079.6039 | 69.7017 | 168.2747 | 506.6568
0.67151 | 994699.5301 | 63.1182 | 152.3809 | 472.502
0.70841 | 1099895.1441 | 75.7948 | 182.9849 | 485.8017
0.68358 | 1035509.4877 | 77.3913 | 186.8392 | 495.1273
0.6752 | 1026482.1374 | 60.5951 | 146.2895 | 497.1497
0.69404 | 1051274.4777 | 73.944 | 178.5165 | 465.0684
0.6806 | 1014694.3014 | 79.4109 | 191.7149 | 508.4635
0.6868 | 1005346.5615 | 76.3936 | 184.4304 | 504.679
0.69105 | 1042618.1412 | 80.3625 | 194.0123 | 479.3658
0.69416 | 1042972.6539 | 71.5795 | 172.8083 | 491.1351
0.67859 | 1007861.9016 | 79.1098 | 190.9879 | 474.4536
0.68706 | 1033253.9426 | 67.9688 | 164.0912 | 494.8589
0.67883 | 1026560.5751 | 70.2198 | 169.5257 | 501.2342
0.68468 | 1035658.1709 | 66.4372 | 160.3935 | 476.5396
0.68914 | 1033993.8054 | 71.4804 | 172.569 | 497.2361
0.69252 | 1044570.7417 | 71.8589 | 173.4828 | 520.4516
0.68156 | 1020908.2347 | 67.6984 | 163.4385 | 511.3479
0.70156 | 1067914.5553 | 79.4408 | 191.7871 | 494.2109
0.71193 | 1094617.6619 | 71.3945 | 172.3616 | 473.3225
0.68468 | 1044900.3655 | 68.974 | 166.5179 | 473.9307
0.67399 | 991693.7513 | 62.6959 | 151.3614 | 525.9913
0.68632 | 1033850.7532 | 71.7642 | 173.2541 | 526.8585
0.69344 | 1047997.4102 | 70.1759 | 169.4196 | 514.955
0.68944 | 1057625.5036 | 71.2317 | 171.9685 | 491.2762
0.69462 | 1055875.5362 | 68.7007 | 165.8581 | 470.2787
0.72357 | 1134972.0362 | 81.8617 | 197.6317 | 494.0897
0.67629 | 997141.2778 | 83.1149 | 200.657 | 502.2114
0.69566 | 1059566.0791 | 74.3545 | 179.5077 | 490.1219
0.68817 | 1040838.114 | 77.2714 | 186.5498 | 475.223
0.68854 | 1055231.0352 | 83.2489 | 200.9807 | 477.8201
0.69156 | 1071062.0347 | 98.9971 | 239.0001 | 481.2327
0.67034 | 999809.1573 | 74.524 | 179.9168 | 489.1574
0.70183 | 1079780.4413 | 85.1639 | 205.6037 | 477.6578
0.68726 | 1014139.7778 | 81.6832 | 197.2008 | 494.3702
0.70862 | 1091608.4825 | 82.0531 | 198.0938 | 509.2375
0.69306 | 1043685.5073 | 71.7421 | 173.2007 | 488.3362
0.67727 | 1001091.0695 | 60.9838 | 147.228 | 468.8481
0.69797 | 1054709.9386 | 55.7622 | 134.6218 | 497.7246
0.68744 | 1018339.7755 | 72.3151 | 174.584 | 499.8637
0.70163 | 1053438.8728 | 73.2952 | 176.9503 | 490.5853
0.69209 | 1053197.9269 | 72.4975 | 175.0244 | 490.7512
0.68149 | 1012924.4788 | 74.2671 | 179.2966 | 548.9247
0.68285 | 1021375.3937 | 73.0186 | 176.2825 | 504.975
0.67135 | 993468.4278 | 78.8217 | 190.2923 | 499.9444
0.72989 | 1152247.8262 | 65.1312 | 157.2407 | 529.1865
0.69696 | 1055809.7823 | 66.5361 | 160.6324 | 452.1727
0.68905 | 1025063.9816 | 79.2393 | 191.3007 | 484.3201
0.68721 | 1039081.1709 | 82.133 | 198.2867 | 488.8881
0.68635 | 1026166.1893 | 92.8741 | 224.218 | 465.7572
0.67878 | 1003745.6024 | 80.3258 | 193.9237 | 512.2399
0.69892 | 1062094.7935 | 84.1176 | 203.0778 | 521.953
0.68654 | 1029227.1652 | 69.6122 | 168.0588 | 474.0479
0.67199 | 998944.3516 | 70.82 | 170.9747 | 485.3609
0.68796 | 1038348.5956 | 68.2142 | 164.6835 | 491.7108
0.67892 | 1015600.9156 | 74.7449 | 180.4501 | 475.3095
0.71233 | 1108559.2142 | 87.7266 | 211.7906 | 528.3336
0.68616 | 1037890.9457 | 70.2363 | 169.5655 | 490.7419
0.70062 | 1049926.9852 | 73.1558 | 176.6137 | 500.1381
0.7251 | 1120605.6514 | 69.8473 | 168.6263 | 512.4032
0.67897 | 1025642.1587 | 66.3483 | 160.179 | 455.6958
0.68848 | 1027171.0886 | 65.5292 | 158.2014 | 504.9106
0.69692 | 1071693.8562 | 61.2922 | 147.9725 | 489.8136
0.70228 | 1092824.7166 | 89.2724 | 215.5227 | 474.5549
0.6583 | 981838.5079 | 71.0552 | 171.5425 | 510.1963
0.6854 | 1026238.509 | 79.7524 | 192.5394 | 477.3313
0.68226 | 1030226.711 | 77.1458 | 186.2464 | 480.7236
0.70203 | 1060489.0664 | 66.3487 | 160.1799 | 485.7409
0.6959 | 1048329.649 | 66.1654 | 159.7374 | 495.4079
0.67432 | 996305.5862 | 78.6988 | 189.9958 | 491.5598
0.68567 | 1039488.533 | 68.5177 | 165.4164 | 482.1664
0.67153 | 997012.2489 | 82.1845 | 198.4109 | 503.062
0.68723 | 1024419.491 | 63.517 | 153.3437 | 495.4948
0.67214 | 1008259.5826 | 76.6376 | 185.0195 | 506.27
0.71126 | 1112662.014 | 83.3251 | 201.1646 | 449.0188
0.69791 | 1066687.7892 | 69.0767 | 166.7658 | 469.6852
0.68883 | 1037467.9601 | 68.0702 | 164.336 | 482.0354
0.69544 | 1045040.2253 | 87.6824 | 211.6841 | 504.1119
0.68326 | 1021449.3267 | 67.5656 | 163.1178 | 460.2855
0.71127 | 1084417.4595 | 72.6343 | 175.3546 | 486.2387
0.70955 | 1085153.282 | 72.8595 | 175.8985 | 477.7294
0.68498 | 1021717.0985 | 67.4138 | 162.7512 | 473.2826
0.67658 | 1006463.5642 | 77.6338 | 187.4247 | 500.6145
0.69375 | 1050549.3676 | 79.0288 | 190.7924 | 470.9427
0.67428 | 989320.4016 | 68.1295 | 164.4792 | 496.9588
0.6644 | 986279.5155 | 79.0702 | 190.8923 | 485.571
0.68674 | 1044108.1417 | 66.7785 | 161.2176 | 487.3493
0.69346 | 1052410.7564 | 51.1226 | 123.4208 | 496.1808
0.69462 | 1061956.4203 | 76.3903 | 184.4226 | 473.3847
0.69984 | 1065238.8937 | 93.1504 | 224.8849 | 512.8671
0.67743 | 1019451.5929 | 57.7996 | 139.5406 | 485.0594
0.69917 | 1071434.7266 | 67.308 | 162.4959 | 509.6457
0.70978 | 1069617.663 | 73.5666 | 177.6055 | 502.2031
0.6979 | 1069017.5242 | 81.5886 | 196.9722 | 476.312
0.69142 | 1039397.0437 | 62.0214 | 149.7329 | 489.8794
0.68891 | 1038449.1126 | 71.8127 | 173.3711 | 506.468
0.68559 | 1014788.0861 | 77.2475 | 186.492 | 525.4635
0.69083 | 1044236.1704 | 70.1886 | 169.4503 | 488.3705
0.71307 | 1085545.2339 | 66.865 | 161.4264 | 465.5377
0.69483 | 1053192.0216 | 66.5746 | 160.7253 | 465.6269
0.69114 | 1038494.9964 | 71.5841 | 172.8193 | 503.8189
0.69926 | 1060415.5964 | 74.0303 | 178.7249 | 468.2111
0.68213 | 1033081.4192 | 82.267 | 198.61 | 460.6596
0.71964 | 1122701.2159 | 79.7034 | 192.4211 | 469.1456
0.68982 | 1045572.0122 | 77.4365 | 186.9483 | 477.6981
0.67545 | 997822.8894 | 87.6228 | 211.54 | 513.1897
0.68014 | 1004723.5714 | 70.3763 | 169.9034 | 519.4422
0.68864 | 1035261.973 | 71.6197 | 172.9053 | 493.3761
0.69188 | 1036202.1623 | 84.4246 | 203.8191 | 492.6628
0.67928 | 1029145.686 | 68.4741 | 165.311 | 497.7455
0.70291 | 1075394.6237 | 79.8435 | 192.7593 | 487.5209
0.67871 | 1010079.9056 | 77.6683 | 187.5079 | 506.1744
0.67218 | 992763.0199 | 60.9345 | 147.109 | 493.0034
0.69417 | 1041872.5891 | 71.5034 | 172.6246 | 486.0451
0.68797 | 1035919.4515 | 66.136 | 159.6665 | 506.1046
0.68268 | 1016734.8996 | 65.414 | 157.9234 | 471.0053
0.69802 | 1065673.7608 | 72.7329 | 175.5927 | 484.7761
0.66879 | 994107.7493 | 67.1798 | 162.1864 | 482.2795
0.66292 | 978992.4645 | 78.6607 | 189.9037 | 485.5604
0.68065 | 1024021.6732 | 83.2799 | 201.0555 | 505.4112
0.66591 | 992547.9353 | 79.3916 | 191.6683 | 485.8686
0.70278 | 1081105.2819 | 69.5259 | 167.8503 | 507.188
0.68659 | 1030599.1111 | 67.8403 | 163.781 | 476.7693
0.71945 | 1100092.9045 | 69.1074 | 166.8401 | 499.2944
0.699 | 1049393.1458 | 63.9698 | 154.4368 | 462.9602
0.69255 | 1054269.8364 | 87.3322 | 210.8386 | 487.5273
0.69767 | 1061073.9439 | 72.4556 | 174.9232 | 517.8324
0.704 | 1079236.4118 | 67.3329 | 162.556 | 496.2675
0.68664 | 1019148.7976 | 80.6238 | 194.6431 | 476.9776
0.70394 | 1063669.1535 | 71.9855 | 173.7885 | 494.445
0.68221 | 1023077.7346 | 57.5986 | 139.0553 | 467.6089
0.6908 | 1041744.9977 | 64.2577 | 155.1319 | 452.1736
0.6843 | 1029961.3292 | 75.2061 | 181.5635 | 497.9105
0.69851 | 1056171.2255 | 80.5386 | 194.4374 | 487.4504
0.6911 | 1041913.5671 | 66.2173 | 159.8627 | 491.3627
0.66259 | 969764.1043 | 81.0172 | 195.5927 | 489.2524
0.68441 | 1031421.569 | 81.8605 | 197.6288 | 516.5594
0.67267 | 990498.7804 | 79.9849 | 193.1006 | 470.3555
0.67169 | 994142.5059 | 66.8925 | 161.4927 | 469.2066
0.65471 | 957245.6871 | 58.2762 | 140.6913 | 495.0675
0.70769 | 1083705.8482 | 83.1809 | 200.8164 | 477.3045
0.67111 | 994812.7224 | 68.8314 | 166.1736 | 479.8305
0.68856 | 1044255.6015 | 75.7773 | 182.9425 | 479.3818
0.69518 | 1051929.4609 | 78.9705 | 190.6517 | 487.5855
0.69815 | 1078467.0391 | 65.823 | 158.9109 | 514.3153
0.70655 | 1084448.2614 | 80.1013 | 193.3817 | 499.2994
0.68311 | 1023642.2025 | 84.4711 | 203.9313 | 469.6836
0.71036 | 1082489.1223 | 79.5275 | 191.9964 | 473.5638
0.68465 | 1020260.4828 | 70.9934 | 171.3933 | 485.2678
0.71817 | 1116467.6388 | 84.4451 | 203.8685 | 510.6671
0.70024 | 1072966.061 | 74.8016 | 180.5871 | 502.1167
0.69327 | 1046769.3034 | 78.7991 | 190.2379 | 483.7972
0.69221 | 1035615.6373 | 79.6891 | 192.3864 | 491.8415
0.70211 | 1068935.9487 | 79.4543 | 191.8197 | 488.1646
0.70387 | 1083039.4695 | 69.5419 | 167.889 | 503.7555
0.69815 | 1047419.1803 | 63.9273 | 154.3341 | 464.6393
0.69731 | 1060337.7996 | 68.7247 | 165.9162 | 507.3346
0.69487 | 1057090.3548 | 74.2641 | 179.2894 | 507.3152
0.67515 | 1026456.0972 | 74.2186 | 179.1796 | 492.7309
0.67836 | 996842.3666 | 73.4252 | 177.2642 | 462.9814
0.69555 | 1052324.5584 | 72.0774 | 174.0102 | 487.5174
0.68435 | 1018075.7727 | 79.2392 | 191.3004 | 487.4915
0.70074 | 1055864.6649 | 83.4073 | 201.3632 | 518.6791
0.69981 | 1060150.8242 | 68.4668 | 165.2936 | 485.2978
0.69749 | 1048291.7663 | 78.4367 | 189.363 | 489.3133
0.67487 | 1007698.324 | 76.2856 | 184.1696 | 487.0112
0.71021 | 1092346.3848 | 75.0841 | 181.2691 | 468.374
0.70809 | 1100170.383 | 68.2427 | 164.7526 | 519.514
0.66772 | 985229.4876 | 78.5933 | 189.7409 | 503.8436
0.70366 | 1068855.6336 | 87.143 | 210.3818 | 474.9786
0.69642 | 1056052.095 | 79.139 | 191.0585 | 502.9507
0.66827 | 995797.1449 | 83.9991 | 202.7917 | 462.725
0.67585 | 1007327.967 | 88.406 | 213.4308 | 488.9658
0.70826 | 1079449.7017 | 70.7695 | 170.8526 | 496.6815
0.69227 | 1035161.598 | 74.7153 | 180.3787 | 479.8081
0.68327 | 1031781.2531 | 73.7168 | 177.9682 | 499.3037
0.68077 | 1022688.9789 | 90.5309 | 218.5609 | 491.7034
0.71244 | 1111340.6186 | 64.3575 | 155.3726 | 496.5304
0.71243 | 1096230.7622 | 73.3053 | 176.9747 | 481.5905
0.68017 | 1019808.5637 | 81.4442 | 196.6237 | 497.8755
0.71099 | 1083098.2711 | 68.849 | 166.2161 | 475.1962
0.68433 | 1031624.9491 | 80.2969 | 193.8538 | 501.0607
0.70426 | 1080837.504 | 71.263 | 172.0442 | 493.9889
0.70944 | 1096394.6313 | 91.6668 | 221.3032 | 472.5873
0.70197 | 1071929.8055 | 84.5279 | 204.0684 | 479.4251
0.68772 | 1046361.8513 | 68.7652 | 166.0139 | 506.123
0.68805 | 1030382.4295 | 78.1547 | 188.6821 | 491.2994
0.6788 | 1033229.6462 | 70.6789 | 170.634 | 525.7131
0.67965 | 1014609.3467 | 74.5918 | 180.0806 | 463.9156
0.70325 | 1068296.6287 | 77.6093 | 187.3653 | 493.6511
0.68277 | 1014608.3724 | 79.1527 | 191.0915 | 453.5864
0.68569 | 1043903.8965 | 56.1791 | 135.6283 | 482.7166
0.69067 | 1032352.6881 | 65.6507 | 158.4948 | 472.0925
0.65528 | 971038.0171 | 75.7082 | 182.7757 | 505.6667
0.67925 | 1026142.6633 | 75.484 | 182.2345 | 494.5702
0.69145 | 1051577.1794 | 74.2431 | 179.2388 | 494.4362
0.69953 | 1051760.9975 | 59.9278 | 144.6785 | 496.8039
0.65823 | 950210.9096 | 67.0319 | 161.8294 | 443.1537
0.68481 | 1030693.5299 | 81.756 | 197.3764 | 537.7014
0.68282 | 1022209.2219 | 57.4983 | 138.8131 | 519.7786
0.69047 | 1043461.2382 | 71.9966 | 173.8151 | 465.7662
0.68308 | 1023672.4062 | 85.9732 | 207.5577 | 497.9685
0.69442 | 1039359.8106 | 59.657 | 144.0248 | 447.5099
0.66074 | 973486.709 | 76.8901 | 185.6291 | 507.8158
0.68858 | 1031470.3042 | 67.7015 | 163.4459 | 496.5553
0.70145 | 1078311.4577 | 83.5034 | 201.5949 | 488.1581
0.69024 | 1047530.6572 | 79.203 | 191.2129 | 492.0412
0.70445 | 1092860.9185 | 85.8753 | 207.3213 | 495.0795
0.67749 | 1009067.8342 | 65.4182 | 157.9334 | 492.0843
0.68119 | 1025957.5184 | 73.5249 | 177.5048 | 515.0291
0.70386 | 1092653.7184 | 81.0159 | 195.5896 | 503.232
0.68499 | 1021797.8118 | 67.4185 | 162.7626 | 484.4147
0.70241 | 1082025.0711 | 69.6213 | 168.0806 | 521.6994
0.68546 | 1034190.5465 | 57.9488 | 139.9009 | 503.9935
0.69219 | 1036851.1416 | 85.9342 | 207.4634 | 506.715
0.68323 | 1022193.3715 | 76.4362 | 184.5333 | 480.2334
0.68341 | 1047164.9525 | 76.561 | 184.8346 | 497.355
0.67516 | 1012659.6192 | 78.2758 | 188.9744 | 504.6139
0.65175 | 950266.7701 | 71.1763 | 171.8348 | 474.3889
0.68998 | 1034430.5898 | 73.1879 | 176.6913 | 490.4911
0.70191 | 1051006.9463 | 71.3346 | 172.2169 | 506.2339
0.6823 | 1009904.7982 | 73.9577 | 178.5496 | 502.7678
0.67737 | 1009115.6703 | 65.4332 | 157.9697 | 448.7434
0.68808 | 1035753.2255 | 83.3245 | 201.1632 | 509.225
0.66737 | 995893.2141 | 79.4856 | 191.8952 | 495.2805
0.69268 | 1039053.9729 | 69.654 | 168.1595 | 503.8491
0.67965 | 1006101.5713 | 73.9664 | 178.5707 | 505.3805
0.71037 | 1097516.6925 | 65.833 | 158.9349 | 535.1969
0.67404 | 997572.1696 | 86.3562 | 208.4824 | 532.5939
0.68528 | 1025230.7931 | 82.8153 | 199.9338 | 478.1383
0.67672 | 1018100.0979 | 84.808 | 204.7446 | 482.4055
0.69633 | 1046932.6625 | 69.8139 | 168.5456 | 486.7143
0.70773 | 1095056.3949 | 67.9594 | 164.0685 | 472.7624
0.70474 | 1081850.8807 | 67.4251 | 162.7786 | 486.1383
0.68036 | 1012241.9776 | 61.3831 | 148.1918 | 471.9373
0.68927 | 1048724.0537 | 64.8322 | 156.5187 | 501.59
0.70372 | 1066618.0364 | 60.4121 | 145.8477 | 464.8096
0.68944 | 1051849.1579 | 91.9418 | 221.9671 | 489.1595
0.66788 | 985237.5928 | 60.8618 | 146.9334 | 512.1022
0.69674 | 1068325.9125 | 65.3351 | 157.7329 | 480.4322
0.71274 | 1101072.7162 | 82.2853 | 198.6542 | 470.7532
0.68674 | 1045173.3698 | 76.0455 | 183.59 | 500.8193
0.67311 | 990363.7852 | 76.7865 | 185.3789 | 518.7379
0.689 | 1033037.6567 | 74.9154 | 180.8618 | 493.813
0.69212 | 1052233.6188 | 82.583 | 199.373 | 490.6984
0.69389 | 1052614.0728 | 75.7973 | 182.9908 | 475.6566
0.67235 | 1011115.2369 | 73.414 | 177.237 | 512.7223
0.6854 | 1030577.1957 | 76.819 | 185.4575 | 478.4992
0.70086 | 1084001.6846 | 73.685 | 177.8913 | 480.4644
0.68248 | 1023473.2117 | 69.6346 | 168.1129 | 480.9217
0.68755 | 1038276.1002 | 71.9425 | 173.6847 | 505.4129
0.67658 | 1015900.3949 | 65.9498 | 159.2168 | 506.483
0.66695 | 986058.9961 | 78.7496 | 190.1184 | 496.025
0.66282 | 983274.5205 | 67.0466 | 161.8648 | 523.9144
0.69528 | 1038191.9708 | 63.6258 | 153.6063 | 503.5736
0.6842 | 1041306.4623 | 68.7842 | 166.0598 | 498.597
0.70043 | 1068027.8049 | 82.8282 | 199.9649 | 497.6096
0.68481 | 1036955.1965 | 70.3112 | 169.7464 | 476.89
0.69764 | 1051748.3613 | 77.0208 | 185.9446 | 478.8492
0.68389 | 1018723.449 | 80.9144 | 195.3445 | 497.5722
0.71053 | 1111569.4114 | 64.5437 | 155.8223 | 502.9066
0.67176 | 1013056.6888 | 78.7038 | 190.0079 | 495.6614
0.67577 | 1011880.8634 | 69.529 | 167.8578 | 483.1947
0.68512 | 1027014.7311 | 65.8403 | 158.9526 | 518.9386
0.68417 | 1021184.0529 | 76.2558 | 184.0978 | 488.9551
0.7058 | 1077544.7941 | 84.5104 | 204.0261 | 502.5938
0.67531 | 993113.0693 | 66.4643 | 160.4591 | 503.1915
0.70525 | 1085345.9968 | 76.8954 | 185.6419 | 505.0067
0.6942 | 1056209.9027 | 68.7644 | 166.0118 | 464.0475
0.67479 | 1007406.6097 | 72.8802 | 175.9484 | 525.4626
0.695 | 1066030.6997 | 76.641 | 185.0278 | 519.7249
0.67938 | 1014754.6667 | 74.6316 | 180.1766 | 481.7236
0.689 | 1042040.1748 | 75.5688 | 182.4391 | 491.2293
0.69873 | 1047115.4584 | 63.8556 | 154.161 | 519.3798
0.67744 | 1016130.2804 | 76.6316 | 185.005 | 460.7577
0.72115 | 1125441.7474 | 79.7312 | 192.4881 | 469.7029
0.6687 | 977954.3236 | 74.7167 | 180.382 | 508.1041
0.69128 | 1042434.4059 | 62.215 | 150.2002 | 498.0225
0.66391 | 986533.5032 | 70.7916 | 170.9061 | 491.7382
0.70115 | 1072420.6515 | 72.8671 | 175.9168 | 493.4451
0.65999 | 961337.1473 | 60.0948 | 145.0816 | 501.5807
0.68846 | 1042030.5089 | 66.4787 | 160.4937 | 496.789
0.69333 | 1066488.2846 | 73.281 | 176.916 | 500.0196
0.70032 | 1053413.4967 | 78.5009 | 189.5178 | 496.4422
0.69527 | 1060555.7639 | 68.9407 | 166.4376 | 518.7036
0.70012 | 1069426.0349 | 81.3607 | 196.4221 | 487.6567
0.68529 | 1017818.399 | 72.505 | 175.0427 | 474.2822
0.69668 | 1061147.7532 | 79.4911 | 191.9086 | 498.7212
0.68149 | 1014705.1162 | 72.686 | 175.4795 | 485.9586
0.68842 | 1034315.8338 | 73.3451 | 177.0707 | 502.2307
0.67822 | 1016466.5674 | 74.8857 | 180.79 | 525.9442
0.69363 | 1054715.2102 | 85.7153 | 206.9351 | 480.317
0.69424 | 1056049.0854 | 74.2586 | 179.2762 | 506.7278
0.67664 | 1029931.6498 | 74.3059 | 179.3902 | 467.4198
0.69486 | 1051799.2761 | 75.6334 | 182.5951 | 504.8678
0.70179 | 1060540.1409 | 78.8672 | 190.4021 | 479.7613
0.69427 | 1058145.8902 | 62.8806 | 151.8073 | 471.9392
0.67871 | 1015581.7928 | 73.0478 | 176.353 | 493.5977
0.69528 | 1049462.9595 | 73.685 | 177.8914 | 473.0158
0.70426 | 1071446.9095 | 85.7606 | 207.0444 | 456.5544
0.70418 | 1090262.7221 | 75.5826 | 182.4725 | 505.742
0.66981 | 973377.4525 | 69.2325 | 167.142 | 484.7683
0.68496 | 1030947.8129 | 86.347 | 208.46 | 499.3783
0.68849 | 1031553.7908 | 71.379 | 172.3241 | 477.0145
0.67448 | 1006175.085 | 65.5216 | 158.1832 | 496.2368
0.67748 | 1002960.4122 | 65.0235 | 156.9807 | 471.113
0.70809 | 1091028.8286 | 63.5698 | 153.4711 | 486.5505
0.70032 | 1077350.61 | 46.3501 | 111.8991 | 464.2912
0.72147 | 1116180.1965 | 73.7042 | 177.9376 | 489.0731
0.69935 | 1061241.7731 | 79.1945 | 191.1925 | 483.0809
0.685 | 1032236.8541 | 78.6429 | 189.8608 | 466.6572
0.67878 | 1020071.0786 | 57.7198 | 139.3479 | 483.9473
0.68624 | 1051764.0998 | 76.5807 | 184.8822 | 473.3259
0.67149 | 992157.8336 | 62.9589 | 151.9962 | 473.9606
0.6926 | 1044279.6767 | 84.9941 | 205.1939 | 475.6847
0.6874 | 1025256.1997 | 65.5097 | 158.1543 | 496.322
0.70706 | 1062672.9871 | 75.0961 | 181.298 | 468.6688
0.69233 | 1057124.5212 | 70.9003 | 171.1685 | 468.337
0.69214 | 1045420.5264 | 77.1658 | 186.2948 | 498.8277
0.68867 | 1028119.776 | 71.1229 | 171.7058 | 502.3809
0.6869 | 1041714.5739 | 79.1457 | 191.0746 | 472.9872
0.68276 | 1032128.754 | 68.3225 | 164.945 | 527.0157
0.6705 | 998648.2732 | 82.4451 | 199.0402 | 491.7722
0.70076 | 1064694.6835 | 74.1699 | 179.0621 | 464.3269
0.66244 | 975652.2369 | 76.8637 | 185.5655 | 481.8623
0.68963 | 1031845.4373 | 81.2749 | 196.215 | 491.3816
0.68902 | 1045367.5286 | 96.9787 | 234.1272 | 503.5313
0.691 | 1038491.1836 | 66.0101 | 159.3624 | 520.6059
0.69042 | 1053503.5011 | 68.9634 | 166.4925 | 508.3163
0.67901 | 1018558.5444 | 67.7967 | 163.6757 | 500.1156
0.68691 | 1022056.4254 | 67.2467 | 162.3478 | 519.1985
0.68298 | 1033544.2888 | 75.6131 | 182.5462 | 494.4876
0.69065 | 1052014.104 | 93.2182 | 225.0486 | 516.8783
0.67988 | 1023500.5616 | 90.7213 | 219.0207 | 499.1475
0.68656 | 1040806.8033 | 74.0061 | 178.6665 | 515.361
0.71042 | 1092948.0659 | 76.8712 | 185.5835 | 483.9871
0.67904 | 1016418.6468 | 81.3083 | 196.2956 | 519.7886
0.67504 | 1002310.9027 | 83.6998 | 202.0693 | 505.3174
0.7021 | 1055420.2308 | 66.0254 | 159.3993 | 470.1943
0.69211 | 1051254.704 | 74.1488 | 179.011 | 490.8873
0.68307 | 1010577.1398 | 78.803 | 190.2474 | 503.5226
0.6762 | 997630.2544 | 68.507 | 165.3906 | 487.1766
0.69304 | 1043125.4462 | 71.706 | 173.1135 | 507.5362
0.68368 | 1026071.2182 | 90.4437 | 218.3504 | 472.0436
0.67526 | 1021478.8568 | 91.1616 | 220.0835 | 480.8886
0.68582 | 1025056.6184 | 82.7359 | 199.7421 | 499.5033
0.67795 | 1029212.4853 | 80.8619 | 195.2179 | 476.4849
0.67803 | 1019977.6374 | 67.9891 | 164.1402 | 491.58
0.71125 | 1098646.8315 | 71.7256 | 173.161 | 480.0221
0.69542 | 1048004.987 | 73.5682 | 177.6093 | 498.1399
0.68975 | 1041389.3557 | 75.439 | 182.1259 | 507.435
0.69799 | 1062898.1771 | 74.3391 | 179.4704 | 525.0355
0.6902 | 1048219.4993 | 75.8846 | 183.2017 | 467.895
0.68553 | 1033349.9253 | 71.8126 | 173.371 | 468.2732
0.67468 | 1007991.8281 | 77.9713 | 188.2394 | 501.5826
0.69571 | 1035243.0499 | 69.0961 | 166.8128 | 513.4345
0.68514 | 1020232.3267 | 69.1443 | 166.9291 | 510.3465
0.68991 | 1044300.9159 | 64.4985 | 155.7132 | 502.4517
0.66637 | 999217.9796 | 65.8612 | 159.003 | 495.4894
0.7065 | 1069616.2875 | 82.2378 | 198.5397 | 508.5016
0.68454 | 1029608.4222 | 71.6561 | 172.993 | 499.5144
0.70679 | 1083198.8866 | 69.2652 | 167.221 | 477.8248
0.68487 | 1016568.589 | 77.4649 | 187.0169 | 510.6818
0.70562 | 1074022.9237 | 74.3042 | 179.3863 | 491.6211
0.69077 | 1049565.4843 | 56.0687 | 135.3618 | 524.7414
0.69417 | 1051959.3486 | 70.3678 | 169.8829 | 483.8661
0.66164 | 962383.9904 | 63.8865 | 154.2356 | 479.0291
0.697 | 1065196.7694 | 78.0774 | 188.4955 | 462.35
0.67494 | 1007352.8554 | 67.4544 | 162.8493 | 480.6419
0.66408 | 979688.5644 | 58.8013 | 141.9589 | 496.9652
0.71762 | 1112082.565 | 61.7671 | 149.119 | 473.9137
0.69385 | 1064099.3271 | 76.6285 | 184.9977 | 461.7442
0.68079 | 1003212.7537 | 54.378 | 131.2802 | 476.9242
0.66858 | 996432.6455 | 80.9563 | 195.4457 | 474.7545
0.69451 | 1051543.6367 | 64.5153 | 155.7538 | 505.5972
0.67597 | 987967.1825 | 73.0288 | 176.307 | 507.6226
0.68052 | 1027610.6692 | 71.939 | 173.676 | 498.8555
0.68917 | 1045095.784 | 72.245 | 174.4149 | 489.4703
0.6905 | 1034741.0844 | 61.8254 | 149.2597 | 483.1601
0.69911 | 1070889.9002 | 81.5899 | 196.9755 | 494.6403
0.70303 | 1058853.9986 | 83.3718 | 201.2772 | 495.1273
0.68947 | 1034757.6089 | 53.0233 | 128.0095 | 500.9773
0.70649 | 1083332.2955 | 67.3497 | 162.5966 | 503.5717
0.67906 | 1013880.0289 | 82.6486 | 199.5313 | 510.7532
0.70342 | 1079585.2457 | 74.9228 | 180.8797 | 494.577
0.66939 | 1010194.2561 | 64.3048 | 155.2454 | 528.5904
0.67264 | 988653.3722 | 64.5571 | 155.8546 | 486.5883
0.67716 | 1011373.4264 | 76.3049 | 184.2164 | 495.67
0.70521 | 1082182.4492 | 80.0854 | 193.3433 | 476.9332
0.66024 | 978181.2912 | 74.0274 | 178.7179 | 500.4667
0.70401 | 1076436.3446 | 87.7171 | 211.7678 | 494.4617
0.68773 | 1018947.324 | 72.328 | 174.6153 | 494.0666
0.71339 | 1101790.4151 | 69.8019 | 168.5168 | 507.5398
0.69256 | 1029836.2553 | 72.5908 | 175.2496 | 512.4537
0.68133 | 1029580.8885 | 70.1682 | 169.4009 | 497.9035
0.69263 | 1048150.6632 | 66.4665 | 160.4643 | 478.1679
0.68059 | 1021627.4122 | 59.8304 | 144.4434 | 494.4817
0.71608 | 1097945.7971 | 71.1961 | 171.8827 | 484.9392
0.69773 | 1052054.8819 | 77.0334 | 185.975 | 502.7877
0.66783 | 991790.8962 | 54.8018 | 132.3032 | 503.0897
0.70604 | 1079324.6962 | 72.8286 | 175.8239 | 505.441
0.67591 | 999941.7887 | 75.5818 | 182.4706 | 486.044
0.67415 | 1014387.7862 | 75.1835 | 181.5091 | 499.1246
0.67247 | 999081.6099 | 77.5354 | 187.1869 | 516.5248
0.67392 | 996195.5666 | 78.7364 | 190.0866 | 486.7244
0.65524 | 957458.2903 | 74.6534 | 180.2292 | 493.8931
0.6749 | 1000085.8277 | 61.1359 | 147.595 | 486.1732
0.6813 | 1024365.8863 | 73.3992 | 177.2012 | 480.4003
0.70559 | 1070826.5876 | 75.8308 | 183.0718 | 547.2043
0.6895 | 1050348.6485 | 68.8482 | 166.2143 | 455.8823
0.71686 | 1105439.978 | 59.2274 | 142.9875 | 469.2472
0.67506 | 992500.0464 | 64.5757 | 155.8996 | 498.2404
0.67554 | 994624.0199 | 70.143 | 169.3402 | 493.9416
0.70309 | 1072511.3128 | 72.6724 | 175.4466 | 525.5684
0.6658 | 975277.9082 | 64.3382 | 155.3262 | 526.0684
0.71846 | 1098986.9939 | 81.4755 | 196.6992 | 525.2681
0.68059 | 1004893.4868 | 83.2317 | 200.939 | 480.7154
0.67316 | 999758.0847 | 65.2319 | 157.4836 | 498.3767
0.68936 | 1042219.1769 | 85.2243 | 205.7497 | 517.7597
0.70063 | 1069094.7447 | 56.3078 | 135.9389 | 485.5181
0.70275 | 1086503.3096 | 87.1531 | 210.4061 | 492.424
0.70299 | 1078058.9055 | 87.9766 | 212.3943 | 465.6531
0.70588 | 1075766.7625 | 53.8435 | 129.9897 | 499.7618
0.68662 | 1061208.4779 | 80.6603 | 194.7313 | 487.5549
0.69226 | 1038746.2618 | 87.5549 | 211.3761 | 511.5474
0.70025 | 1052120.205 | 83.1696 | 200.7891 | 500.087
0.68684 | 1050792.261 | 89.269 | 215.5144 | 492.3909
0.68256 | 1030917.4435 | 73.7324 | 178.0057 | 514.4048
0.70315 | 1072126.0846 | 62.9066 | 151.8699 | 489.6227
0.68893 | 1031389.3261 | 81.3212 | 196.3269 | 484.4953
0.68247 | 1027977.2735 | 75.2619 | 181.6984 | 505.5279
0.68262 | 1021133.3985 | 74.7442 | 180.4485 | 476.2279
0.67538 | 1002223.1667 | 70.6956 | 170.6743 | 478.1018
0.69058 | 1056327.7808 | 76.4298 | 184.5179 | 478.6864
0.71028 | 1075299.6579 | 64.5083 | 155.7367 | 490.6789
0.70087 | 1072156.8245 | 79.8351 | 192.739 | 478.8901
0.67385 | 994768.3878 | 60.9056 | 147.039 | 494.9511
0.66478 | 978799.5944 | 76.8401 | 185.5083 | 485.7057
0.69375 | 1042211.0294 | 81.6033 | 197.0078 | 503.0559
0.66343 | 973892.9017 | 71.6626 | 173.0087 | 482.4534
0.70156 | 1094557.4062 | 76.1634 | 183.8747 | 496.3742
0.68205 | 1028657.9976 | 70.0313 | 169.0705 | 480.5223
0.66499 | 984610.8083 | 86.3946 | 208.5749 | 486.1988
0.67611 | 1006909.8338 | 77.7222 | 187.638 | 486.3934
0.67967 | 1014889.5504 | 69.3362 | 167.3925 | 490.9583
0.68969 | 1041343.8031 | 66.3164 | 160.1021 | 506.2264
0.69389 | 1055610.1449 | 66.8186 | 161.3144 | 490.0164
0.6805 | 1024972.9227 | 71.7563 | 173.235 | 514.4474
0.67536 | 1023188.7428 | 85.4034 | 206.182 | 466.7995
0.67577 | 1003638.0047 | 77.5093 | 187.1239 | 530.4715
0.67953 | 1028148.6481 | 66.455 | 160.4366 | 467.7913
0.69741 | 1079413.0109 | 84.073 | 202.9702 | 482.77
0.69809 | 1070145.242 | 71.1816 | 171.8476 | 494.3
0.6898 | 1039970.8465 | 81.8951 | 197.7122 | 491.7083
0.70558 | 1089346.8133 | 69.7776 | 168.4579 | 499.2844
0.71214 | 1082568.5882 | 75.9565 | 183.3753 | 493.6309
0.66885 | 1004929.8347 | 84.695 | 204.4718 | 509.9465
0.69628 | 1066555.5457 | 69.2297 | 167.1353 | 510.4483
0.67896 | 1012606.3908 | 76.1949 | 183.9508 | 480.7436
0.67178 | 994231.1109 | 89.1892 | 215.3217 | 471.8367
0.68604 | 1019248.6177 | 63.3062 | 152.8347 | 486.5909
0.70352 | 1086681.2116 | 63.7275 | 153.8517 | 476.022
0.68248 | 1020867.9953 | 76.4201 | 184.4943 | 490.9968
0.65851 | 980219.0637 | 69.1195 | 166.8691 | 484.4413
0.6877 | 1021433.5729 | 59.2014 | 142.9248 | 485.9033
0.69365 | 1042117.5387 | 73.3415 | 177.0621 | 533.4625
0.69112 | 1034830.186 | 78.1427 | 188.653 | 519.9974
0.71319 | 1077056.1357 | 68.2544 | 164.7806 | 484.6477
0.69387 | 1028692.0685 | 68.841 | 166.1969 | 480.1171
0.69228 | 1053384.9608 | 82.6542 | 199.5448 | 468.3757
0.703 | 1063078.3419 | 73.8217 | 178.2213 | 488.734
0.69302 | 1040595.4099 | 81.5626 | 196.9095 | 496.7471
0.69005 | 1031245.8966 | 71.1968 | 171.8842 | 513.2271
0.69218 | 1043108.9817 | 86.4538 | 208.7179 | 484.5589
0.68534 | 1044907.9486 | 77.8941 | 188.053 | 497.182
0.68555 | 1034640.6339 | 64.308 | 155.2532 | 481.5029
0.69789 | 1063003.0333 | 84.3138 | 203.5515 | 496.004
0.69351 | 1050066.7423 | 73.9157 | 178.4483 | 519.5991
0.68213 | 1016627.2326 | 63.5059 | 153.3167 | 492.0115
0.6967 | 1053097.0216 | 64.4075 | 155.4935 | 475.0537
0.69636 | 1058625.115 | 77.6679 | 187.5069 | 473.218
0.69063 | 1057012.4606 | 61.0031 | 147.2745 | 478.1818
0.65483 | 949373.6689 | 65.5246 | 158.1904 | 476.8365
0.68276 | 1032716.0744 | 64.4505 | 155.5974 | 505.7098
0.69628 | 1047119.3174 | 71.646 | 172.9687 | 505.6609
0.66403 | 982910.4119 | 70.5181 | 170.2457 | 492.6827
0.68128 | 1032901.3197 | 74.0127 | 178.6826 | 493.077
0.6792 | 1013610.308 | 67.4477 | 162.8332 | 492.271
0.70494 | 1073770.9078 | 58.5041 | 141.2415 | 479.7316
0.67126 | 1000689.9051 | 74.4882 | 179.8303 | 482.4289
0.6874 | 1051418.9879 | 71.0233 | 171.4655 | 479.2491
0.68917 | 1048377.4724 | 77.7176 | 187.6269 | 490.8036
0.68131 | 1037558.2456 | 62.8307 | 151.6866 | 471.9438
0.67918 | 1014429.9691 | 55.1162 | 133.0623 | 490.2785
0.69229 | 1046566.6061 | 68.3241 | 164.949 | 475.6478
0.70385 | 1062463.1593 | 71.9134 | 173.6143 | 488.6403
0.68693 | 1028242.0059 | 71.3111 | 172.1602 | 518.9984
0.69753 | 1059030.453 | 80.8698 | 195.2369 | 480.0655
0.689 | 1038524.6043 | 69.9903 | 168.9715 | 503.2471
0.6641 | 982849.1785 | 72.2485 | 174.4232 | 480.2123
0.6621 | 968097.5699 | 73.0583 | 176.3784 | 480.2753
0.67994 | 1012319.1152 | 54.9397 | 132.6361 | 515.0115
0.69157 | 1065731.8414 | 75.2289 | 181.6185 | 471.0696
0.70694 | 1061741.5672 | 69.7384 | 168.3635 | 487.7912
0.70544 | 1060173.7475 | 71.5968 | 172.8501 | 476.5733
0.70562 | 1071898.5775 | 62.6734 | 151.3069 | 509.2095
0.69502 | 1033898.3025 | 79.2356 | 191.2917 | 488.1871
0.67839 | 994741.5807 | 82.6573 | 199.5523 | 487.2825
0.69244 | 1043982.0545 | 55.6352 | 134.3153 | 473.7532
0.67734 | 1012722.8613 | 74.7063 | 180.3571 | 510.7617
0.67436 | 998740.1017 | 77.2925 | 186.6006 | 468.3519
0.6951 | 1057787.3971 | 68.7775 | 166.0436 | 483.7203
0.67326 | 1006970.0615 | 79.6664 | 192.3316 | 506.3656
0.68462 | 1029718.9657 | 75.1531 | 181.4357 | 488.1765
0.6891 | 1042647.3817 | 85.2914 | 205.9118 | 501.4128
0.66255 | 970939.9735 | 86.9222 | 209.8488 | 469.9832
0.6961 | 1052026.0694 | 58.0467 | 140.1371 | 468.8195
0.69026 | 1021478.932 | 70.5013 | 170.2051 | 499.1415
0.68916 | 1047322.4898 | 74.1876 | 179.1047 | 480.3962
0.69661 | 1059390.5344 | 64.8015 | 156.4447 | 494.9008
0.66001 | 969875.5524 | 68.2341 | 164.7317 | 484.6739
0.72589 | 1117836.8013 | 92.8028 | 224.0457 | 473.4644
0.69553 | 1047088.4202 | 55.5533 | 134.1174 | 498.7232
0.68913 | 1035973.4296 | 71.6187 | 172.9027 | 503.0724
0.67664 | 997655.2637 | 80.0899 | 193.3542 | 495.8473
0.69204 | 1048173.7848 | 70.33 | 169.7916 | 509.1154
0.68459 | 1028833.9048 | 73.365 | 177.1188 | 483.7317
0.72402 | 1128439.7903 | 92.446 | 223.1843 | 496.7043
0.68943 | 1047832.9546 | 74.1952 | 179.1231 | 506.8076
0.67906 | 1015940.5002 | 63.7499 | 153.9059 | 490.2535
0.68163 | 1032845.1935 | 60.3956 | 145.808 | 485.7279
0.67311 | 1001529.0978 | 85.3594 | 206.0758 | 519.682
0.66211 | 966426.4484 | 65.9683 | 159.2616 | 515.6857
0.70654 | 1082615.7023 | 74.8021 | 180.5882 | 530.1332
0.70199 | 1047154.9607 | 61.5437 | 148.5796 | 498.0722
0.69794 | 1064303.1325 | 72.6478 | 175.3873 | 505.7658
0.69274 | 1040384.6012 | 84.6597 | 204.3865 | 492.2293
0.68881 | 1039164.5727 | 73.6475 | 177.8009 | 485.5634
0.6816 | 1014018.1539 | 86.8061 | 209.5685 | 484.825
0.68066 | 1007509.6363 | 66.8985 | 161.5072 | 494.8429
0.69202 | 1062501.9733 | 74.9525 | 180.9513 | 499.3435
0.70608 | 1075043.0547 | 82.704 | 199.6651 | 482.2454
0.70986 | 1082916.2004 | 72.6778 | 175.4598 | 459.6723
0.68584 | 1037293.9107 | 68.3554 | 165.0245 | 498.8684
0.69092 | 1040156.7838 | 81.7768 | 197.4267 | 476.3725
0.70942 | 1082805.1361 | 72.7152 | 175.5499 | 482.9603
0.70956 | 1097830.8535 | 75.5306 | 182.3469 | 482.4091
0.67869 | 1012407.8669 | 88.4796 | 213.6086 | 493.0868
0.68118 | 1027200.9446 | 66.2336 | 159.902 | 491.1513
0.68445 | 1040338.7932 | 68.696 | 165.8469 | 525.6923
0.70519 | 1075856.197 | 58.5962 | 141.4638 | 484.793
0.69448 | 1058501.1611 | 72.612 | 175.3009 | 496.4968
0.70591 | 1095613.3757 | 68.1701 | 164.5771 | 498.6626
0.67711 | 1025243.6208 | 77.3569 | 186.756 | 505.3458
0.69186 | 1040349.648 | 66.0457 | 159.4484 | 454.2498
0.67778 | 1020163.2554 | 76.897 | 185.6458 | 497.3283
0.7055 | 1060432.3518 | 67.9329 | 164.0046 | 508.8793
0.71497 | 1106937.9681 | 71.8905 | 173.559 | 518.8417
0.67211 | 984676.5215 | 68.0287 | 164.2357 | 506.706
0.69454 | 1049009.4723 | 78.8231 | 190.2958 | 486.6936
0.65737 | 980667.913 | 71.0703 | 171.5789 | 495.9734
0.68135 | 1027843.0848 | 70.0477 | 169.11 | 515.0342
0.68123 | 1007165.9122 | 80.3087 | 193.8824 | 498.8727
0.68288 | 1013644.2809 | 74.168 | 179.0574 | 510.1204
0.68895 | 1049273.1283 | 66.8937 | 161.4957 | 477.3089
0.68554 | 1040306.9437 | 84.001 | 202.7963 | 469.4958
0.67665 | 991579.7769 | 68.0459 | 164.2773 | 468.9629
0.71745 | 1122570.322 | 60.0962 | 145.0852 | 500.4784
0.67245 | 999729.5694 | 72.5763 | 175.2147 | 470.3016
0.68181 | 1028945.544 | 64.3052 | 155.2464 | 466.186
0.67848 | 1009432.0577 | 67.2415 | 162.3353 | 466.9223
0.67748 | 1006566.8917 | 77.5394 | 187.1966 | 486.287
0.68579 | 1032350.3263 | 64.1431 | 154.8551 | 481.263
0.70398 | 1068893.925 | 80.8751 | 195.2497 | 491.3063
0.66163 | 962926.5092 | 82.041 | 198.0646 | 482.2225
0.68299 | 1028926.859 | 75.2747 | 181.7293 | 490.1183
0.68862 | 1043944.2663 | 77.451 | 186.9833 | 486.9496
0.68759 | 1026695.5657 | 76.2852 | 184.1688 | 492.2161
0.6838 | 1020357.0543 | 91.3189 | 220.4632 | 501.4752
0.6647 | 982209.8284 | 68.6151 | 165.6514 | 507.2969
0.69862 | 1066011.0674 | 65.0183 | 156.9681 | 533.2196
0.70205 | 1073875.1238 | 74.6722 | 180.2747 | 485.3017
0.69524 | 1046138.6781 | 73.4563 | 177.3392 | 467.3709
0.69466 | 1076515.8491 | 68.0658 | 164.3255 | 481.1316
0.68968 | 1036043.4214 | 75.0597 | 181.2103 | 505.9565
0.69691 | 1062937.5132 | 56.2823 | 135.8775 | 506.789
0.70184 | 1060068.2808 | 85.1431 | 205.5537 | 514.1782
0.69031 | 1041364.2257 | 68.1795 | 164.5998 | 510.62
0.69812 | 1052726.0827 | 57.9174 | 139.8251 | 467.8989
0.69656 | 1052889.7548 | 62.363 | 150.5577 | 483.8402
0.69751 | 1062554.3252 | 68.8486 | 166.2153 | 498.3789
0.65578 | 968438.0101 | 66.7435 | 161.133 | 500.1946
0.66468 | 973207.2743 | 56.2364 | 135.7666 | 482.556
0.66613 | 983905.5655 | 66.756 | 161.1632 | 497.238
0.67485 | 1007515.5499 | 65.5735 | 158.3084 | 514.5475
0.69253 | 1054265.7821 | 70.6887 | 170.6576 | 479.7887
0.68793 | 1019082.9596 | 70.5735 | 170.3795 | 490.1513
0.68903 | 1046916.1986 | 70.5519 | 170.3274 | 478.4627
0.68457 | 1031855.3102 | 69.99 | 168.9708 | 511.81
0.69959 | 1067625.7813 | 62.9614 | 152.0023 | 474.8862
0.69166 | 1043188.027 | 57.9284 | 139.8514 | 520.5861
0.6967 | 1066756.0178 | 67.2511 | 162.3586 | 513.8189
0.68426 | 1028930.0669 | 78.4766 | 189.4594 | 475.14
0.66041 | 972125.4895 | 90.0832 | 217.48 | 489.9535
0.67243 | 999833.7916 | 83.8169 | 202.3518 | 468.9503
0.69769 | 1050035.776 | 73.4705 | 177.3734 | 511.8114
0.69385 | 1025215.0143 | 72.1308 | 174.1392 | 529.7315
0.68507 | 1031638.1636 | 69.9245 | 168.8126 | 519.1116
0.69305 | 1047734.4824 | 78.8968 | 190.4737 | 486.4049
0.70347 | 1088726.045 | 80.7691 | 194.9938 | 491.5078
0.69722 | 1049697.5968 | 81.7807 | 197.436 | 493.7866
0.68187 | 1021725.6986 | 82.944 | 200.2446 | 446.3048
0.69608 | 1051749.1637 | 71.983 | 173.7823 | 481.8278
0.68606 | 1043961.7014 | 79.4136 | 191.7215 | 510.2622
0.68331 | 1023946.1424 | 79.8172 | 192.6957 | 513.8539
0.67787 | 1010998.804 | 69.254 | 167.194 | 527.3158
0.68374 | 1026756.4084 | 84.6505 | 204.3643 | 469.9506
0.68279 | 1018791.9441 | 61.5603 | 148.6196 | 467.5221
0.70325 | 1083369.0698 | 94.2769 | 227.6046 | 498.2515
0.67202 | 986349.2061 | 76.5991 | 184.9265 | 472.3301
0.6885 | 1032941.9679 | 61.8974 | 149.4335 | 490.0002
0.67684 | 1010488.9657 | 76.2742 | 184.1422 | 491.9992
0.64559 | 929504.4985 | 55.2993 | 133.5042 | 542.8154
0.66418 | 991873.0387 | 77.9375 | 188.1577 | 464.0278
0.68223 | 1021921.6476 | 79.7859 | 192.6203 | 497.7732
0.68839 | 1037044.4919 | 75.2726 | 181.7241 | 510.2034
0.67649 | 999290.1535 | 80.2393 | 193.7147 | 491.5067
0.69279 | 1045256.8784 | 62.2474 | 150.2784 | 485.7718
0.71179 | 1104833.4534 | 70.1517 | 169.3611 | 466.4648
0.6768 | 1002242.875 | 70.5489 | 170.32 | 511.3789
0.67987 | 1001019.7968 | 56.5513 | 136.5268 | 504.7321
0.68977 | 1039172.4678 | 71.7734 | 173.2764 | 500.0724
0.68829 | 1037475.431 | 68.1243 | 164.4666 | 490.7998
0.69191 | 1042040.9188 | 71.7482 | 173.2155 | 465.175
0.68745 | 1047631.212 | 72.6017 | 175.2761 | 467.9895
0.69151 | 1066648.9708 | 83.7879 | 202.2818 | 511.9737
0.69799 | 1063417.6012 | 68.857 | 166.2355 | 473.8562
0.68262 | 1019808.8193 | 82.6975 | 199.6495 | 473.7934
0.66922 | 994399.6189 | 70.7896 | 170.9012 | 504.9515
0.68876 | 1051046.6367 | 70.858 | 171.0664 | 468.8325
0.7057 | 1074985.3088 | 74.3631 | 179.5284 | 497.1454
0.67971 | 1022838.1406 | 71.6899 | 173.0746 | 472.6309
0.68561 | 1028784.5082 | 67.8179 | 163.7268 | 500.157
0.69009 | 1052075.2399 | 74.4247 | 179.677 | 513.5385
0.68883 | 1029753.9488 | 63.6995 | 153.7841 | 487.3639
0.69855 | 1077363.3775 | 71.615 | 172.8939 | 493.7599
0.69908 | 1066496.8098 | 76.2267 | 184.0275 | 470.6847
0.66965 | 995444.6005 | 70.818 | 170.9697 | 497.613
0.70331 | 1064530.079 | 80.6217 | 194.638 | 517.1551
0.66693 | 974437.8314 | 66.0341 | 159.4205 | 526.8667
0.69646 | 1060836.5934 | 88.8767 | 214.5674 | 504.5985
0.6912 | 1049490.4116 | 80.8754 | 195.2506 | 472.4877
0.71116 | 1092410.3195 | 89.63 | 216.386 | 504.4616
0.67846 | 1018585.6886 | 81.5508 | 196.8811 | 505.2632
0.70069 | 1057905.1709 | 64.3334 | 155.3145 | 492.2255
0.69044 | 1029568.4859 | 72.7951 | 175.743 | 475.2217
0.67161 | 996738.9892 | 65.185 | 157.3706 | 472.1148
0.70255 | 1062253.6099 | 80.5358 | 194.4305 | 488.8172
0.70309 | 1058662.8535 | 75.2351 | 181.6335 | 493.6922
0.6851 | 1021155.2669 | 67.3648 | 162.633 | 491.1538
0.68269 | 1027792.9214 | 84.8659 | 204.8844 | 508.6406
0.67569 | 1007698.0296 | 63.5475 | 153.4172 | 489.919
0.67042 | 1005102.354 | 74.9102 | 180.8492 | 482.811
0.68526 | 1032140.65 | 78.6062 | 189.7721 | 496.9166
0.70038 | 1057460.5771 | 60.1792 | 145.2856 | 483.5546
0.67042 | 997907.0005 | 74.3742 | 179.5553 | 495.5275
0.68185 | 1015108.0176 | 79.2977 | 191.4415 | 484.7915
0.66805 | 984487.7567 | 58.7378 | 141.8057 | 494.4366
0.70267 | 1067267.5769 | 74.1469 | 179.0064 | 493.32
0.67117 | 984107.4281 | 66.2684 | 159.9861 | 479.3741
0.69212 | 1074541.5719 | 79.3186 | 191.492 | 475.0864
0.67875 | 1006116.0521 | 63.1623 | 152.4873 | 503.0627
0.70983 | 1074456.8135 | 85.3274 | 205.9986 | 451.7394
0.68636 | 1038788.4821 | 83.7783 | 202.2586 | 491.5611
0.67402 | 983955.697 | 64.1185 | 154.7958 | 514.2329
0.70885 | 1087433.3695 | 71.2334 | 171.9726 | 520.5182
0.64954 | 949287.3543 | 73.0241 | 176.2958 | 491.1317
0.68797 | 1024294.235 | 77.7009 | 187.5864 | 489.9365
0.71989 | 1117317.7749 | 73.9415 | 178.5107 | 476.1787
0.69189 | 1043129.9128 | 64.242 | 155.094 | 488.7037
0.6972 | 1051798.9459 | 62.2412 | 150.2635 | 460.5076
0.69958 | 1078767.1774 | 102.5922 | 247.6796 | 497.4911
0.70303 | 1071091.2695 | 72.5822 | 175.229 | 489.2868
0.68831 | 1043501.7707 | 82.3502 | 198.8109 | 492.2815
0.69896 | 1073131.2308 | 74.95 | 180.9454 | 482.8277
0.69835 | 1060390.9432 | 85.5945 | 206.6433 | 480.1141
0.66981 | 982071.2741 | 60.4916 | 146.0396 | 495.1994
0.67698 | 1017912.0858 | 64.0695 | 154.6775 | 497.0757
0.68576 | 1029249.1772 | 84.6062 | 204.2574 | 476.3761
0.66072 | 968522.5927 | 84.0941 | 203.0212 | 488.0798
0.68784 | 1032365.1212 | 69.692 | 168.2513 | 488.2642
0.68495 | 1037504.1631 | 70.3346 | 169.8028 | 475.3209
0.69957 | 1044603.5522 | 71.1372 | 171.7403 | 507.6118
0.68561 | 1019822.2955 | 70.8641 | 171.081 | 485.382
0.67946 | 1014693.1525 | 65.592 | 158.353 | 494.0684
0.67931 | 1005516.0081 | 81.936 | 197.8109 | 486.1498
0.69823 | 1058345.6082 | 68.5057 | 165.3874 | 504.0823
0.70756 | 1096485.0645 | 79.1716 | 191.1372 | 506.4906
0.70966 | 1101738.1436 | 61.8798 | 149.391 | 483.8339
0.6844 | 1025089.3463 | 57.5279 | 138.8846 | 481.3615
0.68406 | 1029126.5832 | 76.8611 | 185.559 | 481.5592
0.69452 | 1055205.6714 | 70.5489 | 170.3202 | 486.9998
0.68222 | 1012717.2059 | 75.839 | 183.0916 | 473.8366
0.69228 | 1021700.7709 | 60.8898 | 147.0009 | 475.9866
0.71559 | 1111615.942 | 89.118 | 215.1499 | 481.9442
0.69802 | 1052851.742 | 78.7174 | 190.0405 | 525.3046
0.6916 | 1041178.7801 | 81.7765 | 197.4259 | 523.374
0.7084 | 1071195.5259 | 64.433 | 155.5551 | 507.5114
0.71564 | 1113533.1629 | 72.2514 | 174.4303 | 487.4182
0.69294 | 1057725.8916 | 70.879 | 171.117 | 491.5531
0.69481 | 1056496.6845 | 74.2291 | 179.2048 | 505.1288
0.69137 | 1044841.8502 | 68.3019 | 164.8954 | 484.4248
0.6897 | 1035950.7079 | 59.8686 | 144.5357 | 517.8166
0.70494 | 1086289.1382 | 71.5537 | 172.7459 | 511.5157
0.66335 | 990057.8823 | 67.4546 | 162.8497 | 509.5301
0.6996 | 1057969.7519 | 72.0446 | 173.931 | 471.7219
0.69803 | 1059449.9491 | 66.6636 | 160.9401 | 492.6127
0.69329 | 1039626.1765 | 67.7737 | 163.6201 | 521.0262
0.71058 | 1093546.5024 | 76.8951 | 185.6412 | 472.8064
0.74117 | 1171193.1993 | 75.2817 | 181.7461 | 491.9149
0.70084 | 1075968.5596 | 71.288 | 172.1044 | 508.1152
0.67233 | 1006952.5383 | 73.1136 | 176.5118 | 451.8608
0.69781 | 1049603.3388 | 78.4981 | 189.5111 | 504.2154
0.7004 | 1069978.4821 | 70.9361 | 171.2548 | 473.6719
0.7091 | 1103295.4428 | 77.7435 | 187.6893 | 482.7337
0.69297 | 1053252.2893 | 88.6861 | 214.1071 | 507.5414
0.71778 | 1102062.6759 | 74.9529 | 180.9524 | 499.3367
0.71392 | 1111964.1563 | 86.218 | 208.1487 | 502.594
0.69038 | 1047415.1502 | 80.8107 | 195.0942 | 466.7142
0.68755 | 1039306.0711 | 60.25 | 145.4563 | 479.3675
0.69505 | 1051680.6918 | 72.085 | 174.0287 | 498.8461
0.68989 | 1044582.3222 | 77.3559 | 186.7536 | 464.3533
0.67597 | 1032401.6707 | 92.0401 | 222.2044 | 497.1375
0.67149 | 978859.0657 | 69.4477 | 167.6617 | 500.1041
0.69289 | 1044312.3477 | 62.1824 | 150.1215 | 494.7878
0.67296 | 995724.0222 | 73.9312 | 178.4857 | 480.1745
0.68981 | 1018371.7693 | 62.906 | 151.8685 | 516.0215
0.70519 | 1100278.7984 | 81.427 | 196.5822 | 505.716
0.69185 | 1036040.0999 | 73.1039 | 176.4883 | 479.4034
0.68931 | 1029212.6148 | 77.9232 | 188.1233 | 500.1259
0.69782 | 1046565.7708 | 81.4664 | 196.6773 | 470.8726
0.70204 | 1067207.6648 | 75.9565 | 183.3751 | 466.1039
0.673 | 1002989.6243 | 74.4662 | 179.7774 | 517.668
0.68096 | 1018165.243 | 78.0321 | 188.3861 | 485.694
0.68076 | 1009775.884 | 80.5728 | 194.5199 | 509.4619
0.6989 | 1067349.3076 | 82.9565 | 200.2747 | 490.7275
0.6937 | 1050108.3468 | 58.142 | 140.3671 | 476.3889
0.68853 | 1049304.0042 | 74.3969 | 179.6099 | 454.9372
0.67346 | 994117.5082 | 64.8348 | 156.5251 | 476.9787
0.71684 | 1101842.2622 | 83.4938 | 201.5717 | 504.0344
0.70002 | 1074376.3901 | 78.4111 | 189.3011 | 468.0953
0.67985 | 1006077.206 | 70.501 | 170.2045 | 530.9668
0.70459 | 1063211.4324 | 66.2776 | 160.0082 | 479.8741
0.688 | 1035526.6514 | 69.8896 | 168.7284 | 469.2615
0.67744 | 1014318.8671 | 76.4951 | 184.6756 | 504.6905
0.68556 | 1015360.5261 | 65.0517 | 157.0487 | 493.006
0.69586 | 1064947.496 | 63.1403 | 152.4341 | 488.8473
0.70194 | 1055134.8376 | 83.2077 | 200.881 | 465.6715
0.68015 | 1011013.4275 | 77.5762 | 187.2856 | 511.1155
0.67511 | 1003764.0939 | 65.3038 | 157.6573 | 506.3171
0.6898 | 1053585.1507 | 69.0311 | 166.6557 | 491.0588
0.70407 | 1090742.6017 | 90.3952 | 218.2332 | 502.7552
0.69827 | 1063365.305 | 77.8014 | 187.8292 | 495.0463
0.68854 | 1033256.5513 | 74.9813 | 181.0208 | 498.233
0.70153 | 1066374.4798 | 70.5835 | 170.4035 | 486.2685
0.69847 | 1081621.8387 | 69.9878 | 168.9655 | 498.9068
0.69323 | 1060304.9771 | 58.7463 | 141.826 | 482.9251
0.67455 | 1005900.3793 | 77.8241 | 187.884 | 508.1213
0.6911 | 1045585.0349 | 58.109 | 140.2875 | 486.0045
0.68152 | 1020755.198 | 50.4534 | 121.8053 | 479.4087
0.71726 | 1091747.7865 | 79.4361 | 191.7756 | 491.0964
0.67619 | 982497.5666 | 83.356 | 201.2393 | 501.9709
0.68767 | 1032913.2243 | 78.3889 | 189.2476 | 516.0554
0.6907 | 1031837.7285 | 72.9286 | 176.0651 | 502.5963
0.68429 | 1028652.8243 | 73.3839 | 177.1645 | 513.4947
0.65189 | 948382.5762 | 69.3081 | 167.3246 | 468.9644
0.67527 | 990573.9389 | 68.1158 | 164.446 | 492.7188
0.67502 | 1005803.3771 | 72.7392 | 175.608 | 508.3834
0.6852 | 1031652.4569 | 71.729 | 173.169 | 480.7335
0.66563 | 982925.7895 | 73.7846 | 178.1317 | 477.9726
0.70254 | 1066521.4645 | 84.0332 | 202.8742 | 495.7263
0.67719 | 1013240.7824 | 92.944 | 224.3867 | 481.1505
0.70997 | 1096660.6358 | 63.7283 | 153.8537 | 489.0477
0.70378 | 1080054.8894 | 73.112 | 176.508 | 470.4074
0.70189 | 1068164.476 | 66.8421 | 161.3711 | 502.6967
0.69311 | 1063514.6254 | 71.2495 | 172.0115 | 477.5266
0.68066 | 1029539.3247 | 68.3612 | 165.0387 | 468.4913
0.65034 | 943919.7459 | 63.7491 | 153.9041 | 488.5084
0.67099 | 987394.5645 | 76.7981 | 185.407 | 496.5373
0.67465 | 995371.2797 | 72.0246 | 173.8827 | 481.2396
0.6885 | 1042344.627 | 85.3416 | 206.0328 | 480.4666
0.68364 | 1013163.1752 | 56.9212 | 137.4199 | 478.265
0.70022 | 1072716.6675 | 61.0616 | 147.4156 | 495.7235
0.69886 | 1049680.5651 | 71.5558 | 172.751 | 481.4809
0.69107 | 1042514.0275 | 62.2386 | 150.2572 | 506.7748
0.6858 | 1035500.4621 | 68.2413 | 164.749 | 489.3588
0.68541 | 1022242.8121 | 89.8791 | 216.9873 | 507.3443
0.69289 | 1057202.7483 | 67.0157 | 161.7901 | 472.0383
0.6985 | 1058486.4568 | 66.5586 | 160.6867 | 477.3558
0.71116 | 1094238.5447 | 76.881 | 185.6071 | 498.219
0.69046 | 1023361.3496 | 72.3541 | 174.6781 | 483.44
0.70894 | 1074712.2213 | 75.7456 | 182.866 | 493.6557
0.68412 | 1008380.9676 | 68.4433 | 165.2368 | 484.3725
0.68883 | 1044922.2538 | 80.7996 | 195.0676 | 491.6973
0.68836 | 1032668.2145 | 63.924 | 154.3261 | 506.8296
0.68204 | 1030892.8935 | 72.0078 | 173.8422 | 516.4363
0.69757 | 1038397.4524 | 76.0513 | 183.604 | 469.3267
0.6728 | 1001921.7503 | 80.8915 | 195.2894 | 468.3935
0.69579 | 1036453.4853 | 69.1691 | 166.9889 | 499.7999
0.6964 | 1048571.6721 | 75.2344 | 181.632 | 474.4079
0.70734 | 1083405.8776 | 61.0495 | 147.3865 | 468.1107
0.69861 | 1058035.2957 | 66.5197 | 160.5927 | 520.1611
0.67879 | 1015408.1759 | 59.6246 | 143.9466 | 505.4523
0.69847 | 1056203.1617 | 78.9179 | 190.5246 | 545.948
0.69683 | 1055728.3496 | 68.4734 | 165.3093 | 473.5101
0.68453 | 1017496.9672 | 77.5736 | 187.2793 | 504.3336
0.66808 | 991954.7498 | 59.1807 | 142.8748 | 480.6464
0.72882 | 1140910.4864 | 85.0336 | 205.2894 | 502.2016
0.67429 | 1005505.0683 | 79.4281 | 191.7563 | 490.3581
0.6957 | 1054700.7787 | 62.547 | 151.0019 | 483.3361
0.68511 | 1036060.8135 | 55.8039 | 134.7226 | 491.3841
0.71155 | 1089424.05 | 61.0256 | 147.3289 | 528.9561
0.66714 | 984632.7363 | 80.17 | 193.5475 | 491.3154
0.66066 | 995472.3675 | 68.0995 | 164.4067 | 485.4358
0.70005 | 1054594.6378 | 73.5409 | 177.5435 | 479.7477
0.69037 | 1052490.2866 | 72.6292 | 175.3423 | 467.0676
0.66647 | 1003440.9142 | 80.1952 | 193.6084 | 493.1873
0.68889 | 1042864.8791 | 73.901 | 178.4127 | 499.0942
0.70693 | 1059016.6881 | 73.1311 | 176.554 | 507.1729
0.69354 | 1038782.2063 | 90.2619 | 217.9114 | 488.3677
0.68604 | 1013386.9074 | 62.942 | 151.9553 | 482.6114
0.69939 | 1067221.7001 | 81.2784 | 196.2234 | 489.8971
0.6847 | 1027265.4925 | 59.8002 | 144.3705 | 479.6091
0.65993 | 980767.7801 | 72.5511 | 175.1539 | 468.3346
0.69481 | 1068753.1815 | 85.1461 | 205.5609 | 520.0979
0.6659 | 980171.2401 | 75.2009 | 181.5511 | 492.4808
0.68925 | 1038615.1462 | 76.985 | 185.8581 | 494.9143
0.70037 | 1061725.6741 | 68.5146 | 165.4089 | 492.6246
0.67037 | 996247.1125 | 70.7996 | 170.9253 | 522.7601
0.70016 | 1057503.2816 | 70.1329 | 169.3158 | 500.138
0.70088 | 1066768.5781 | 74.3015 | 179.3798 | 480.5975
0.67948 | 1006746.4032 | 70.5865 | 170.4108 | 479.9239
0.67672 | 1009743.6439 | 69.2848 | 167.2683 | 495.6593
0.68489 | 1037948.3827 | 70.3706 | 169.8897 | 472.2327
0.68195 | 1007769.072 | 66.7891 | 161.2432 | 491.8139
0.70571 | 1078351.7365 | 70.9529 | 171.2955 | 507.3161
0.681 | 1031390.0679 | 64.5342 | 155.7994 | 472.5836
0.68408 | 1042394.8539 | 82.7714 | 199.8278 | 500.2213
0.70122 | 1063951.2661 | 75.8132 | 183.0294 | 514.1352
0.69634 | 1048981.2764 | 64.1896 | 154.9673 | 492.0622
0.68137 | 993788.0095 | 74.5148 | 179.8947 | 466.0381
0.69357 | 1061077.4979 | 83.1024 | 200.6269 | 487.0954
0.6911 | 1054080.3369 | 68.933 | 166.4191 | 510.9723
0.65951 | 966403.7382 | 62.4391 | 150.7413 | 491.8713
0.66702 | 1006166.974 | 53.294 | 128.663 | 485.3745
0.68429 | 1026237.3874 | 59.7762 | 144.3125 | 479.1219
0.674 | 1012160.4722 | 78.3721 | 189.2069 | 496.7296
0.69165 | 1040457.4806 | 90.6555 | 218.8617 | 461.0857
0.6669 | 981296.9131 | 58.6484 | 141.5897 | 457.8765
0.67255 | 1000036.753 | 77.6003 | 187.3436 | 497.8403
0.66699 | 976304.4447 | 60.3902 | 145.7948 | 480.1967
0.6993 | 1072776.0322 | 76.6515 | 185.053 | 487.6831
0.70538 | 1080948.4522 | 79.975 | 193.0767 | 515.2328
0.67332 | 983238.9563 | 67.8071 | 163.7009 | 527.4101
0.68079 | 1029404.8442 | 53.4218 | 128.9718 | 488.4891
0.69271 | 1038878.2921 | 74.9355 | 180.9104 | 473.532
0.69579 | 1047408.2115 | 92.1244 | 222.408 | 483.8232
0.68956 | 1037211.8958 | 67.982 | 164.1231 | 472.4412
0.69563 | 1059653.0111 | 66.9061 | 161.5257 | 480.1517
0.66715 | 987316.027 | 63.0595 | 152.2392 | 474.6456
0.66694 | 998216.6456 | 73.0657 | 176.3962 | 503.4531
0.69721 | 1061774.0524 | 76.0931 | 183.705 | 479.6624
0.69991 | 1051484.5589 | 80.0202 | 193.1858 | 471.8782
0.68773 | 1025314.8285 | 77.8061 | 187.8404 | 510.4454
0.68735 | 1019634.1434 | 82.115 | 198.2431 | 490.2074
0.69176 | 1030397.2053 | 79.3389 | 191.5411 | 476.6111
0.68149 | 1016041.2881 | 67.383 | 162.6769 | 492.2597
0.69344 | 1060529.6917 | 74.66 | 180.2452 | 496.8467
0.69394 | 1051928.6097 | 68.511 | 165.4003 | 506.844
0.6911 | 1048605.0394 | 70.4546 | 170.0925 | 481.772
0.67304 | 1006530.789 | 84.3021 | 203.5232 | 494.0671
0.69288 | 1059919.348 | 78.1527 | 188.6774 | 511.6921
0.70479 | 1070943.0704 | 88.6634 | 214.0524 | 491.579
0.6886 | 1047492.6581 | 72.4708 | 174.96 | 487.3678
0.70373 | 1084451.3648 | 75.2275 | 181.6153 | 477.4921
0.69173 | 1041142.7861 | 73.4763 | 177.3876 | 502.5082
0.67679 | 1000655.0207 | 61.0007 | 147.2686 | 497.0657
0.65902 | 962942.2068 | 66.0385 | 159.4311 | 502.8354
0.68336 | 1039018.2349 | 74.2248 | 179.1945 | 478.0373
0.68398 | 1037207.9189 | 60.4422 | 145.9205 | 470.669
0.71887 | 1109758.1808 | 87.0227 | 210.0915 | 473.4888
0.68123 | 1029263.6404 | 64.3802 | 155.4276 | 501.509
0.67721 | 1015426.236 | 65.8574 | 158.9939 | 486.4474
0.67378 | 1005030.4162 | 77.8456 | 187.9359 | 493.6849
0.71613 | 1122037.061 | 64.6421 | 156.0599 | 475.3479
0.70065 | 1070877.2345 | 84.604 | 204.2522 | 480.0337
0.68709 | 1034137.589 | 78.5486 | 189.633 | 468.6331
0.71095 | 1074646.0104 | 60.2483 | 145.4522 | 497.7904
0.70153 | 1078177.6032 | 78.5189 | 189.5615 | 480.8812
0.7128 | 1089845.2762 | 76.3969 | 184.4383 | 479.8245
0.68679 | 1023231.2528 | 77.7538 | 187.7142 | 458.688
0.69714 | 1059717.8109 | 80.9672 | 195.472 | 483.475
0.68873 | 1027851.5798 | 77.8847 | 188.0303 | 470.5967
0.68667 | 1036423.1383 | 75.4168 | 182.0723 | 490.5556
0.70714 | 1086415.5073 | 71.3389 | 172.2272 | 487.6658
0.69948 | 1066030.3362 | 77.8623 | 187.9763 | 483.3594
0.68021 | 1014900.1941 | 85.5966 | 206.6485 | 481.7775
0.67493 | 1010814.077 | 74.8321 | 180.6607 | 501.0922
0.67607 | 1010045.08 | 76.327 | 184.2696 | 497.9147
0.67349 | 1003468.2758 | 70.9828 | 171.3675 | 466.6517
0.71573 | 1098199.3194 | 80.0763 | 193.3213 | 487.8902
0.69295 | 1047525.3348 | 95.2754 | 230.0153 | 465.0895
0.68702 | 1025862.0811 | 63.6266 | 153.6081 | 494.5118
0.66253 | 972130.0246 | 64.4471 | 155.5891 | 499.1486
0.69174 | 1038893.511 | 73.3168 | 177.0025 | 476.9282
0.70427 | 1075703.0456 | 74.564 | 180.0134 | 489.9455
0.6896 | 1039146.8792 | 75.2928 | 181.7728 | 472.9872
0.66735 | 983700.4348 | 58.7524 | 141.8409 | 487.7562
0.68866 | 1045426.1423 | 72.3211 | 174.5986 | 490.5485
0.69062 | 1047715.833 | 75.8016 | 183.0013 | 485.8144
0.71151 | 1082684.2482 | 76.032 | 183.5575 | 523.486
0.68869 | 1035482.5197 | 76.8155 | 185.449 | 498.5847
0.70079 | 1069344.246 | 79.6346 | 192.255 | 471.507
0.68882 | 1031247.9544 | 76.4875 | 184.6571 | 489.2894
0.69642 | 1059203.2143 | 81.012 | 195.5802 | 513.5034
0.70589 | 1075085.8621 | 76.0999 | 183.7214 | 502.5169
0.68239 | 1021757.688 | 73.0947 | 176.4663 | 478.8728
0.699 | 1054828.0119 | 75.4012 | 182.0345 | 475.0501
0.70442 | 1076044.5276 | 63.0229 | 152.1508 | 499.644
0.6605 | 973161.3233 | 66.5897 | 160.7618 | 501.971
0.70658 | 1089910.2542 | 75.301 | 181.7928 | 521.9879
0.70044 | 1067211.6441 | 82.763 | 199.8075 | 504.4456
0.69231 | 1066405.6142 | 78.6965 | 189.9902 | 473.7127
0.70487 | 1081606.793 | 63.3083 | 152.8397 | 472.7775
0.69729 | 1066514.1153 | 76.4242 | 184.5043 | 484.6637
0.68217 | 1033282.6272 | 66.5286 | 160.6141 | 493.0824
0.68171 | 1013785.2583 | 77.6108 | 187.369 | 519.2108
0.64873 | 944428.59 | 77.5439 | 187.2075 | 509.7142
0.6969 | 1060310.2055 | 76.0222 | 183.5337 | 500.4831
0.67935 | 1019971.6118 | 75.0188 | 181.1114 | 486.7578
0.68017 | 1025543.7825 | 70.0118 | 169.0235 | 499.823
0.68006 | 1008301.9937 | 70.6347 | 170.5273 | 501.6972
0.65565 | 973811.8268 | 80.6787 | 194.7757 | 498.9969
0.67854 | 1012777.9695 | 77.8953 | 188.056 | 499.8004
0.68987 | 1037555.4592 | 73.4202 | 177.2521 | 491.1919
0.69865 | 1041565.0935 | 69.2254 | 167.1249 | 493.7798
0.66551 | 990998.0521 | 59.3524 | 143.2895 | 489.9802
0.71099 | 1103079.3495 | 66.1088 | 159.6007 | 467.9387
0.69041 | 1040094.5268 | 66.1679 | 159.7434 | 483.0703
0.68775 | 1018875.2034 | 68.7904 | 166.0747 | 513.2608
0.68968 | 1040927.5184 | 86.5858 | 209.0367 | 493.5915
0.69246 | 1034053.8533 | 72.8989 | 175.9936 | 486.6428
0.68059 | 1027473.7264 | 68.2305 | 164.7231 | 480.0236
0.68472 | 1025601.7797 | 61.7972 | 149.1915 | 477.6829
0.65565 | 948230.3141 | 67.1553 | 162.1273 | 481.5216
0.67836 | 1014223.2203 | 79.6372 | 192.2611 | 462.5362
0.66116 | 973176.7101 | 68.3473 | 165.0049 | 486.3343
0.68 | 1020772.5876 | 78.3418 | 189.1337 | 476.6013
0.70718 | 1075735.7705 | 81.0241 | 195.6095 | 495.7666
0.69091 | 1052484.3217 | 62.8482 | 151.729 | 480.252
0.68707 | 1031685.7656 | 75.0274 | 181.1322 | 503.4015
0.67627 | 988604.7527 | 83.8642 | 202.4661 | 508.3799
0.71531 | 1080051.0016 | 71.9331 | 173.6619 | 467.1544
0.67334 | 998667.1499 | 59.1161 | 142.7189 | 477.7803
0.67668 | 1006797.8469 | 83.8709 | 202.4823 | 499.6002
0.67595 | 1004830.3447 | 82.2873 | 198.6591 | 505.1342
0.68331 | 1026750.1841 | 73.3533 | 177.0904 | 456.1057
0.68037 | 1032832.8268 | 74.1069 | 178.9098 | 487.6584
0.71301 | 1098562.4557 | 67.6723 | 163.3754 | 500.3915
0.68685 | 1033549.5558 | 80.1509 | 193.5013 | 480.1812
0.69684 | 1052966.8612 | 73.7659 | 178.0867 | 486.4774
0.66889 | 987983.5558 | 56.7305 | 136.9595 | 472.16
0.65511 | 959677.9002 | 74.8412 | 180.6827 | 493.5834
0.69225 | 1051937.0784 | 87.1772 | 210.4643 | 466.4348
0.68448 | 1039937.216 | 80.9253 | 195.371 | 482.6168
0.71144 | 1077240.0457 | 68.4337 | 165.2135 | 499.6766
0.6824 | 1029882.4007 | 60.1543 | 145.2253 | 471.932
0.68389 | 1032073.603 | 85.0699 | 205.3768 | 489.4089
0.68656 | 1018412.7213 | 83.6171 | 201.8697 | 481.9095
0.68217 | 1033864.6024 | 73.9854 | 178.6165 | 491.5251
0.69102 | 1031028.2204 | 76.2273 | 184.0289 | 488.7525
0.69286 | 1053512.4906 | 72.4393 | 174.884 | 498.9427
0.67042 | 974688.7485 | 69.2617 | 167.2126 | 507.5757
0.68732 | 1025427.1618 | 63.5717 | 153.4757 | 497.3889
0.68106 | 1024190.7675 | 62.0434 | 149.7861 | 479.5326
0.69123 | 1050142.6817 | 68.6625 | 165.7659 | 488.3032
0.67446 | 1000801.3678 | 79.0375 | 190.8133 | 459.3728
0.70458 | 1091985.9367 | 66.0391 | 159.4324 | 470.7397
0.67671 | 1013135.0317 | 65.7573 | 158.7523 | 504.7974
0.69667 | 1067360.5597 | 74.7925 | 180.565 | 502.2204
0.66213 | 970983.5256 | 64.4101 | 155.4998 | 525.0815
0.69141 | 1034330.5716 | 65.7058 | 158.6278 | 496.5733
0.70293 | 1063013.7372 | 75.5616 | 182.4219 | 484.2758
0.68887 | 1034900.8368 | 69.7584 | 168.4117 | 478.8314
0.68704 | 1025129.4265 | 74.5539 | 179.9891 | 500.8402
0.69478 | 1043771.5003 | 83.1598 | 200.7656 | 477.4967
0.6917 | 1029667.6411 | 63.4298 | 153.133 | 468.3112
0.68841 | 1034335.7611 | 83.1702 | 200.7906 | 485.3211
0.69185 | 1029846.207 | 82.3972 | 198.9244 | 509.0843
0.68149 | 1016931.3435 | 79.4827 | 191.8883 | 487.2547
0.69196 | 1028046.8033 | 74.235 | 179.2192 | 495.5834
0.70845 | 1083697.0599 | 76.4323 | 184.5239 | 478.8585
0.66936 | 994585.4907 | 74.2434 | 179.2394 | 493.5356
0.69052 | 1033456.8016 | 76.4618 | 184.5951 | 511.0756
0.70458 | 1080496.737 | 56.5894 | 136.6189 | 498.3736
0.67945 | 1016622.754 | 85.8373 | 207.2296 | 473.3821
0.66813 | 980566.7205 | 73.3323 | 177.0398 | 466.0866
0.69199 | 1029798.3106 | 77.665 | 187.5 | 490.4181
0.68271 | 1036695.2072 | 70.5098 | 170.2258 | 516.6333
0.70204 | 1079639.3796 | 54.3328 | 131.171 | 484.0035
0.67179 | 989444.2005 | 71.9005 | 173.5832 | 496.147
0.67888 | 1028089.1918 | 73.9282 | 178.4785 | 473.4803
0.68434 | 1034288.4929 | 72.0022 | 173.8287 | 492.5074
0.71324 | 1105900.502 | 71.9974 | 173.8172 | 505.9464
0.68625 | 1019078.0786 | 67.1154 | 162.0309 | 516.5296
0.70162 | 1076213.2561 | 74.8803 | 180.7771 | 527.2099
0.69086 | 1053840.564 | 70.8311 | 171.0014 | 496.0981
0.68526 | 1029658.8423 | 87.6746 | 211.6652 | 495.064
0.69131 | 1047061.8109 | 66.5248 | 160.605 | 531.6052
0.67715 | 1013938.986 | 69.5288 | 167.8573 | 503.4206
0.69218 | 1058442.9199 | 69.1101 | 166.8466 | 505.5914
0.67956 | 1018533.0991 | 73.1685 | 176.6443 | 495.7826
0.68707 | 1038398.2021 | 73.7798 | 178.1202 | 467.8038
0.71511 | 1107085.1194 | 79.0932 | 190.9478 | 499.5162
0.66662 | 985030.416 | 84.7712 | 204.6557 | 474.8271
0.69025 | 1035902.9888 | 67.8279 | 163.7509 | 484.8516
0.68344 | 1022745.9462 | 76.4537 | 184.5755 | 480.4535
0.67406 | 993414.7438 | 56.6052 | 136.6571 | 518.0849
0.69779 | 1054724.9318 | 68.3136 | 164.9237 | 475.6234
0.6651 | 977630.6909 | 70.0266 | 169.0592 | 498.4701
0.7034 | 1054276.6062 | 57.5677 | 138.9808 | 463.0493
0.70109 | 1074098.1418 | 76.5505 | 184.8091 | 485.5595
0.68428 | 1020535.2085 | 76.1952 | 183.9514 | 477.2978
0.69286 | 1066224.2082 | 80.3114 | 193.8889 | 490.9624
0.65005 | 952821.4421 | 71.555 | 172.749 | 486.3751
0.66351 | 991471.754 | 74.6635 | 180.2537 | 521.0059
0.66031 | 964951.7095 | 69.6206 | 168.079 | 488.3047
0.69691 | 1076336.8837 | 59.3192 | 143.2093 | 471.8923
0.70629 | 1071920.3082 | 62.6149 | 151.1658 | 522.8121
0.70431 | 1075771.3228 | 78.0342 | 188.3912 | 450.5053
0.72222 | 1108233.9813 | 95.3195 | 230.1216 | 476.8671
0.7019 | 1059229.7521 | 55.6876 | 134.4417 | 492.8394
0.67505 | 1003833.2696 | 72.5941 | 175.2577 | 485.3518
0.68044 | 1021505.9279 | 78.3476 | 189.1478 | 488.528
0.68986 | 1050127.5865 | 81.0808 | 195.7464 | 506.9931
0.68434 | 1023833.5758 | 73.0345 | 176.321 | 525.1998
0.71216 | 1102350.4443 | 77.3422 | 186.7205 | 498.7143
0.7065 | 1086420.7244 | 78.5629 | 189.6677 | 488.1384
0.69246 | 1034126.1404 | 81.1212 | 195.8439 | 480.8246
0.68561 | 1031813.0455 | 64.1267 | 154.8156 | 486.6911
0.68662 | 1033904.97 | 84.8825 | 204.9245 | 484.8691
0.69724 | 1064949.6607 | 58.6639 | 141.6271 | 526.1179
0.70248 | 1061874.3639 | 80.5154 | 194.3813 | 470.2025
0.6778 | 1019327.2092 | 62.0457 | 149.7916 | 476.9988
0.6819 | 1018820.2574 | 65.6231 | 158.4282 | 493.9208
0.71307 | 1114483.7499 | 52.6491 | 127.1062 | 484.7462
0.7159 | 1112628.2019 | 74.042 | 178.7531 | 504.0028
0.66049 | 969413.9218 | 68.1524 | 164.5344 | 498.6917
0.68463 | 1031148.599 | 66.1524 | 159.7061 | 479.898
0.68045 | 1039383.6445 | 72.7707 | 175.6839 | 482.6925
0.70439 | 1067718.9946 | 72.2137 | 174.3393 | 513.0561
0.68944 | 1056344.1704 | 79.9618 | 193.0448 | 490.2363
0.69663 | 1061829.3467 | 84.3734 | 203.6955 | 455.328
0.67281 | 1005441.8561 | 76.3472 | 184.3183 | 469.8977
0.67179 | 1008809.8599 | 63.9873 | 154.479 | 461.8101
0.68673 | 1028451.8884 | 67.6854 | 163.4069 | 511.4987
0.68864 | 1034334.6962 | 83.1424 | 200.7234 | 498.2747
0.65683 | 954501.6673 | 59.955 | 144.7441 | 489.7641
0.68721 | 1019811.2997 | 70.6981 | 170.6804 | 516.6562
0.6887 | 1034967.3276 | 84.7133 | 204.516 | 495.1437
0.67843 | 1018705.8768 | 86.142 | 207.9652 | 495.3343
0.6666 | 995170.0511 | 71.1228 | 171.7057 | 474.2235
0.69333 | 1040246.6294 | 76.6527 | 185.056 | 462.8297
0.70053 | 1070468.5153 | 72.7985 | 175.751 | 483.6091
0.68682 | 1035479.0066 | 78.6808 | 189.9523 | 463.656
0.6932 | 1062955.3055 | 58.8953 | 142.1859 | 474.1258
0.71233 | 1091574.9239 | 65.296 | 157.6386 | 493.1133
0.70556 | 1078155.5027 | 76.3531 | 184.3328 | 521.0069
0.67334 | 996916.9143 | 66.9146 | 161.5463 | 491.8657
0.68647 | 1038473.9522 | 77.2862 | 186.5854 | 489.4837
0.69337 | 1053115.2837 | 70.5263 | 170.2655 | 486.1415
0.68941 | 1028743.9498 | 82.6003 | 199.4147 | 470.2606
0.69526 | 1057621.3771 | 81.0257 | 195.6134 | 470.7023
0.66858 | 999554.2674 | 72.9843 | 176.1996 | 502.9332
0.65979 | 970692.8178 | 76.7809 | 185.3654 | 466.4084
0.69572 | 1054727.9287 | 68.5171 | 165.415 | 487.0503
0.69248 | 1043277.3173 | 73.5468 | 177.5576 | 513.5427
0.68322 | 1028098.9583 | 66.0929 | 159.5623 | 480.7101
0.70322 | 1085105.1169 | 83.8181 | 202.3548 | 490.4698
0.71653 | 1106019.7601 | 82.2178 | 198.4914 | 482.9236
0.67222 | 1002376.444 | 82.5422 | 199.2746 | 494.8075
0.68921 | 1057981.9216 | 83.3839 | 201.3066 | 475.148
0.70161 | 1050682.0828 | 85.9116 | 207.4091 | 495.8254
0.68841 | 1020161.0321 | 66.9758 | 161.6939 | 502.8261
0.69968 | 1071463.7148 | 86.3232 | 208.4028 | 489.2047
0.69457 | 1058083.5561 | 76.1165 | 183.7616 | 460.1651
0.70365 | 1083924.3444 | 73.3867 | 177.1711 | 500.4917
0.70996 | 1104766.8726 | 77.7524 | 187.711 | 482.036
0.6921 | 1061065.2456 | 71.189 | 171.8654 | 504.2478
0.69612 | 1075211.2428 | 69.8081 | 168.5317 | 504.8567
0.7084 | 1069675.6445 | 68.2448 | 164.7575 | 483.7393
0.6857 | 1041966.5947 | 62.6936 | 151.3557 | 484.6288
0.67559 | 1015916.2331 | 73.4091 | 177.2253 | 474.4824
0.68078 | 1010326.3226 | 82.1499 | 198.3275 | 473.9146
0.67887 | 1014089.7307 | 74.6394 | 180.1955 | 479.7934
0.69922 | 1065929.3358 | 74.4196 | 179.6648 | 497.9295
0.68657 | 1045922.4282 | 72.5758 | 175.2135 | 513.8737
0.69857 | 1057048.3505 | 72.0881 | 174.036 | 502.8354
0.6942 | 1070020.1847 | 67.7001 | 163.4425 | 494.2837
0.69511 | 1066441.3751 | 81.7192 | 197.2876 | 497.2483
0.70286 | 1078386.0661 | 67.3893 | 162.6922 | 467.8497
0.68266 | 1034304.8779 | 79.0713 | 190.895 | 526.7651
0.67449 | 1003141.7307 | 77.6182 | 187.3869 | 539.1863
0.6927 | 1063514.0954 | 76.7142 | 185.2044 | 484.6169
0.70517 | 1056973.9796 | 81.4194 | 196.5638 | 509.384
0.72092 | 1103063.5184 | 69.1529 | 166.9499 | 502.1282
0.67537 | 1025238.7633 | 82.4596 | 199.0752 | 485.0494
0.72113 | 1112076.8629 | 80.4816 | 194.2998 | 512.6561
0.68419 | 1041152.1242 | 81.0542 | 195.6822 | 507.5989
0.69916 | 1069813.3223 | 78.174 | 188.7288 | 531.56
0.68594 | 1055833.9866 | 83.6116 | 201.8563 | 476.8333
0.68674 | 1036188.2804 | 92.3392 | 222.9266 | 501.5534
0.68973 | 1036335.3177 | 71.581 | 172.8118 | 505.5652
0.6798 | 1004769.1554 | 78.7272 | 190.0642 | 464.8137
0.71872 | 1112376.9202 | 80.7726 | 195.0024 | 536.1982
0.68005 | 1010747.2436 | 70.8074 | 170.9441 | 490.2472
0.69056 | 1030742.9765 | 84.1394 | 203.1304 | 480.8129
0.67327 | 1005411.3363 | 71.1425 | 171.7531 | 485.1131
0.689 | 1040070.8124 | 75.4261 | 182.0947 | 482.8171
0.71339 | 1086655.7109 | 62.8444 | 151.7197 | 460.4304
0.69001 | 1045922.5567 | 66.5774 | 160.7321 | 489.7923
0.68291 | 1019865.3392 | 81.1212 | 195.8438 | 479.6441
0.6755 | 1018111.1357 | 71.8037 | 173.3495 | 507.9925
0.69012 | 1044229.3577 | 72.086 | 174.031 | 474.8349
0.69301 | 1047305.2744 | 68.3013 | 164.8939 | 501.2597
0.68264 | 1038736.8973 | 74.2827 | 179.3343 | 483.451
0.67481 | 1010835.8457 | 79.788 | 192.6252 | 504.4512
0.69866 | 1054008.7322 | 83.509 | 201.6085 | 472.1821
0.68088 | 1029921.936 | 60.2913 | 145.556 | 472.5709
0.67276 | 975873.0749 | 72.4784 | 174.9783 | 511.7691
0.6553 | 959144.2015 | 69.7307 | 168.3448 | 510.2275
0.68197 | 1004724.0026 | 73.6134 | 177.7185 | 481.8344
0.69075 | 1050767.1585 | 72.4709 | 174.9602 | 505.6277
0.71469 | 1074631.4978 | 71.6337 | 172.9391 | 481.6658
0.68521 | 1028533.6991 | 73.2767 | 176.9055 | 489.9607
0.67034 | 990718.0991 | 73.8468 | 178.282 | 498.1302
0.689 | 1045152.9919 | 74.0514 | 178.7758 | 508.3209
0.67741 | 1003051.3055 | 72.2847 | 174.5106 | 497.3706
0.67811 | 1006519.394 | 79.0603 | 190.8684 | 507.4182
0.68764 | 1041461.8761 | 79.0419 | 190.824 | 504.4932
0.67716 | 1011966.9008 | 69.3921 | 167.5273 | 495.6109
0.71039 | 1073905.2026 | 80.5214 | 194.3959 | 494.0853
0.68965 | 1049419.9718 | 82.6569 | 199.5514 | 460.7688
0.68282 | 1025323.3732 | 63.9839 | 154.4708 | 497.1112
0.68898 | 1041364.7811 | 75.5221 | 182.3265 | 484.2718
0.70154 | 1077755.6916 | 69.4323 | 167.6243 | 493.421
0.67901 | 1010260.2162 | 77.6486 | 187.4604 | 482.5337
0.68846 | 1020855.7287 | 67.0169 | 161.7932 | 468.426
0.68761 | 1033714.2218 | 76.8052 | 185.4241 | 509.8642
0.67292 | 996709.6915 | 66.9421 | 161.6126 | 481.3315
0.6788 | 1009412.1573 | 72.5942 | 175.258 | 476.9197
0.72309 | 1122850.4156 | 89.0846 | 215.0693 | 472.2363
0.68658 | 1033851.763 | 69.9204 | 168.8027 | 511.1177
0.68885 | 1033259.8913 | 81.4786 | 196.7067 | 486.0712
0.69747 | 1048824.2685 | 90.6218 | 218.7803 | 476.3179
0.71852 | 1112295.5988 | 84.0893 | 203.0095 | 479.0858
0.69717 | 1069556.2945 | 67.3828 | 162.6766 | 491.4865
0.72006 | 1120341.1046 | 82.8751 | 200.0781 | 522.3591
0.66123 | 968259.3708 | 56.2422 | 135.7807 | 490.7359
0.68859 | 1041513.7765 | 72.0579 | 173.9632 | 496.2589
0.68387 | 1021806.5237 | 81.1621 | 195.9426 | 505.424
0.70211 | 1077934.3098 | 78.4365 | 189.3624 | 497.3583
0.7145 | 1111539.0646 | 75.9446 | 183.3465 | 493.3588
0.69523 | 1054137.946 | 72.2346 | 174.3898 | 501.4979
0.69308 | 1062599.1405 | 81.6625 | 197.1508 | 477.2003
0.68717 | 1030843.5674 | 76.6401 | 185.0257 | 481.1088
0.69507 | 1052264.1529 | 72.1231 | 174.1205 | 488.9603
0.66059 | 956671.7087 | 80.1648 | 193.535 | 470.6225
0.67323 | 1004521.8564 | 67.4362 | 162.8053 | 502.5849
0.66686 | 989351.1133 | 65.1624 | 157.3159 | 461.8909
0.68685 | 1042813.8774 | 77.5672 | 187.2639 | 471.2186
0.68444 | 1034232.6689 | 88.1698 | 212.8606 | 471.3176
0.66983 | 991767.8181 | 66.9177 | 161.5537 | 504.7113
0.69752 | 1053624.3242 | 68.269 | 164.816 | 510.4902
0.68593 | 1022487.1373 | 63.5181 | 153.3464 | 504.5739
0.69739 | 1050152.5002 | 81.7968 | 197.4749 | 516.1508
0.68477 | 1020816.3313 | 65.4763 | 158.0738 | 485.5685
0.71116 | 1083389.3687 | 74.3689 | 179.5424 | 496.4722
0.69913 | 1059259.2933 | 92.7215 | 223.8494 | 515.6028
0.68669 | 1031162.9241 | 53.0531 | 128.0814 | 508.8928
0.68334 | 1028589.4225 | 73.4816 | 177.4003 | 511.9342
0.68477 | 1010016.2719 | 72.0042 | 173.8335 | 482.1356
0.70408 | 1081995.1406 | 81.8543 | 197.6137 | 516.1386
0.68408 | 1035433.7222 | 70.2837 | 169.6799 | 527.4199
0.71303 | 1094505.6096 | 73.1287 | 176.5484 | 506.8966
0.69343 | 1051181.151 | 66.5822 | 160.7438 | 495.0669
0.68793 | 1040270.5165 | 78.9177 | 190.5242 | 465.0224
0.6875 | 1047864.1724 | 81.1838 | 195.995 | 499.2304
0.66956 | 980126.8734 | 62.375 | 150.5866 | 512.7461
0.68688 | 1040785.0327 | 72.1863 | 174.2731 | 478.9104
0.68275 | 1019035.0399 | 82.6194 | 199.4609 | 477.0822
0.69063 | 1027462.8897 | 47.5436 | 114.7803 | 481.6957
0.69412 | 1037287.3673 | 67.5398 | 163.0555 | 470.7008
0.68579 | 1019920.156 | 72.6019 | 175.2766 | 470.6809
0.68658 | 1035451.5588 | 70.0287 | 169.0641 | 492.4991
0.69046 | 1036076.7946 | 67.8188 | 163.7292 | 492.8594
0.68937 | 1050750.6279 | 72.6144 | 175.3066 | 488.9923
0.69564 | 1055349.1265 | 66.6336 | 160.8677 | 511.9808
0.69468 | 1062645.9541 | 74.6753 | 180.2822 | 495.8071
0.6908 | 1035673.5628 | 76.595 | 184.9166 | 481.0838
0.69032 | 1039468.7506 | 75.2378 | 181.6401 | 498.3954
0.67069 | 992623.7817 | 61.0612 | 147.4148 | 463.6775
0.7079 | 1091183.0884 | 83.7308 | 202.1439 | 489.627
0.66201 | 982555.3731 | 65.1886 | 157.3793 | 509.3559
0.7061 | 1078934.5526 | 79.7453 | 192.5222 | 505.5581
0.69023 | 1041358.1894 | 70.0561 | 169.1305 | 498.9488
0.6854 | 1026223.0523 | 79.7508 | 192.5355 | 480.0885
0.69994 | 1075366.8814 | 67.4804 | 162.912 | 476.1544
0.70792 | 1078274.9054 | 66.8999 | 161.5106 | 504.8351
0.70395 | 1082240.0444 | 75.0503 | 181.1874 | 514.3824
0.68579 | 1038856.893 | 58.1818 | 140.4632 | 466.5825
0.70438 | 1064898.1783 | 66.4024 | 160.3097 | 470.8838
0.69649 | 1055507.1387 | 79.0896 | 190.9392 | 474.6116
0.69154 | 1028796.4679 | 79.2408 | 191.3043 | 519.6396
0.68558 | 1021478.0356 | 61.4714 | 148.405 | 476.8779
0.69552 | 1066289.6229 | 83.2763 | 201.0468 | 499.0772
0.71998 | 1120473.149 | 82.8934 | 200.1224 | 475.7262
0.68759 | 1027179.0825 | 81.147 | 195.9062 | 483.9058
0.69257 | 1059249.7405 | 78.1383 | 188.6425 | 475.2192
0.66 | 976413.3506 | 75.582 | 182.4711 | 488.6136
0.66577 | 974748.4171 | 71.4731 | 172.5513 | 483.9151
0.71233 | 1096967.817 | 78.6766 | 189.9422 | 491.3698
0.68373 | 1007963.4392 | 83.1023 | 200.6267 | 490.872
0.70514 | 1071115.9356 | 62.6702 | 151.2992 | 473.3874
0.68415 | 1028503.5987 | 69.8059 | 168.5264 | 500.5118
0.70015 | 1068278.6246 | 62.9501 | 151.975 | 514.3748
0.70715 | 1080038.0131 | 87.6199 | 211.5332 | 480.6176
0.70807 | 1085138.1944 | 73.0103 | 176.2626 | 506.6766
0.69905 | 1070525.8753 | 78.2379 | 188.8829 | 507.9176
0.69075 | 1040670.3314 | 81.8372 | 197.5726 | 483.4544
0.66445 | 981193.2547 | 90.3712 | 218.1754 | 485.2648
0.69656 | 1051788.3543 | 64.341 | 155.3329 | 516.1148
0.70663 | 1082314.4051 | 81.5829 | 196.9585 | 487.44
0.69082 | 1024811.4675 | 82.1174 | 198.2489 | 487.2536
0.71055 | 1087789.7356 | 61.0192 | 147.3134 | 514.9293
0.70513 | 1079586.1385 | 79.9024 | 192.9014 | 501.0299
0.68755 | 1033668.4431 | 69.8097 | 168.5356 | 474.7164
0.68373 | 1019798.4505 | 76.2012 | 183.9661 | 484.095
0.69321 | 1039383.2793 | 73.1957 | 176.7099 | 497.2678
0.68557 | 1030612.3499 | 67.9421 | 164.0269 | 493.9364
0.70714 | 1082710.2298 | 65.2417 | 157.5074 | 497.1713
0.68949 | 1042372.1827 | 64.419 | 155.5212 | 512.6337
0.68733 | 1027445.5576 | 61.6731 | 148.8919 | 475.9169
0.67772 | 1015252.7377 | 74.8516 | 180.7077 | 511.1762
0.7032 | 1063016.5824 | 70.1938 | 169.4629 | 498.074
0.70305 | 1059434.4983 | 66.1865 | 159.7884 | 488.0432
0.68355 | 1036076.1551 | 77.4376 | 186.9508 | 474.9231
0.67315 | 989125.1595 | 75.071 | 181.2374 | 491.1991
0.67861 | 1013373.9993 | 72.8993 | 175.9944 | 484.6378
0.72502 | 1126711.1578 | 87.6027 | 211.4915 | 514.3408
0.69255 | 1059104.9179 | 72.8555 | 175.8888 | 505.5251
0.70041 | 1072731.5059 | 81.5784 | 196.9477 | 494.6534
0.69532 | 1040396.8018 | 61.7329 | 149.0365 | 485.3122
0.69317 | 1063811.618 | 80.0933 | 193.3624 | 492.4653
0.68963 | 1048737.4127 | 91.6444 | 221.2493 | 487.5622
0.6792 | 1009869.034 | 80.7657 | 194.9857 | 494.293
0.71123 | 1083159.9485 | 76.0955 | 183.7109 | 498.04
0.66935 | 991129.6971 | 73.9866 | 178.6195 | 500.8681
0.71198 | 1105049.3768 | 72.0699 | 173.992 | 476.0081
0.71757 | 1102533.1935 | 78.4982 | 189.5114 | 522.3358
0.68159 | 1043910.7553 | 79.9304 | 192.969 | 499.6175
0.67977 | 1010975.7913 | 95.0633 | 229.5031 | 489.734
0.68678 | 1041683.3956 | 74.0445 | 178.7591 | 508.7391
0.67084 | 991154.3309 | 58.8895 | 142.1719 | 502.0854
0.69515 | 1045881.4532 | 75.1765 | 181.4921 | 466.9751
0.67769 | 1019875.5532 | 64.1263 | 154.8145 | 498.2213
0.71852 | 1109342.4513 | 69.7785 | 168.4601 | 479.6368
0.70647 | 1084064.805 | 71.2522 | 172.018 | 490.8483
0.71106 | 1081059.1804 | 79.3444 | 191.5542 | 501.2394
0.67468 | 1001350.2061 | 65.1881 | 157.3781 | 487.088
0.68937 | 1043531.4021 | 66.4866 | 160.5129 | 483.8002
0.67747 | 1005882.7393 | 59.1801 | 142.8734 | 479.3322
0.69641 | 1053278.9136 | 82.155 | 198.3398 | 496.915
0.67499 | 1003005.1691 | 75.9165 | 183.2786 | 452.2837
0.69272 | 1051571.6583 | 58.3047 | 140.7601 | 495.2564
0.66443 | 978372.0867 | 56.5557 | 136.5374 | 466.6428
0.68806 | 1054805.5738 | 76.5984 | 184.925 | 490.6707
0.66819 | 997719.6044 | 67.4842 | 162.9213 | 499.2344
0.69848 | 1056445.2022 | 75.5739 | 182.4516 | 471.507
0.71305 | 1109139.8565 | 74.1048 | 178.9049 | 482.6819
0.69373 | 1033077.4986 | 83.945 | 202.6611 | 467.5823
0.67408 | 998296.1417 | 66.9331 | 161.5907 | 495.296
0.69702 | 1043975.5041 | 76.5206 | 184.737 | 499.4455
0.68151 | 1024509.9616 | 59.9183 | 144.6556 | 493.6511
0.69159 | 1040945.6093 | 86.3481 | 208.4627 | 495.8413
0.65722 | 949579.6872 | 85.6991 | 206.896 | 483.7104
0.69788 | 1067730.3024 | 79.8457 | 192.7645 | 509.7446
0.70718 | 1073572.923 | 72.3231 | 174.6034 | 491.7981
0.67246 | 1002971.1591 | 65.5091 | 158.153 | 485.894
0.68251 | 1023985.5668 | 79.914 | 192.9295 | 500.2956
0.7002 | 1064173.5663 | 79.3165 | 191.487 | 469.3672
0.68848 | 1045264.2789 | 66.6834 | 160.988 | 494.1688
0.68641 | 1023211.8559 | 71.0169 | 171.4498 | 486.396
0.68345 | 1008917.883 | 70.3275 | 169.7856 | 465.6556
0.69747 | 1048667.3653 | 80.0855 | 193.3436 | 476.1794
0.69635 | 1075816.1428 | 78.93 | 190.5539 | 502.2542
0.71297 | 1090890.2889 | 78.1703 | 188.7199 | 493.0978
0.68396 | 1021450.6546 | 76.2989 | 184.2018 | 497.9534
0.68359 | 1038530.3549 | 68.6622 | 165.7651 | 488.0233
0.68196 | 1016511.4681 | 69.2135 | 167.0963 | 482.508
0.69013 | 1046378.3797 | 60.4336 | 145.8995 | 492.5635
0.68088 | 1018402.233 | 79.6693 | 192.3387 | 499.7605
0.68372 | 1028700.9057 | 71.6782 | 173.0465 | 482.322
0.69274 | 1042018.4505 | 71.6612 | 173.0055 | 494.6393
0.67865 | 1015396.8019 | 67.6221 | 163.2542 | 447.9061
0.72185 | 1113494.0733 | 67.7519 | 163.5676 | 505.3147
0.69706 | 1056765.0136 | 75.7501 | 182.877 | 486.343
0.66443 | 1006911.2142 | 68.4919 | 165.3541 | 494.0862
0.72406 | 1130857.4971 | 62.2514 | 150.2881 | 510.0543
0.68928 | 1037441.2443 | 78.5491 | 189.6344 | 469.0478
0.68179 | 1021882.4714 | 69.5968 | 168.0215 | 500.2911
0.69473 | 1053129.4492 | 70.3888 | 169.9335 | 498.1305
0.68074 | 1022193.7302 | 71.5371 | 172.7057 | 472.2959
0.68875 | 1032277.4758 | 79.8318 | 192.731 | 494.3488
0.65052 | 947062.198 | 67.6014 | 163.2043 | 504.4374
0.68058 | 1012734.7169 | 77.6586 | 187.4845 | 455.1954
0.6866 | 1034179.441 | 84.9077 | 204.9854 | 501.8737
0.71907 | 1107478.1512 | 78.6858 | 189.9642 | 533.8346
0.7018 | 1064806.4947 | 72.2826 | 174.5055 | 482.8235
0.70473 | 1078559.2045 | 79.8724 | 192.8291 | 516.1706
0.67015 | 992889.9429 | 70.5845 | 170.4061 | 467.4272
0.70941 | 1079764.6923 | 72.5118 | 175.0589 | 498.5961
0.69904 | 1049963.2218 | 73.3235 | 177.0187 | 473.7016
0.68175 | 1033149.4895 | 70.3685 | 169.8845 | 508.3671
0.69762 | 1060874.6475 | 62.7401 | 151.468 | 502.3835
0.69755 | 1053823.0557 | 71.9734 | 173.7591 | 504.0657
0.68896 | 1032042.9795 | 63.8296 | 154.0982 | 474.5856
0.70576 | 1064774.668 | 71.875 | 173.5216 | 496.9737
0.68586 | 1040788.5541 | 74.0796 | 178.8439 | 475.2713
0.67583 | 1019593.128 | 85.0442 | 205.3149 | 493.0638
0.69682 | 1066418.7516 | 58.7804 | 141.9084 | 519.3635
0.70329 | 1081509.2371 | 76.8373 | 185.5016 | 507.6365
0.66581 | 985569.076 | 70.5205 | 170.2516 | 512.2522
0.66854 | 982343.7862 | 54.2219 | 130.9033 | 503.3822
0.6942 | 1042698.9338 | 67.8842 | 163.887 | 470.5857
0.6885 | 1039800.979 | 64.3517 | 155.3587 | 512.6263
0.69071 | 1038215.4142 | 66.0196 | 159.3855 | 472.2255
0.69566 | 1058536.4243 | 68.7711 | 166.0281 | 513.8969
0.69522 | 1042975.0237 | 71.4707 | 172.5456 | 470.7347
0.7185 | 1107260.4589 | 75.2309 | 181.6236 | 484.4094
0.71174 | 1087608.1494 | 46.0407 | 111.1521 | 479.6193
0.6961 | 1054870.2647 | 64.5722 | 155.8912 | 472.4093
0.69291 | 1069616.8184 | 63.687 | 153.7539 | 498.4309
0.70412 | 1088619.7439 | 75.4751 | 182.213 | 484.3841
0.69387 | 1044352.7401 | 69.8885 | 168.7257 | 487.0276
0.68284 | 1018845.698 | 72.8385 | 175.8477 | 477.3979
0.69603 | 1073644.1585 | 67.7511 | 163.5655 | 514.4154
0.72181 | 1119342.2085 | 57.2247 | 138.1525 | 487.0108
0.68318 | 1037270.9528 | 85.588 | 206.6277 | 468.4465
0.68785 | 1061254.3963 | 63.6538 | 153.6739 | 488.0342
0.68579 | 1008249.5557 | 62.646 | 151.2408 | 494.6313
0.70967 | 1093088.9205 | 65.6321 | 158.4499 | 546.9524
0.70434 | 1082582.2265 | 91.1669 | 220.0964 | 499.691
0.6995 | 1053240.1923 | 75.2338 | 181.6305 | 486.1237
0.70092 | 1076875.8986 | 76.7668 | 185.3314 | 505.6178
0.71554 | 1116424.2335 | 94.026 | 226.9988 | 503.487
0.67581 | 1002594.1435 | 61.2072 | 147.7673 | 500.0805
0.70804 | 1092132.4459 | 90.0021 | 217.2842 | 475.3561
0.70088 | 1053573.4738 | 66.0247 | 159.3977 | 511.7153
0.66598 | 973525.7773 | 73.0406 | 176.3357 | 499.5209
0.7098 | 1093896.8606 | 73.421 | 177.254 | 509.5811
0.68653 | 1046507.5589 | 70.7812 | 170.8809 | 479.4501
0.69613 | 1052776.5563 | 85.2511 | 205.8144 | 508.5142
0.68452 | 1029171.0504 | 64.0642 | 154.6648 | 486.7619
0.68578 | 1033001.4381 | 94.9364 | 229.1966 | 512.5334
0.68861 | 1033533.8171 | 69.6925 | 168.2526 | 470.5708
0.67921 | 1010711.9152 | 76.0251 | 183.5408 | 497.7733
0.67478 | 1016011.8137 | 87.8558 | 212.1027 | 491.98
0.69247 | 1070590.5435 | 61.6223 | 148.7693 | 502.7061
0.70178 | 1063726.7622 | 85.4439 | 206.2798 | 487.2907
0.68102 | 1022713.8833 | 75.0361 | 181.1532 | 477.6793
0.68646 | 1039902.3049 | 77.3945 | 186.8468 | 493.2098
0.68049 | 1036758.5348 | 64.9193 | 156.729 | 499.0903
0.69289 | 1040017.9429 | 73.2738 | 176.8986 | 487.4399
0.69616 | 1061037.5349 | 77.8667 | 187.9869 | 498.8477
0.68242 | 1023984.8879 | 71.4859 | 172.5822 | 496.7801
0.69856 | 1073390.2529 | 65.4743 | 158.069 | 505.7064
0.66313 | 974791.6796 | 66.4364 | 160.3916 | 482.343
0.66258 | 980951.3251 | 80.4209 | 194.1533 | 485.4754
0.65818 | 966648.9465 | 56.4094 | 136.1843 | 506.1469
0.68549 | 1037700.3033 | 77.3396 | 186.7142 | 499.9944
0.69468 | 1056664.7537 | 98.569 | 237.9667 | 487.5165
0.66917 | 996797.2746 | 52.6276 | 127.0544 | 474.7572
0.68962 | 1041926.0322 | 64.3793 | 155.4253 | 498.0958
0.67725 | 1015957.1654 | 76.6407 | 185.027 | 498.3116
0.67841 | 1015260.1631 | 59.6489 | 144.0051 | 456.7372
0.70326 | 1067843.6901 | 75.8698 | 183.166 | 531.3125
0.68103 | 1026407.5979 | 76.999 | 185.8919 | 458.1768
0.68188 | 1014382.5824 | 74.3311 | 179.4513 | 457.335
0.69517 | 1042632.9056 | 69.643 | 168.1332 | 495.1459
0.70922 | 1090387.8313 | 78.5468 | 189.6289 | 488.4499
0.67076 | 994125.1693 | 74.0542 | 178.7826 | 483.5037
0.66404 | 975276.517 | 64.508 | 155.7361 | 492.9024
0.69209 | 1036722.1826 | 74.8471 | 180.6969 | 509.2364
0.69798 | 1049490.7218 | 81.6756 | 197.1824 | 501.0469
0.71507 | 1092686.5078 | 78.0684 | 188.4739 | 479.0835
0.69062 | 1035483.4096 | 73.1943 | 176.7066 | 484.8486
0.6956 | 1059030.1456 | 70.6943 | 170.671 | 474.338
0.71149 | 1096172.7648 | 63.5642 | 153.4575 | 488.1851
0.67543 | 1013394.1749 | 67.81 | 163.7078 | 508.5767
0.68271 | 1022283.7466 | 52.9029 | 127.7189 | 485.2673
0.69875 | 1062170.5604 | 91.6069 | 221.1587 | 489.5125
0.66629 | 984718.4919 | 64.9127 | 156.7131 | 503.2123
0.68966 | 1045519.0406 | 80.749 | 194.9454 | 493.6961
0.68842 | 1039167.2557 | 66.3002 | 160.0629 | 480.1226
0.69437 | 1047934.4443 | 68.2085 | 164.6699 | 507.9697
0.68135 | 1001270.5018 | 70.0095 | 169.0178 | 521.0171
0.68442 | 1022526.7433 | 69.3731 | 167.4815 | 494.9464
0.71972 | 1106528.4476 | 81.891 | 197.7024 | 460.2407
0.68996 | 1041852.0401 | 86.627 | 209.136 | 486.5203
0.67726 | 1016769.1098 | 81.5503 | 196.8797 | 487.9473
0.7106 | 1096891.5764 | 65.774 | 158.7926 | 502.444
0.67904 | 1015695.5869 | 78.0624 | 188.4594 | 464.2005
0.68815 | 1018073.437 | 63.0397 | 152.1914 | 484.364
0.69982 | 1079161.4463 | 85.3599 | 206.077 | 479.3643
0.68742 | 1055893.2591 | 74.984 | 181.0274 | 472.7791
0.70335 | 1090199.6497 | 71.9733 | 173.7589 | 522.7605
0.69589 | 1064162.9849 | 65.1609 | 157.3123 | 494.2518
0.69356 | 1051014.984 | 53.5392 | 129.2552 | 474.6479
0.68318 | 1018172.7621 | 67.3571 | 162.6144 | 507.0116
0.67938 | 1012018.4798 | 82.4574 | 199.0698 | 505.7562
0.69469 | 1027447.2917 | 78.7783 | 190.1876 | 486.0206
0.69092 | 1055459.9313 | 89.136 | 215.1934 | 509.5956
0.67766 | 1018515.9758 | 62.0089 | 149.7028 | 483.4332
0.68432 | 1042736.8585 | 76.136 | 183.8086 | 510.7397
0.68757 | 1028625.3909 | 67.6138 | 163.2342 | 507.9746
0.67537 | 1013123.746 | 73.2312 | 176.7958 | 528.995
0.68212 | 1017531.4564 | 69.2664 | 167.2238 | 487.1519
0.68195 | 1034050.2209 | 68.5304 | 165.4471 | 466.0341
0.68722 | 1020060.7494 | 74.166 | 179.0526 | 485.9488
0.7005 | 1060809.583 | 66.5134 | 160.5775 | 498.4796
0.7064 | 1061560.9365 | 71.5933 | 172.8414 | 498.3619
0.66621 | 974153.0895 | 74.7049 | 180.3537 | 483.6943
0.70662 | 1063775.4864 | 64.1475 | 154.8657 | 462.4172
0.678 | 1010230.2935 | 83.9931 | 202.7773 | 472.9563
0.69551 | 1051914.0704 | 73.8333 | 178.2494 | 481.6687
0.68292 | 1018770.1891 | 79.4601 | 191.8336 | 485.9729
0.72187 | 1108786.7771 | 78.4725 | 189.4493 | 479.2857
0.70798 | 1089147.2204 | 78.5956 | 189.7466 | 502.9713
0.67286 | 1011920.1453 | 78.487 | 189.4843 | 484.6586
0.71117 | 1100803.1415 | 84.0803 | 202.9878 | 498.3336
0.69014 | 1069701.5534 | 77.4464 | 186.9721 | 458.2317
0.66607 | 997178.7734 | 67.6631 | 163.3533 | 487.6821
0.67758 | 1013578.7364 | 61.7164 | 148.9967 | 503.7948
0.69709 | 1052966.7247 | 75.4751 | 182.2129 | 508.1418
0.69926 | 1070321.3228 | 61.0093 | 147.2896 | 488.4377
0.66926 | 986283.5486 | 66.6043 | 160.797 | 491.2656
0.70414 | 1062050.5914 | 60.1177 | 145.137 | 495.7076
0.68292 | 1034914.4549 | 73.9794 | 178.602 | 520.6092
0.66835 | 999888.3585 | 63.7476 | 153.9004 | 503.5316
0.68372 | 1026125.0672 | 74.9895 | 181.0406 | 476.3268
0.6751 | 993719.2384 | 78.4029 | 189.2812 | 488.466
0.69111 | 1060300.6495 | 81.7188 | 197.2867 | 505.2363
0.70502 | 1078525.3648 | 71.0341 | 171.4916 | 502.9272
0.68075 | 1018681.9972 | 73.0502 | 176.3588 | 461.0326
0.66971 | 1004628.3526 | 74.954 | 180.955 | 518.5557
0.67923 | 1027172.6434 | 73.8244 | 178.2279 | 463.8111
0.70329 | 1072921.9681 | 77.9406 | 188.1653 | 522.0178
0.65179 | 964067.8087 | 64.9657 | 156.8411 | 494.2936
0.70785 | 1085947.0105 | 90.9966 | 219.6852 | 498.1201
0.68668 | 1036637.6724 | 62.2832 | 150.3649 | 509.9665
0.66717 | 1001194.3954 | 79.9319 | 192.9727 | 487.0282
0.68609 | 1041880.2562 | 75.8778 | 183.1853 | 492.3456
0.68081 | 1017845.679 | 67.5697 | 163.1278 | 462.7419
0.71093 | 1104223.3542 | 81.0595 | 195.695 | 498.5948
0.68027 | 1020343.9156 | 61.8826 | 149.3979 | 519.5047
0.69801 | 1080403.8575 | 80.7791 | 195.0179 | 501.2209
0.69082 | 1032123.2473 | 74.6526 | 180.2274 | 455.7762
0.70623 | 1088712.0028 | 63.6014 | 153.5475 | 478.1073
0.68122 | 1014728.2151 | 61.456 | 148.368 | 467.7069
0.68146 | 1025945.7425 | 78.5699 | 189.6845 | 505.9693
0.68976 | 1033671.4861 | 67.7301 | 163.515 | 508.5657
0.67415 | 1009301.3252 | 82.8742 | 200.0761 | 479.8498
0.68418 | 1018577.7585 | 77.6962 | 187.5752 | 514.5818
0.70837 | 1088162.1233 | 78.4809 | 189.4696 | 461.0022
0.67403 | 1010882.2715 | 67.7828 | 163.6421 | 517.1779
0.68288 | 1018961.1943 | 76.2331 | 184.043 | 483.4904
0.70112 | 1068808.08 | 87.4545 | 211.1337 | 479.3416
0.68675 | 1043413.683 | 62.6842 | 151.333 | 527.9965
0.6976 | 1056225.4426 | 70.305 | 169.7312 | 504.0136
0.65747 | 974554.4636 | 66.9922 | 161.7336 | 476.6632
0.69801 | 1075539.688 | 73.408 | 177.2226 | 500.0643
0.70684 | 1079185.5356 | 79.6797 | 192.3638 | 505.6592
0.69445 | 1051396.647 | 73.909 | 178.4322 | 485.3195
0.68085 | 1019137.8972 | 74.7927 | 180.5655 | 482.7295
0.68233 | 1018042.9529 | 69.2806 | 167.2583 | 490.4422
0.68432 | 1027519.061 | 63.9806 | 154.4628 | 496.3444
0.70802 | 1091346.1912 | 69.6645 | 168.185 | 503.3472
0.69041 | 1035165.8199 | 81.4447 | 196.6249 | 468.0738
0.69684 | 1054268.6161 | 75.5954 | 182.5035 | 476.0596
0.68824 | 1053461.1808 | 71.0752 | 171.5907 | 499.1749
0.66852 | 963262.5061 | 76.7479 | 185.2859 | 514.1856
0.68031 | 1025785.741 | 60.0986 | 145.091 | 476.5947
0.69016 | 1050301.8583 | 68.7795 | 166.0483 | 472.0723
0.67075 | 980738.8679 | 66.0826 | 159.5374 | 492.1768
0.68243 | 1040633.7558 | 91.8953 | 221.855 | 476.8053
0.70758 | 1084245.3207 | 78.2852 | 188.9971 | 502.4286
0.69307 | 1058860.903 | 76.3378 | 184.2957 | 449.7282
0.67286 | 999874.1899 | 50.0574 | 120.8492 | 498.7872
0.69413 | 1034841.7471 | 65.4811 | 158.0854 | 490.6315
0.67112 | 990860.4102 | 73.772 | 178.1014 | 468.7644
0.68901 | 1047811.5749 | 70.6147 | 170.479 | 488.7209
0.68419 | 1035983.0503 | 85.3544 | 206.0638 | 496.6162
0.6808 | 1019941.2 | 78.1855 | 188.7565 | 458.6582
0.68856 | 1026701.2029 | 71.0364 | 171.497 | 497.8031
0.68879 | 1048334.5917 | 79.4306 | 191.7624 | 491.769
0.70446 | 1070376.6294 | 88.6581 | 214.0396 | 468.9609
0.7017 | 1071968.9341 | 72.7788 | 175.7036 | 498.286
0.66894 | 993730.1995 | 75.8946 | 183.2258 | 526.9704
0.71894 | 1121132.7602 | 87.9062 | 212.2242 | 461.4047
0.69451 | 1057231.6185 | 68.7999 | 166.0977 | 494.0608
0.69426 | 1043375.5291 | 83.1902 | 200.8388 | 491.2985
0.69199 | 1041708.8231 | 62.1084 | 149.9429 | 512.7663
0.68793 | 1027390.5779 | 76.2991 | 184.2023 | 487.4558
0.70354 | 1082375.1694 | 81.9461 | 197.8353 | 515.985
0.67284 | 1003674.6896 | 74.5344 | 179.9418 | 485.7467
0.70097 | 1061355.3112 | 73.915 | 178.4467 | 470.9287
0.69682 | 1066808.1172 | 76.4968 | 184.6795 | 487.2921
0.69885 | 1071918.4089 | 65.3572 | 157.7863 | 485.5692
0.67749 | 1021374.3598 | 66.2163 | 159.8603 | 496.2238
0.67357 | 1005125.9723 | 67.4428 | 162.8213 | 505.6817
0.70881 | 1094639.7946 | 77.1647 | 186.2921 | 458.8724
0.68918 | 1035352.8982 | 71.5704 | 172.7862 | 485.5263
0.67139 | 993378.3997 | 66.8706 | 161.4398 | 465.6466
0.694 | 1061938.6488 | 84.7017 | 204.488 | 487.2066
0.69707 | 1053106.7061 | 70.1512 | 169.3601 | 485.7271
0.70803 | 1078357.3614 | 72.5588 | 175.1724 | 485.3723
0.69832 | 1067967.5289 | 72.8591 | 175.8974 | 490.4805
0.70947 | 1071719.5316 | 66.3487 | 160.18 | 478.0711
0.68729 | 1028179.6445 | 78.0731 | 188.4851 | 517.9386
0.68639 | 1044172.879 | 72.4738 | 174.9674 | 478.0972
0.68453 | 1041672.9032 | 56.1537 | 135.5669 | 476.9443
0.68788 | 1045291.3452 | 88.6673 | 214.0618 | 475.5804
0.69713 | 1059940.9944 | 80.9855 | 195.5164 | 513.1791
0.70523 | 1071353.6445 | 79.282 | 191.4038 | 490.386
0.69664 | 1044818.7666 | 81.4682 | 196.6817 | 513.2119
0.68837 | 1034004.2941 | 83.1485 | 200.7383 | 498.6159
0.69015 | 1028844.3895 | 69.2219 | 167.1164 | 485.8802
0.67534 | 1020297.6985 | 75.4888 | 182.2461 | 494.1105
0.66413 | 994222.7035 | 74.8007 | 180.5849 | 476.5279
0.69014 | 1037952.4921 | 67.9728 | 164.101 | 451.8598
0.68363 | 1022814.8715 | 65.7141 | 158.6478 | 492.703
0.71211 | 1105255.6189 | 66.1347 | 159.6634 | 480.338
0.69717 | 1044805.8673 | 76.5641 | 184.8421 | 501.2788
0.69748 | 1057719.488 | 79.1427 | 191.0675 | 491.1378
0.68951 | 1039978.1864 | 78.7145 | 190.0337 | 480.5574
0.67441 | 990391.0208 | 68.1901 | 164.6254 | 521.2544
0.69435 | 1043421.0456 | 66.003 | 159.3453 | 491.8244
0.67396 | 995927.6741 | 70.4 | 169.9607 | 487.0139
0.69161 | 1055138.0982 | 74.4771 | 179.8036 | 497.9827
0.66748 | 998591.4297 | 84.3339 | 203.6 | 469.7705
0.71168 | 1100737.3417 | 84.0144 | 202.8286 | 467.1373
0.68613 | 1030794.3508 | 83.1615 | 200.7696 | 503.0719
0.68523 | 1041276.6286 | 82.5448 | 199.2807 | 487.4965
0.67293 | 1006829.2305 | 87.3014 | 210.7643 | 489.8098
0.68055 | 1038874.4537 | 72.7242 | 175.5718 | 495.1119
0.69995 | 1050280.7449 | 59.8079 | 144.389 | 496.7202
0.6933 | 1049801.3602 | 62.4724 | 150.8217 | 492.5739
0.67318 | 994477.0312 | 70.3789 | 169.9097 | 457.6342
0.68703 | 1019592.8507 | 67.0728 | 161.928 | 501.6226
0.67462 | 1003343.2688 | 65.3243 | 157.7068 | 488.0278
0.68769 | 1055687.0472 | 88.0675 | 212.6136 | 495.17
0.67782 | 1001013.1201 | 64.8646 | 156.5969 | 489.1297
0.68363 | 1051182.0799 | 88.2127 | 212.9642 | 466.9307
0.66017 | 975553.0584 | 77.1202 | 186.1846 | 485.5861
0.71217 | 1110618.6813 | 79.6734 | 192.3486 | 468.3567
0.69493 | 1066241.9963 | 78.3875 | 189.2443 | 458.0427
0.69853 | 1067201.2369 | 70.9418 | 171.2686 | 485.2021
0.69978 | 1062031.2418 | 77.5365 | 187.1897 | 478.373
0.68477 | 1015505.4144 | 65.1358 | 157.2518 | 504.1249
0.66014 | 958140.5033 | 81.8172 | 197.5243 | 474.9755
0.70325 | 1080578.3501 | 71.3481 | 172.2495 | 477.8208
0.67282 | 995309.2205 | 66.8584 | 161.4104 | 502.8406
0.70752 | 1076070.5871 | 77.7017 | 187.5885 | 495.3942
0.7034 | 1077812.8487 | 78.2834 | 188.9929 | 493.6261
0.68953 | 1043870.2626 | 72.1225 | 174.1192 | 494.9289
0.68238 | 1029719.0421 | 75.3998 | 182.0312 | 507.1847
0.69759 | 1052869.4788 | 64.3122 | 155.2635 | 502.0266
0.68772 | 1034738.2016 | 76.8687 | 185.5774 | 494.2487
0.68821 | 1029723.6444 | 67.6232 | 163.2569 | 499.5825
0.6841 | 1015170.9864 | 67.0678 | 161.916 | 517.2932
0.69274 | 1058775.509 | 65.1255 | 157.227 | 503.3472
0.70272 | 1083456.796 | 71.592 | 172.8383 | 468.1888
0.7143 | 1086594.4986 | 70.6358 | 170.53 | 488.3976
0.71059 | 1094245.2053 | 73.3619 | 177.1114 | 501.8813
0.67607 | 1018332.1095 | 66.1578 | 159.7192 | 501.8961
0.67065 | 1000861.4398 | 65.5477 | 158.2461 | 503.9544
0.70386 | 1067645.3876 | 74.0478 | 178.7672 | 479.9495
0.67492 | 1016661.2122 | 64.1862 | 154.9591 | 523.6096
0.70061 | 1080457.2506 | 78.7888 | 190.2129 | 499.9908
0.71596 | 1110794.9438 | 89.0057 | 214.8788 | 482.3401
0.69177 | 1043448.0625 | 71.8598 | 173.4848 | 501.1715
0.69171 | 1037068.6804 | 82.9928 | 200.3624 | 463.5728
0.68337 | 1024993.103 | 61.8818 | 149.3959 | 493.01
0.67244 | 1012048.5967 | 57.8056 | 139.555 | 474.9825
0.69027 | 1033826.7552 | 78.1634 | 188.7031 | 468.4216
0.70097 | 1085416.8265 | 80.8111 | 195.0953 | 486.8329
0.66575 | 988959.8036 | 80.6909 | 194.8051 | 499.4467
0.68663 | 1040312.8518 | 60.389 | 145.7921 | 465.0531
0.67345 | 994173.8424 | 77.0422 | 185.9963 | 477.5368
0.69244 | 1036479.4106 | 69.5054 | 167.8009 | 495.264
0.68722 | 1048813.2886 | 70.866 | 171.0856 | 516.73
0.67447 | 1003603.911 | 65.3552 | 157.7813 | 488.1693
0.73035 | 1137132.2999 | 96.7162 | 233.4935 | 503.9948
0.68892 | 1036206.9544 | 59.9508 | 144.734 | 496.141
0.68861 | 1039802.5957 | 80.4299 | 194.175 | 479.3213
0.69726 | 1059152.3676 | 75.8995 | 183.2376 | 495.7814
0.69382 | 1043775.1829 | 78.512 | 189.5447 | 513.3429
0.66572 | 976558.697 | 74.944 | 180.9309 | 471.2999
0.71675 | 1127174.9243 | 85.4244 | 206.2327 | 495.8542
0.69911 | 1072725.5662 | 67.395 | 162.7059 | 476.3114
0.69771 | 1065912.5556 | 46.0298 | 111.1259 | 475.1754
0.68762 | 1044571.5823 | 75.9046 | 183.2499 | 501.0582
0.69458 | 1046155.9571 | 69.9373 | 168.8436 | 504.0997
0.68723 | 1030412.7168 | 63.8887 | 154.241 | 492.9495
0.71898 | 1108432.1841 | 67.7133 | 163.4744 | 484.4169
0.65565 | 940449.6722 | 66.6039 | 160.7961 | 468.3855
0.68706 | 1031318.5804 | 59.8292 | 144.4404 | 511.2731
0.66757 | 975368.4318 | 58.2355 | 140.5928 | 499.1415
0.68642 | 1038535.8287 | 83.7508 | 202.1922 | 482.2749
0.69442 | 1044659.0322 | 69.854 | 168.6426 | 497.6863
0.69092 | 1040319.9402 | 89.309 | 215.6109 | 477.9691
0.6915 | 1053713.0479 | 88.9133 | 214.6556 | 502.9335
0.68261 | 1027715.3985 | 64.1532 | 154.8796 | 478.4708
0.68248 | 1031259.1307 | 77.1984 | 186.3735 | 516.8125
0.67255 | 994987.0406 | 73.9213 | 178.4618 | 495.8255
0.68198 | 1008846.5147 | 75.5763 | 182.4573 | 490.114
0.71661 | 1127250.2196 | 64.8994 | 156.681 | 493.8264
0.69371 | 1052710.5467 | 66.6525 | 160.9133 | 492.3348
0.71074 | 1080725.5061 | 62.7346 | 151.4548 | 476.5724
0.69795 | 1067254.6345 | 76.4047 | 184.4572 | 475.7961
0.69795 | 1057691.9889 | 91.3248 | 220.4775 | 503.4599
0.69691 | 1069756.4213 | 78.4217 | 189.3267 | 502.3986
0.67906 | 1009187.0505 | 75.9272 | 183.3044 | 491.1171
0.68093 | 1016059.3861 | 88.506 | 213.6723 | 520.6153
0.69462 | 1052080.5316 | 77.3809 | 186.8141 | 483.5441
0.68329 | 1015847.5318 | 57.1013 | 137.8547 | 489.4466
0.67587 | 1029174.2704 | 82.7151 | 199.692 | 496.546
0.68709 | 1034034.6036 | 69.8807 | 168.707 | 481.1811
0.68958 | 1044502.1331 | 51.0238 | 123.1822 | 490.2462
0.7077 | 1102810.2323 | 76.0719 | 183.6539 | 489.191
0.67795 | 1004976.2192 | 83.5624 | 201.7376 | 485.7795
0.69087 | 1037811.5949 | 67.892 | 163.9059 | 493.1943
0.68541 | 1030069.2038 | 73.3653 | 177.1195 | 508.3712
0.68944 | 1029767.8756 | 65.6037 | 158.3814 | 495.0274
0.69327 | 1060286.239 | 60.9593 | 147.1687 | 490.7147
0.70477 | 1052345.1895 | 61.604 | 148.7253 | 498.5931
0.6851 | 1051804.4293 | 80.1219 | 193.4315 | 514.9981
0.6796 | 1009607.1414 | 80.6966 | 194.8189 | 473.4809
0.68855 | 1043838.7554 | 64.5976 | 155.9523 | 471.0773
0.71721 | 1101318.4893 | 69.4005 | 167.5476 | 469.4989
0.67978 | 1030501.0152 | 66.5825 | 160.7443 | 487.5962
0.69149 | 1059873.2224 | 58.8696 | 142.1239 | 497.5266
0.71007 | 1058255.4652 | 72.7553 | 175.6468 | 478.4859
0.67794 | 1015973.287 | 76.5631 | 184.8396 | 479.0602
0.67762 | 1016623.7393 | 76.6491 | 185.0472 | 512.3067
0.6878 | 1032989.6618 | 71.5507 | 172.7386 | 464.3482
0.69413 | 1046901.0231 | 73.6269 | 177.751 | 486.1521
0.7005 | 1049771.2271 | 61.8284 | 149.2671 | 482.4166
0.70739 | 1083262.5418 | 65.2519 | 157.532 | 506.3917
0.72768 | 1116263.6403 | 86.4725 | 208.763 | 496.0008
0.69479 | 1051762.5409 | 79.0016 | 190.7267 | 471.7508
0.70088 | 1045246.6605 | 84.067 | 202.9556 | 455.4804
0.67231 | 983946.3005 | 81.0129 | 195.5825 | 473.2004
0.70487 | 1066056.6693 | 77.2678 | 186.5411 | 518.5606
0.67069 | 996888.2765 | 83.7874 | 202.2807 | 500.5928
0.69793 | 1055615.7232 | 72.0564 | 173.9594 | 465.7233
0.69073 | 1030916.8511 | 50.2761 | 121.3773 | 492.9845
0.68598 | 1025717.11 | 59.5987 | 143.8839 | 506.7614
0.68118 | 1023524.7268 | 61.9923 | 149.6627 | 490.6463
0.70394 | 1055990.209 | 61.891 | 149.4181 | 492.6331
0.6973 | 1045288.2259 | 69.6074 | 168.0471 | 494.5667
0.68937 | 1036932.4846 | 76.8476 | 185.5266 | 497.943
0.70169 | 1057749.0643 | 80.2924 | 193.8431 | 469.5683
0.67461 | 998517.4951 | 73.9571 | 178.5483 | 476.8413
0.65755 | 961250.716 | 62.2907 | 150.3832 | 487.6721
0.68474 | 1028201.4302 | 78.3654 | 189.1907 | 494.1775
0.68196 | 1031203.277 | 70.2143 | 169.5124 | 466.6477
0.67935 | 998081.5252 | 54.2141 | 130.8843 | 491.7924
0.67653 | 1012057.1924 | 67.6104 | 163.2259 | 492.6453
0.68446 | 1043440.7522 | 70.7873 | 170.8956 | 517.9689
0.66325 | 982953.499 | 66.981 | 161.7066 | 497.1722
0.70077 | 1056435.1345 | 80.2986 | 193.8579 | 503.9413
0.67555 | 1017649.6936 | 78.6172 | 189.7987 | 475.9435
0.71551 | 1116750.5362 | 79.7394 | 192.5081 | 500.8811
0.69843 | 1044651.274 | 63.733 | 153.865 | 488.4571
0.6807 | 1026319.295 | 80.3094 | 193.884 | 477.7347
0.69766 | 1049976.2641 | 76.889 | 185.6265 | 494.502
0.66805 | 980516.3453 | 64.4653 | 155.6329 | 512.611
0.67789 | 996847.2786 | 75.1271 | 181.3729 | 500.4825
0.68288 | 1026851.3643 | 73.4065 | 177.219 | 494.2363
0.67355 | 1013915.7107 | 71.7145 | 173.1342 | 480.7664
0.69972 | 1070554.5973 | 65.1928 | 157.3894 | 492.0789
0.71308 | 1085592.0444 | 74.319 | 179.4219 | 487.8187
0.69671 | 1061826.9115 | 62.8784 | 151.802 | 464.7298
0.6617 | 987770.2203 | 65.566 | 158.2904 | 503.9189
0.69386 | 1053609.9804 | 64.7032 | 156.2073 | 485.2611
0.6746 | 1009745.8309 | 85.87 | 207.3086 | 540.5813
0.6919 | 1046472.0212 | 68.3563 | 165.0266 | 476.2001
0.69187 | 1043661.062 | 70.0441 | 169.1013 | 500.9561
0.69487 | 1052037.038 | 70.3017 | 169.7233 | 473.5886
0.69865 | 1051700.3239 | 83.3266 | 201.1682 | 497.2834
0.67638 | 1011863.4925 | 74.7493 | 180.4607 | 492.787
0.66303 | 977587.5329 | 83.1143 | 200.6558 | 514.5874
0.69737 | 1053406.0636 | 68.27 | 164.8183 | 481.1266
0.68754 | 1040081.953 | 64.4596 | 155.6193 | 467.4337
0.69733 | 1045160.7231 | 73.1678 | 176.6427 | 493.4897
0.70009 | 1066558.2311 | 70.7401 | 170.7817 | 519.8867
0.6836 | 1011067.334 | 78.7804 | 190.1928 | 514.305
0.68544 | 1031656.2905 | 81.757 | 197.3789 | 496.3196
0.69774 | 1054735.1192 | 73.7938 | 178.1541 | 469.2419
0.69971 | 1066782.3955 | 88.9603 | 214.7693 | 518.2867
0.67091 | 1004585.6584 | 73.0964 | 176.4703 | 493.1594
0.70616 | 1079914.5579 | 69.1163 | 166.8616 | 498.1165
0.72076 | 1120276.7933 | 75.8771 | 183.1835 | 510.7217
0.67711 | 1023558.1063 | 58.06 | 140.1692 | 494.7445
0.71833 | 1101812.0103 | 81.7006 | 197.2428 | 506.8561
0.69227 | 1031207.9902 | 61.4567 | 148.3695 | 477.4476
0.68561 | 1035901.9394 | 70.1587 | 169.3781 | 470.0118
0.7075 | 1086492.2699 | 76.7325 | 185.2486 | 506.9625
0.70295 | 1087273.4416 | 59.4074 | 143.4221 | 480.5702
0.68865 | 1031509.3514 | 61.7977 | 149.193 | 524.8128
0.70614 | 1085598.3402 | 78.5431 | 189.6199 | 490.6621
0.67314 | 992014.8163 | 60.8011 | 146.7867 | 473.9266
0.69804 | 1053885.7198 | 73.7027 | 177.9341 | 486.7439
0.67356 | 982706.9361 | 67.7461 | 163.5534 | 499.6465
0.67305 | 995391.3927 | 66.8408 | 161.3679 | 492.6379
0.6917 | 1041970.9705 | 66.1634 | 159.7326 | 501.991
0.68845 | 1030140.5024 | 78.0908 | 188.5279 | 506.444
0.67211 | 1006656.99 | 85.9241 | 207.439 | 472.574
0.70171 | 1077071.3575 | 65.4036 | 157.8981 | 489.6537
0.6946 | 1053012.202 | 86.9709 | 209.9664 | 501.4481
0.6706 | 986687.1601 | 73.5175 | 177.4869 | 522.7395
0.67567 | 1009542.8657 | 76.3347 | 184.2883 | 473.8883
0.69143 | 1048084.0962 | 77.4425 | 186.9628 | 465.8789
0.68952 | 1028755.6891 | 76.2243 | 184.0219 | 508.1592
0.70125 | 1059064.1305 | 80.4436 | 194.208 | 517.451
0.68984 | 1053517.0529 | 74.5534 | 179.9878 | 511.3446
0.68566 | 1016629.4702 | 72.3816 | 174.7447 | 480.9734
0.69292 | 1050540.3495 | 70.3995 | 169.9594 | 512.4786
0.68123 | 1013006.2193 | 61.3507 | 148.1137 | 518.5215
0.70741 | 1084389.4435 | 78.3149 | 189.0688 | 468.2691
0.68119 | 1021198.8034 | 57.5791 | 139.0081 | 477.8405
0.69722 | 1076361.1246 | 69.772 | 168.4446 | 491.2027
0.69058 | 1044289.3966 | 62.3887 | 150.6195 | 491.7917
0.69592 | 1066636.1285 | 67.3195 | 162.5237 | 497.6705
0.68093 | 1021301.5465 | 74.9422 | 180.9264 | 485.2816
0.70289 | 1084314.3792 | 77.0806 | 186.089 | 501.8126
0.67646 | 1016706.1858 | 71.6032 | 172.8653 | 487.6563
0.68944 | 1038846.2994 | 62.1665 | 150.0833 | 483.1593
0.68591 | 1052576.5595 | 73.1077 | 176.4977 | 525.8371
0.70073 | 1056412.6193 | 70.0031 | 169.0025 | 501.1681
0.69649 | 1052099.1505 | 78.8347 | 190.3238 | 500.9591
0.67609 | 993869.8962 | 73.452 | 177.3288 | 498.5374
0.69709 | 1060650.3045 | 74.2773 | 179.3214 | 464.8644
0.67031 | 998632.0616 | 72.7283 | 175.5815 | 514.2442
0.66656 | 986035.3847 | 86.3157 | 208.3846 | 474.0252
0.69019 | 1049675.5695 | 79.3705 | 191.6173 | 489.5908
0.71405 | 1084732.0225 | 74.1591 | 179.0359 | 547.4718
0.67566 | 1007248.3956 | 72.7744 | 175.693 | 472.2237
0.67876 | 1005045.722 | 56.8715 | 137.2999 | 470.9802
0.6702 | 997973.8913 | 83.9394 | 202.6475 | 458.9959
0.6896 | 1041949.4125 | 70.1597 | 169.3805 | 483.2127
0.68731 | 1026110.1036 | 67.4745 | 162.8978 | 509.6344
0.69421 | 1054743.9486 | 98.4563 | 237.6945 | 500.616
0.69988 | 1051523.9456 | 78.4099 | 189.2982 | 459.8602
0.68331 | 1032174.9496 | 60.2076 | 145.3541 | 495.5372
0.67352 | 1004918.6318 | 59.4699 | 143.573 | 481.1428
0.67569 | 1014540.6857 | 57.6693 | 139.226 | 527.3622
0.69459 | 1032181.4015 | 74.2512 | 179.2584 | 480.1298
0.67619 | 998350.054 | 60.9141 | 147.0597 | 483.1377
0.7089 | 1077745.0246 | 85.701 | 206.9005 | 501.7902
0.65473 | 957968.3603 | 82.4791 | 199.1222 | 456.6429
0.70641 | 1075619.6845 | 74.332 | 179.4533 | 461.4948
0.68824 | 1041738.7765 | 75.6306 | 182.5883 | 501.5686
0.68845 | 1028265.8825 | 67.504 | 162.969 | 481.3902
0.68703 | 1027755.7964 | 59.6251 | 143.9478 | 501.2274
0.67512 | 1004478.6219 | 77.6483 | 187.4597 | 500.0015
0.69848 | 1069041.4928 | 63.1458 | 152.4474 | 497.1432
0.68574 | 1025815.9006 | 61.7183 | 149.0011 | 467.7285
0.67789 | 1016662.8671 | 69.6398 | 168.1252 | 509.7243
0.67122 | 1000859.6213 | 52.681 | 127.1831 | 451.7011
0.69752 | 1043517.2289 | 71.2725 | 172.0671 | 472.225
0.68619 | 1034686.8482 | 77.0363 | 185.982 | 500.6095
0.69752 | 1030916.7672 | 78.7244 | 190.0575 | 474.3405
0.68286 | 1014582.8762 | 74.2395 | 179.2301 | 503.8578
0.67555 | 1020854.8585 | 82.0853 | 198.1714 | 502.7819
0.69205 | 1050340.3475 | 72.3051 | 174.56 | 497.9767
0.68782 | 1040522.8523 | 82.1743 | 198.3863 | 489.8868
0.68975 | 1035407.5163 | 73.2813 | 176.9166 | 502.1583
0.69798 | 1076662.285 | 67.7519 | 163.5676 | 523.323
0.68134 | 1027186.1208 | 86.4883 | 208.8011 | 483.9175
0.68149 | 1028916.9305 | 68.2369 | 164.7384 | 475.8234
0.70214 | 1072374.0168 | 58.6608 | 141.6198 | 474.5817
0.67914 | 1026036.4259 | 68.2806 | 164.8438 | 505.5348
0.65232 | 951837.7494 | 67.7549 | 163.5747 | 471.0086
0.70245 | 1060612.1876 | 70.1101 | 169.2607 | 489.894
0.70471 | 1070444.7513 | 74.1523 | 179.0194 | 483.0865
0.73349 | 1145053.1673 | 86.4138 | 208.6214 | 494.3595
0.6796 | 1015741.9878 | 67.5502 | 163.0807 | 498.7348
0.69795 | 1045261.7599 | 73.1092 | 176.5011 | 464.8417
0.67383 | 1005480.9182 | 69.289 | 167.2784 | 509.0988
0.69613 | 1035188.5595 | 79.2078 | 191.2246 | 494.7936
0.69541 | 1058767.8745 | 64.875 | 156.6222 | 473.804
0.68961 | 1041820.662 | 62.3293 | 150.4763 | 509.7033
0.67981 | 1017009.0402 | 73.0316 | 176.3138 | 482.7201
0.68307 | 1015748.4441 | 75.9718 | 183.4121 | 486.8916
0.70567 | 1085271.6281 | 81.9176 | 197.7666 | 512.2115
0.67885 | 1018020.0601 | 74.931 | 180.8995 | 479.5702
0.68882 | 1034916.6542 | 78.4107 | 189.3001 | 475.2191
0.68575 | 1033935.1304 | 71.8301 | 173.4133 | 500.8323
0.67643 | 1031636.4811 | 58.577 | 141.4175 | 479.5426
0.69357 | 1081557.5027 | 66.4467 | 160.4165 | 521.7701
0.70642 | 1086271.2853 | 59.061 | 142.5858 | 494.1306
0.70576 | 1083798.74 | 71.3065 | 172.1492 | 469.6639
0.67607 | 1006611.3488 | 72.6844 | 175.4757 | 492.4704
0.67745 | 1005793.7117 | 68.9395 | 166.4346 | 479.7266
0.71715 | 1116290.9303 | 77.7755 | 187.7666 | 480.0755
0.69066 | 1038040.7898 | 73.3705 | 177.1321 | 485.0604
0.69436 | 1055845.5158 | 77.6871 | 187.5531 | 504.7684
0.70015 | 1070205.86 | 78.0921 | 188.5311 | 495.5899
0.66773 | 1004617.2356 | 69.8618 | 168.6613 | 494.3846
0.66996 | 992726.5644 | 68.8045 | 166.1086 | 495.58
0.68727 | 1030345.5778 | 74.9083 | 180.8446 | 491.943
0.69382 | 1059008.9375 | 76.266 | 184.1225 | 459.3108
0.69302 | 1052088.3685 | 70.4924 | 170.1838 | 495.7245
0.69942 | 1078167.6572 | 71.5795 | 172.8082 | 502.3052
0.6912 | 1054478.742 | 77.9406 | 188.1653 | 500.3103
0.7018 | 1063278.7989 | 77.4039 | 186.8695 | 466.9465
0.67315 | 993366.6447 | 62.8799 | 151.8056 | 459.2267
0.71501 | 1093871.9506 | 72.8835 | 175.9565 | 506.5642
0.67579 | 1012668.7896 | 82.9487 | 200.256 | 479.4485
0.70228 | 1065705.0873 | 58.2842 | 140.7105 | 466.6414
0.67678 | 1020329.9089 | 81.8938 | 197.7091 | 457.3649
0.68915 | 1044156.7034 | 66.5476 | 160.6601 | 504.8227
0.67175 | 1015554.316 | 73.8023 | 178.1744 | 478.2103
0.708 | 1079793.8622 | 72.6581 | 175.4123 | 468.5859
0.69452 | 1060348.5882 | 53.9401 | 130.223 | 508.7409
0.69239 | 1049984.809 | 83.9428 | 202.6559 | 481.5799
0.68804 | 1036081.5975 | 71.7393 | 173.1939 | 505.4749
0.68537 | 1039196.8487 | 86.9852 | 210.0008 | 484.6096
0.70386 | 1074459.1696 | 68.9919 | 166.5613 | 478.4961
0.70628 | 1084332.3352 | 83.3955 | 201.3345 | 512.1917
0.68948 | 1049938.9667 | 85.8411 | 207.2388 | 492.6001
0.69106 | 1039370.4933 | 73.4224 | 177.2575 | 515.0712
0.68836 | 1053433.0362 | 67.216 | 162.2738 | 507.3793
0.71385 | 1095056.9572 | 78.3721 | 189.2071 | 477.884
0.69103 | 1037455.1028 | 73.2904 | 176.9386 | 487.6249
0.68308 | 1023163.3003 | 85.9302 | 207.4538 | 486.6278
0.68585 | 1031711.1625 | 71.6648 | 173.0141 | 478.8602
0.69482 | 1040650.0198 | 67.6903 | 163.4189 | 486.9661
0.69223 | 1029734.8033 | 63.3854 | 153.0258 | 476.8868
0.69781 | 1065383.8389 | 78.0012 | 188.3115 | 502.4457
0.66693 | 1005293.4085 | 87.9531 | 212.3376 | 489.5913
0.6982 | 1053190.8033 | 57.936 | 139.8699 | 476.9697
0.67732 | 1026906.893 | 79.1243 | 191.023 | 501.926
0.67148 | 985563.9254 | 82.7374 | 199.7456 | 485.5577
0.68932 | 1047436.7527 | 62.691 | 151.3494 | 507.9149
0.67539 | 1010108.1072 | 61.704 | 148.9666 | 508.0012
0.6944 | 1057658.8513 | 77.8161 | 187.8646 | 495.707
0.67945 | 1006021.5608 | 75.6445 | 182.6219 | 503.5905
0.66871 | 997234.1935 | 59.4397 | 143.5 | 494.1255
0.69756 | 1069034.2707 | 81.6299 | 197.0719 | 495.2317
0.67614 | 1014379.3202 | 73.2383 | 176.813 | 507.5691
0.66811 | 986603.1923 | 81.7428 | 197.3446 | 500.4776
0.69252 | 1022174.0588 | 80.1772 | 193.565 | 492.1029
0.67155 | 1004519.0222 | 79.675 | 192.3524 | 478.8638
0.67362 | 991677.2638 | 71.8667 | 173.5015 | 485.4267
0.67364 | 1009725.418 | 67.7438 | 163.5479 | 489.147
0.6959 | 1066107.9011 | 74.7868 | 180.5512 | 493.5386
0.68804 | 1020008.6273 | 61.1633 | 147.6612 | 473.5807
0.68376 | 1018168.1656 | 76.0759 | 183.6635 | 497.7152
0.71249 | 1101639.1365 | 65.8831 | 159.056 | 493.1503
0.68673 | 1046208.0389 | 74.3714 | 179.5484 | 501.4997
0.68504 | 1035965.0259 | 60.2763 | 145.5198 | 458.895
0.67985 | 1025239.7375 | 68.1568 | 164.5451 | 494.325
0.70167 | 1072332.0043 | 69.0702 | 166.7502 | 499.0913
0.68988 | 1046940.6416 | 75.8273 | 183.0633 | 473.2975
0.6879 | 1029870.7389 | 79.7442 | 192.5195 | 495.3253
0.68984 | 1056476.093 | 61.0417 | 147.3676 | 469.313
0.71177 | 1103247.0214 | 84.196 | 203.2672 | 510.6931
0.67104 | 999205.6244 | 80.8842 | 195.2716 | 493.3639
0.66023 | 975846.423 | 68.6312 | 165.6904 | 500.1342
0.6857 | 1025806.012 | 76.4295 | 184.5172 | 498.1497
0.68225 | 1028051.982 | 71.7874 | 173.3101 | 482.0898
0.69178 | 1038319.4983 | 63.9555 | 154.4023 | 476.9696
0.70624 | 1083820.105 | 74.9166 | 180.8646 | 528.4791
0.7121 | 1092172.9899 | 63.2778 | 152.7662 | 500.0944
0.72211 | 1131290.734 | 85.0993 | 205.4479 | 474.4547
0.6736 | 1000921.5733 | 83.7631 | 202.222 | 470.7508
0.661 | 981418.5852 | 79.0843 | 190.9265 | 468.1491
0.68339 | 1023396.1944 | 74.8257 | 180.6453 | 474.5544
0.69306 | 1054078.0512 | 74.2467 | 179.2474 | 470.3553
0.69018 | 1046687.7481 | 89.9523 | 217.1641 | 477.8305
0.68174 | 1018207.6523 | 67.5013 | 162.9626 | 493.6309
0.67524 | 1002130.1703 | 61.2308 | 147.8242 | 492.2486
0.69219 | 1037100.8874 | 74.8636 | 180.7367 | 497.5288
0.68748 | 1055402.7933 | 71.2847 | 172.0966 | 486.6987
0.70119 | 1073464.4576 | 71.087 | 171.6192 | 481.4505
0.70242 | 1050646.4977 | 52.8455 | 127.5804 | 465.2854
0.68353 | 1019568.6537 | 74.5312 | 179.9343 | 503.0604
0.68255 | 1031724.8923 | 68.3162 | 164.9299 | 473.525
0.68556 | 1032985.4582 | 66.1806 | 159.774 | 499.3955
0.69996 | 1073000.2261 | 58.8775 | 142.1428 | 500.9353
0.69805 | 1041880.6581 | 69.306 | 167.3195 | 488.9446
0.67517 | 1014596.9746 | 81.6277 | 197.0666 | 500.5453
0.66348 | 995053.4614 | 69.6391 | 168.1237 | 489.9627
0.67761 | 1014892.5538 | 76.5195 | 184.7343 | 473.1332
0.6771 | 1000659.1232 | 77.1268 | 186.2005 | 462.8467
0.71488 | 1099220.5444 | 71.3981 | 172.3703 | 513.8535
0.67394 | 994979.3284 | 78.6377 | 189.8481 | 492.5529
0.68168 | 1025605.9836 | 76.8651 | 185.5687 | 487.8049
0.68992 | 1028598.8428 | 74.4949 | 179.8466 | 501.1074
0.68615 | 1027349.4533 | 45.1117 | 108.9094 | 489.724
0.66329 | 985164.5475 | 57.0463 | 137.7219 | 499.5069
0.69019 | 1058499.3751 | 93.8553 | 226.5866 | 521.4541
0.68963 | 1024176.3875 | 79.1038 | 190.9734 | 507.5607
0.68626 | 1047257.7532 | 86.0231 | 207.6781 | 493.2557
0.70149 | 1069235.6355 | 84.3738 | 203.6963 | 504.6637
0.71247 | 1085816.3902 | 56.2379 | 135.7703 | 480.5265
0.70501 | 1066723.0095 | 80.5932 | 194.5692 | 474.4042
0.69352 | 1066718.0713 | 73.2772 | 176.9067 | 515.7846
0.68805 | 1043598.6781 | 80.7892 | 195.0423 | 480.9656
0.67771 | 1000116.2242 | 52.1376 | 125.8713 | 485.5507
0.69928 | 1071841.264 | 69.2749 | 167.2444 | 486.8972
0.69668 | 1064582.7255 | 79.7481 | 192.529 | 484.9759
0.67067 | 979775.1045 | 77.8143 | 187.8604 | 490.149
0.69531 | 1052383.9208 | 75.6262 | 182.5779 | 513.1398
0.69886 | 1057336.8887 | 68.3782 | 165.0795 | 479.9167
0.6967 | 1057653.0519 | 70.4915 | 170.1816 | 511.4738
0.68479 | 1036995.4874 | 88.3605 | 213.321 | 462.853
0.68697 | 1044063.992 | 72.405 | 174.801 | 485.9649
0.70148 | 1058128.3593 | 71.8623 | 173.491 | 486.6293
0.69271 | 1052239.073 | 74.154 | 179.0237 | 471.6512
0.67382 | 1014276.0591 | 71.7117 | 173.1273 | 503.4952
0.68361 | 1032128.0026 | 73.7054 | 177.9406 | 493.7672
0.67302 | 991634.3152 | 70.1944 | 169.4643 | 486.7651
0.68493 | 1054839.1119 | 82.0313 | 198.0411 | 471.7443
0.67648 | 1012060.4378 | 82.8142 | 199.9312 | 497.7584
0.68106 | 1006234.1824 | 83.2846 | 201.0669 | 487.7512
0.70172 | 1060578.9683 | 75.5187 | 182.3184 | 474.5789
0.68495 | 1024862.2259 | 71.2827 | 172.0916 | 496.7055
0.68929 | 1036210.3121 | 76.8032 | 185.4193 | 487.073
0.71028 | 1101202.1335 | 84.2158 | 203.3149 | 479.5967
0.71692 | 1101884.9432 | 76.7964 | 185.4029 | 485.2849
0.7109 | 1085050.3676 | 77.9774 | 188.2542 | 494.6551
0.68349 | 1036420.93 | 79.1364 | 191.0521 | 510.2238
0.71351 | 1094910.8008 | 78.3985 | 189.2707 | 486.9526
0.71638 | 1101727.2211 | 80.2615 | 193.7684 | 488.7726
0.6866 | 1039769.3822 | 66.5146 | 160.5804 | 487.6689
0.71837 | 1118664.4102 | 76.0192 | 183.5266 | 512.9012
0.70695 | 1071729.0382 | 75.748 | 182.8719 | 520.8156
0.68705 | 1039567.3156 | 48.354 | 116.7368 | 497.2011
0.69423 | 1041062.9528 | 81.4577 | 196.6564 | 468.7317
0.6932 | 1042019.352 | 75.1095 | 181.3304 | 486.2766
0.67839 | 1022097.819 | 64.1994 | 154.991 | 446.7758
0.67582 | 1005489.828 | 79.2481 | 191.3218 | 473.7348
0.67705 | 999284.8862 | 77.0274 | 185.9606 | 511.4344
0.66807 | 977328.6933 | 76.3472 | 184.3185 | 510.4927
0.71472 | 1117568.6796 | 74.4934 | 179.8431 | 496.1228
0.6994 | 1065780.4435 | 82.7745 | 199.8353 | 523.8686
0.69756 | 1083277.2884 | 73.9834 | 178.6117 | 507.6892
0.6606 | 982215.5293 | 74.2928 | 179.3588 | 498.3184
0.68577 | 1039737.236 | 45.6807 | 110.2831 | 505.2457
0.68734 | 1045055.5652 | 77.6775 | 187.53 | 511.8821
0.69104 | 1058896.0705 | 79.9694 | 193.0632 | 454.4997
0.6972 | 1055683.8765 | 75.658 | 182.6546 | 474.8167
0.69575 | 1057280.4829 | 74.1842 | 179.0964 | 459.2108
0.69381 | 1055259.9732 | 81.0132 | 195.5831 | 496.7833
0.69426 | 1059344.8782 | 77.9551 | 188.2002 | 484.136
0.66865 | 988729.3711 | 66.8305 | 161.3432 | 462.3543
0.65643 | 968211.2325 | 84.6166 | 204.2825 | 518.0518
0.66068 | 968002.2749 | 74.8541 | 180.7137 | 479.94
0.7154 | 1098216.0375 | 92.5111 | 223.3417 | 458.5498
0.70095 | 1059211.0496 | 88.1728 | 212.8681 | 490.8791
0.71811 | 1122880.1058 | 88.144 | 212.7985 | 503.1159
0.71352 | 1110007.151 | 81.1879 | 196.0049 | 442.1458
0.66197 | 964768.4086 | 72.8213 | 175.8061 | 475.2651
0.68546 | 1025858.0721 | 65.7334 | 158.6944 | 470.4555
0.68159 | 1008839.7235 | 72.2557 | 174.4407 | 467.4305
0.68844 | 1047823.9707 | 79.4315 | 191.7646 | 496.31
0.69501 | 1037897.2906 | 72.9011 | 175.9988 | 500.6461
0.67617 | 994624.2775 | 79.9024 | 192.9016 | 466.5552
0.68702 | 1039061.0204 | 62.3985 | 150.6432 | 511.9484
0.68101 | 1042601.4568 | 71.0886 | 171.6231 | 480.8433
0.68013 | 1014469.0922 | 72.8148 | 175.7905 | 496.9493
0.6808 | 1019982.2029 | 71.3759 | 172.3168 | 530.1154
0.68844 | 1039400.1876 | 62.2896 | 150.3803 | 483.4908
0.69523 | 1056180.4291 | 64.7327 | 156.2786 | 471.5596
0.70223 | 1053268.7552 | 79.8908 | 192.8734 | 511.924
0.67555 | 999809.6751 | 72.2489 | 174.4244 | 505.2899
0.6899 | 1033471.2191 | 59.7074 | 144.1465 | 484.8317
0.70298 | 1063488.7276 | 55.8251 | 134.7736 | 481.1858
0.67336 | 1010615.1966 | 73.2673 | 176.883 | 486.7869
0.69089 | 1038060.0739 | 65.9928 | 159.3208 | 494.427
0.7115 | 1103431.9005 | 79.232 | 191.283 | 480.6776
0.70244 | 1061367.237 | 80.481 | 194.2984 | 499.4922
0.69027 | 1038458.4969 | 67.9929 | 164.1493 | 497.8065
0.71528 | 1091129.3282 | 76.2211 | 184.014 | 502.6314
0.68689 | 1026973.7217 | 63.7073 | 153.8031 | 487.1539
0.68225 | 1022998.0672 | 69.6257 | 168.0914 | 471.0463
0.67756 | 1002275.6806 | 81.8835 | 197.6842 | 505.1407
0.68472 | 1033834.8218 | 73.707 | 177.9445 | 512.4899
0.69536 | 1065408.4513 | 84.8132 | 204.7571 | 498.6583
0.71584 | 1098053.1836 | 67.3739 | 162.655 | 502.6467
0.7161 | 1099696.764 | 71.3081 | 172.153 | 467.7475
0.68957 | 1047884.2839 | 72.3952 | 174.7775 | 479.6979
0.69359 | 1044256.0002 | 73.4978 | 177.4395 | 499.4649
0.71525 | 1092718.9833 | 96.2872 | 232.4577 | 500.2186
0.68298 | 1031758.0748 | 64.3709 | 155.4051 | 488.8059
0.69221 | 1058234.5187 | 83.0426 | 200.4826 | 490.3522
0.70426 | 1073076.6032 | 82.7666 | 199.8163 | 493.6929
0.68973 | 1055039.7148 | 74.6724 | 180.2752 | 498.7799
0.69293 | 1031116.9555 | 69.0967 | 166.8143 | 477.9095
0.67152 | 1012125.3145 | 73.5778 | 177.6325 | 508.5649
0.69578 | 1046805.0087 | 67.9974 | 164.1603 | 476.6481
0.67239 | 1015398.5923 | 80.4369 | 194.1918 | 515.0193
0.68248 | 1007048.3266 | 72.0335 | 173.9042 | 491.7203
0.67817 | 1015263.4465 | 84.3904 | 203.7365 | 501.3075
0.69954 | 1077667.03 | 76.9753 | 185.8347 | 491.4689
0.68001 | 1028411.7235 | 73.8291 | 178.2393 | 485.8842
0.70359 | 1079192.9478 | 78.3626 | 189.1841 | 492.3472
0.70166 | 1072774.3349 | 74.6376 | 180.1911 | 522.0232
0.67331 | 986460.5662 | 81.0996 | 195.7918 | 481.5911
0.70639 | 1091898.8823 | 67.8926 | 163.9072 | 488.3726
0.68667 | 1043813.5855 | 88.6981 | 214.1361 | 514.7629
0.70097 | 1069096.7649 | 81.2377 | 196.1252 | 480.6222
0.67962 | 1017055.336 | 81.2898 | 196.2509 | 471.441
0.70703 | 1112494.0372 | 76.8126 | 185.4421 | 484.793
0.69054 | 1029546.1266 | 79.4141 | 191.7225 | 472.2701
0.68418 | 1030478.109 | 69.9372 | 168.8434 | 500.3291
0.67346 | 1008417.3737 | 78.1453 | 188.6595 | 492.1712
0.69761 | 1043433.0631 | 63.7341 | 153.8677 | 501.1624
0.68791 | 1020204.9886 | 68.8644 | 166.2533 | 501.9167
0.67499 | 996881.2609 | 70.3601 | 169.8642 | 480.5694
0.68275 | 1016717.3966 | 69.1473 | 166.9363 | 500.1254
0.69892 | 1055618.0831 | 71.9538 | 173.7119 | 484.144
0.68153 | 1014033.8254 | 79.2508 | 191.3284 | 540.3554
0.68074 | 1012311.5023 | 67.2093 | 162.2576 | 495.5293
0.69289 | 1048550.8985 | 78.9772 | 190.6677 | 509.9496
0.69019 | 1032625.1154 | 61.7271 | 149.0225 | 492.3521
0.68074 | 1015321.8734 | 77.8394 | 187.9208 | 506.0218
0.68224 | 1021434.3222 | 79.7472 | 192.5269 | 497.0426
0.68274 | 1018987.7303 | 71.1032 | 171.6582 | 519.0254
0.69965 | 1069443.8752 | 79.7722 | 192.5871 | 504.8685
0.69256 | 1061533.196 | 84.846 | 204.8363 | 501.003
0.69266 | 1034483.2614 | 79.5507 | 192.0524 | 489.8982
0.70081 | 1058426.2777 | 71.9508 | 173.7047 | 497.477
0.6728 | 1005363.8727 | 72.9473 | 176.1104 | 490.7872
0.68979 | 1043209.6804 | 78.928 | 190.5491 | 484.5537
0.68856 | 1045603.1027 | 68.6311 | 165.6902 | 510.8079
0.68058 | 1024054.9555 | 80.1459 | 193.4892 | 484.4331
0.69421 | 1045447.9953 | 69.9277 | 168.8203 | 488.3454
0.67624 | 1021191.0646 | 80.435 | 194.1872 | 460.278
0.67015 | 1006233.8456 | 61.9484 | 149.5566 | 474.4608
0.6902 | 1032224.3506 | 59.6095 | 143.9102 | 477.5336
0.69824 | 1075247.3375 | 97.0559 | 234.3137 | 507.7507
0.71188 | 1111644.6923 | 84.8233 | 204.7814 | 493.106
0.69538 | 1073297.2693 | 85.4384 | 206.2665 | 489.0872
0.68871 | 1044650.2897 | 87.0183 | 210.0807 | 472.3998
0.65811 | 959369.0256 | 69.4492 | 167.6653 | 492.5316
0.69657 | 1050752.2189 | 68.1756 | 164.5906 | 480.9528
0.70483 | 1075109.7395 | 64.9962 | 156.9146 | 454.6197
0.70109 | 1056687.5023 | 73.5774 | 177.6315 | 495.0768
0.69143 | 1048875.6518 | 66.6285 | 160.8555 | 488.2147
0.70284 | 1071727.293 | 85.9571 | 207.5188 | 463.1782
0.67426 | 988826.218 | 49.4019 | 119.2667 | 501.0662
0.6651 | 973289.5557 | 62.3546 | 150.5374 | 495.2146
0.68338 | 1036566.1754 | 68.5533 | 165.5024 | 508.4877
0.66669 | 987485.697 | 65.0561 | 157.0592 | 506.2007
0.70174 | 1062877.7628 | 66.5253 | 160.6063 | 481.8799
0.69248 | 1065556.5666 | 78.6134 | 189.7896 | 484.8677
0.6942 | 1054658.3058 | 70.5443 | 170.309 | 483.9004
0.7078 | 1074452.7509 | 75.8497 | 183.1174 | 494.0312
0.69029 | 1038582.6219 | 83.2846 | 201.0669 | 466.5483
0.69745 | 1045384.0773 | 74.8926 | 180.8067 | 478.3957
0.68644 | 1046867.6854 | 77.9151 | 188.1037 | 478.76
0.69921 | 1072963.0492 | 71.255 | 172.0247 | 465.433
0.74229 | 1166589.4904 | 66.967 | 161.6726 | 495.8834
0.71336 | 1112725.7725 | 72.4298 | 174.861 | 489.5541
0.68043 | 1007481.7513 | 86.3949 | 208.5758 | 489.7911
0.68238 | 1033294.1293 | 75.6618 | 182.6638 | 493.4315
0.67215 | 992927.9799 | 66.7643 | 161.1834 | 492.2234
0.72394 | 1118094.0742 | 73.579 | 177.6354 | 476.3809
0.67336 | 1019090.2751 | 80.6132 | 194.6175 | 510.8467
0.69221 | 1040687.0472 | 76.8091 | 185.4336 | 514.5375
0.65595 | 960635.8697 | 73.1753 | 176.6608 | 499.4664
0.70745 | 1083871.0918 | 84.8076 | 204.7436 | 506.1777
0.67611 | 1005755.315 | 75.9983 | 183.4761 | 496.822
0.68318 | 1008935.8841 | 75.4498 | 182.152 | 475.5276
0.7122 | 1089776.9441 | 67.2079 | 162.2543 | 510.7823
0.69276 | 1051110.2702 | 62.5991 | 151.1275 | 518.8288
0.67412 | 1024454.6169 | 88.6736 | 214.0769 | 499.3579
0.70103 | 1065675.6465 | 75.9561 | 183.3743 | 493.042
0.69668 | 1068313.1639 | 73.0537 | 176.3672 | 492.1441
0.67531 | 1000308.1979 | 74.0123 | 178.6816 | 513.6563
0.6634 | 968828.0221 | 58.2093 | 140.5297 | 520.3209
0.67153 | 1008498.2848 | 79.9928 | 193.1198 | 460.4877
0.70737 | 1090617.3665 | 67.719 | 163.4881 | 488.0073
0.72644 | 1142480.4585 | 83.7698 | 202.2382 | 500.2735
0.70368 | 1081096.6577 | 75.0001 | 181.0663 | 490.4406
0.67619 | 1017033.4823 | 67.9766 | 164.11 | 495.195
0.67727 | 1009785.1223 | 57.2654 | 138.2508 | 475.7415
0.69529 | 1053536.8536 | 80.7091 | 194.849 | 518.17
0.70612 | 1079559.5292 | 67.1508 | 162.1164 | 500.2766
0.69604 | 1048527.1455 | 60.0427 | 144.956 | 494.5885
0.71822 | 1099627.593 | 79.9029 | 192.9026 | 475.1667
0.7361 | 1155723.1079 | 72.9044 | 176.0069 | 501.3419
0.69224 | 1021068.8481 | 70.2714 | 169.6502 | 500.616
0.68679 | 1031772.9975 | 67.8979 | 163.9201 | 512.2433
0.70762 | 1081186.5094 | 82.9962 | 200.3706 | 518.2539
0.6779 | 1001817.3849 | 66.7909 | 161.2475 | 496.7439
0.68768 | 1038456.8199 | 66.3265 | 160.1264 | 501.2717
0.69846 | 1072141.9146 | 71.277 | 172.0779 | 498.5789
0.68483 | 1025665.1376 | 65.7819 | 158.8115 | 471.6586
0.65481 | 963233.108 | 68.3049 | 164.9027 | 478.1147
0.71483 | 1099981.7795 | 63.4864 | 153.2698 | 492.0705
0.65876 | 955214.409 | 61.7858 | 149.1641 | 481.6291
0.68538 | 1041145.234 | 72.3696 | 174.7156 | 472.8772
0.67849 | 1019742.9956 | 81.6398 | 197.096 | 518.1733
0.67945 | 1014905.7534 | 59.5369 | 143.7349 | 486.3531
0.69966 | 1069562.27 | 67.1427 | 162.0968 | 472.9022
0.69676 | 1053277.6495 | 83.6789 | 202.0187 | 492.0988
0.67038 | 988813.3209 | 81.6479 | 197.1155 | 494.9799
0.66563 | 982936.6556 | 75.4441 | 182.1383 | 523.9202
0.6971 | 1041272.0772 | 67.5092 | 162.9816 | 471.4385
0.66078 | 968731.776 | 69.8434 | 168.6168 | 485.5733
0.71931 | 1125609.4777 | 68.7311 | 165.9316 | 492.338
0.69307 | 1039716.3122 | 63.923 | 154.3238 | 481.335
0.68391 | 1001986.9726 | 68.0302 | 164.2393 | 485.1498
0.70064 | 1064837.3323 | 75.9385 | 183.3317 | 462.1161
0.6718 | 1016497.2148 | 70.2594 | 169.6211 | 493.8376
0.69188 | 1050111.9341 | 75.8372 | 183.0872 | 508.7457
0.65239 | 948494.4969 | 65.7085 | 158.6344 | 495.0599
0.68668 | 1020982.1573 | 69.0399 | 166.6772 | 487.9832
0.66484 | 972297.1344 | 77.8971 | 188.0603 | 469.0425
0.68534 | 1031583.549 | 81.7634 | 197.3943 | 461.9283
0.68977 | 1036247.8506 | 69.7584 | 168.4116 | 505.4539
0.6815 | 1036370.2018 | 75.9848 | 183.4436 | 499.8156
0.65908 | 963099.3249 | 66.0434 | 159.4428 | 483.5666
0.69834 | 1054087.6865 | 60.1624 | 145.2448 | 480.2958
0.65289 | 950771.0775 | 75.9998 | 183.4797 | 487.6751
0.67244 | 1003751.7443 | 65.5626 | 158.2822 | 518.3602
0.66024 | 965280.0586 | 69.651 | 168.1523 | 487.7386
0.70007 | 1084719.8566 | 75.6393 | 182.6095 | 484.7416
0.67599 | 1010330.2294 | 65.6458 | 158.4829 | 513.6004
0.66724 | 989224.5658 | 61.1667 | 147.6694 | 492.4736
0.66749 | 987425.7221 | 68.6903 | 165.8331 | 488.2202
0.67581 | 1010549.7436 | 85.7836 | 207.1 | 490.6988
0.69488 | 1051246.0154 | 75.5907 | 182.4921 | 459.525
0.69451 | 1054852.8897 | 70.5268 | 170.2667 | 489.2495
0.69491 | 1057529.0683 | 58.4505 | 141.112 | 483.1178
0.68358 | 1026880.1212 | 75.0592 | 181.209 | 504.1221
0.68823 | 1031145.3719 | 61.8144 | 149.2331 | 470.1424
0.69842 | 1045958.3987 | 73.1088 | 176.5003 | 490.0104
0.6942 | 1042016.5749 | 78.3366 | 189.1213 | 471.9902
0.69035 | 1035032.1718 | 71.4265 | 172.4387 | 481.8276
0.71244 | 1120771.9247 | 71.0992 | 171.6485 | 488.6019
0.68892 | 1052179.8802 | 60.8748 | 146.9648 | 507.2367
0.69235 | 1063343.7744 | 58.9891 | 142.4122 | 491.0508
0.70552 | 1081882.4608 | 73.055 | 176.3704 | 556.3507
0.69776 | 1086149.6686 | 86.1666 | 208.0245 | 497.516
0.69091 | 1051218.2928 | 68.7654 | 166.0143 | 488.2998
0.68355 | 1024621.8673 | 74.8978 | 180.8192 | 490.0647
0.67879 | 1025781.3194 | 77.2056 | 186.3908 | 520.1959
0.68003 | 1017618.0909 | 59.6454 | 143.9968 | 490.4505
0.67724 | 1027122.4377 | 64.6246 | 156.0175 | 497.0844
0.70085 | 1050240.919 | 69.5824 | 167.9867 | 464.6936
0.70817 | 1076713.96 | 77.6773 | 187.5295 | 507.2288
0.70459 | 1069407.3001 | 72.3079 | 174.5667 | 486.3929
0.68938 | 1041788.0694 | 78.8666 | 190.4008 | 479.6906
0.68849 | 1030240.7779 | 81.283 | 196.2344 | 472.5322
0.67453 | 1007449.4473 | 77.9458 | 188.1777 | 463.5584
0.66571 | 988497.9023 | 90.8712 | 219.3826 | 502.5017
0.69097 | 1056001.3511 | 79.7589 | 192.5551 | 525.9474
0.68371 | 1032818.4879 | 82.0554 | 198.0993 | 505.0575
0.6801 | 1015119.6563 | 79.5031 | 191.9375 | 513.1907
0.68867 | 1036391.2458 | 78.5389 | 189.6096 | 491.3285
0.70791 | 1065000.6299 | 80.1333 | 193.4588 | 465.5199
0.69894 | 1065410.2415 | 85.9271 | 207.4463 | 461.5682
0.67257 | 1004990.9265 | 72.9456 | 176.1063 | 485.6335
0.69558 | 1048413.6133 | 75.3113 | 181.8176 | 507.1025
0.69501 | 1044149.1037 | 73.3405 | 177.0596 | 491.3021
0.72447 | 1112818.0525 | 83.4375 | 201.436 | 468.2365
0.66576 | 997853.7849 | 71.4043 | 172.3852 | 496.9776
0.67946 | 1014140.4742 | 79.5017 | 191.9342 | 497.4252
0.70994 | 1089797.4352 | 81.7637 | 197.395 | 486.5531
0.67896 | 1021215.7757 | 86.2876 | 208.3167 | 482.8937
0.67686 | 1019522.6392 | 68.0764 | 164.351 | 510.7895
0.66423 | 973560.9976 | 74.8821 | 180.7813 | 514.0751
0.65703 | 942000.9217 | 74.8243 | 180.6419 | 475.7463
0.70013 | 1068444.6177 | 86.0247 | 207.6821 | 483.9189
0.68262 | 1020844.9258 | 82.7814 | 199.852 | 469.9967
0.6905 | 1063540.0868 | 76.9607 | 185.7996 | 525.0018
0.70028 | 1056882.4154 | 75.411 | 182.0582 | 468.9181
0.66423 | 984869.5945 | 65.1241 | 157.2235 | 503.4159
0.67409 | 1003207.7415 | 86.8385 | 209.6468 | 474.1131
0.70899 | 1089830.0894 | 69.4725 | 167.7216 | 495.7307
0.71052 | 1095311.4105 | 80.452 | 194.2284 | 482.0398
0.67771 | 1020402.3285 | 80.1986 | 193.6165 | 487.3315
0.68894 | 1045481.2689 | 72.2957 | 174.5373 | 493.657
0.69973 | 1074980.0199 | 67.4763 | 162.9021 | 485.3537
0.6826 | 1015005.6483 | 65.3104 | 157.6732 | 492.9024
0.67936 | 1024669.2955 | 81.9295 | 197.7953 | 483.5218
0.69095 | 1035496.8028 | 73.1599 | 176.6235 | 476.9746
0.68573 | 1035530.0024 | 62.3036 | 150.4142 | 488.1943
0.70432 | 1067145.4224 | 79.0732 | 190.8997 | 523.2088
0.68022 | 1024302.4536 | 68.0575 | 164.3053 | 495.4557
0.65858 | 969350.3639 | 79.9524 | 193.0221 | 518.7152
0.70101 | 1088928.1853 | 72.1298 | 174.1367 | 493.5777
0.68069 | 999151.8663 | 64.4709 | 155.6466 | 504.2225
0.7057 | 1075715.0468 | 62.8898 | 151.8294 | 494.0775
0.68473 | 1005545.1424 | 71.6897 | 173.0742 | 478.3119
0.68922 | 1046505.7402 | 72.3371 | 174.6372 | 483.0682
0.72828 | 1146461.3944 | 82.1554 | 198.3408 | 530.5086
0.68008 | 1008345.8756 | 70.6363 | 170.5311 | 490.4074
0.70852 | 1079194.2699 | 70.7267 | 170.7493 | 484.9072
0.72906 | 1137241.3602 | 70.499 | 170.1995 | 462.0637
0.6901 | 1029599.4448 | 67.4301 | 162.7907 | 506.952
0.67602 | 1007124.9077 | 83.9802 | 202.7462 | 506.3735
0.68725 | 1030259.6508 | 65.8442 | 158.9619 | 481.8457
0.67589 | 1017866.5846 | 90.7547 | 219.1012 | 467.5032
0.69331 | 1056653.7313 | 70.7686 | 170.8505 | 487.311
0.67789 | 1014705.301 | 61.756 | 149.0922 | 489.0999
0.66869 | 981533.3896 | 66.3405 | 160.1602 | 504.9753
0.69786 | 1053180.5522 | 60.1524 | 145.2207 | 483.6988
0.68045 | 1014594.9108 | 79.4216 | 191.7407 | 475.1506
0.69482 | 1069304.4159 | 75.1281 | 181.3753 | 489.3805
0.66582 | 978122.7682 | 68.2144 | 164.6842 | 492.1984
0.70376 | 1094269.6056 | 79.4381 | 191.7806 | 491.0198
0.68203 | 1025046.5538 | 75.0955 | 181.2965 | 468.0841
0.69052 | 1034635.6089 | 84.4621 | 203.9095 | 524.6666
0.69363 | 1056643.2051 | 87.3926 | 210.9845 | 494.1748
0.6948 | 1057228.4316 | 77.7389 | 187.6783 | 498.3127
0.69253 | 1037177.5307 | 79.7729 | 192.5888 | 497.0354
0.66586 | 986564.5256 | 78.9191 | 190.5276 | 484.0461
0.6836 | 1030225.5997 | 73.5706 | 177.6152 | 507.811
0.71668 | 1103656.1652 | 78.6749 | 189.9381 | 477.2134
0.71105 | 1095497.5615 | 63.564 | 153.457 | 498.0985
0.68481 | 1029346.9634 | 78.4444 | 189.3816 | 478.2572
0.67759 | 996244.8749 | 84.3478 | 203.6335 | 505.2825
0.69982 | 1069758.3535 | 63.0666 | 152.2563 | 511.4903
0.68975 | 1027823.4918 | 79.3722 | 191.6213 | 469.8175
0.66902 | 1003080.9282 | 84.5176 | 204.0435 | 471.2515
0.68803 | 1021766.6175 | 65.2266 | 157.4708 | 460.9521
0.68657 | 1035620.3561 | 73.636 | 177.7731 | 516.6112
0.68977 | 1053353.53 | 86.0845 | 207.8264 | 485.2266
0.68548 | 1049789.0464 | 87.8582 | 212.1084 | 473.0397
0.67813 | 1013974.9909 | 63.7132 | 153.8174 | 455.557
0.67281 | 1012407.8618 | 69.8717 | 168.6853 | 509.3091
0.67751 | 1009159.6636 | 76.0982 | 183.7174 | 460.3651
0.70222 | 1060760.9776 | 66.3478 | 160.1777 | 476.2075
0.72418 | 1129809.2554 | 72.4428 | 174.8925 | 466.351
0.69613 | 1053122.0561 | 55.8254 | 134.7745 | 487.7437
0.67583 | 1004674.7211 | 63.3438 | 152.9254 | 474.1575
0.66532 | 974727.8431 | 73.2026 | 176.7267 | 486.0516
0.70705 | 1071559.4673 | 77.428 | 186.9278 | 493.1138
0.67655 | 1018214.8376 | 76.89 | 185.6289 | 502.8199
0.70112 | 1058576.5722 | 78.7954 | 190.229 | 478.8724
0.67265 | 984644.3563 | 69.7375 | 168.3611 | 478.1975
0.67735 | 1004176.1415 | 67.0029 | 161.7593 | 494.7888
0.68123 | 1029239.7949 | 55.7523 | 134.5979 | 515.2304
0.68395 | 1023061.7592 | 65.6996 | 158.6129 | 493.4142
0.66973 | 993317.5947 | 68.8693 | 166.2652 | 504.0519
0.64463 | 915322.0043 | 62.3656 | 150.5638 | 500.254
0.71131 | 1104966.5685 | 57.3232 | 138.3904 | 497.1846
0.68797 | 1043465.1221 | 64.6287 | 156.0275 | 503.6257
0.7051 | 1079622.1422 | 78.2255 | 188.8532 | 492.4832
0.6849 | 1023091.2584 | 74.6387 | 180.1938 | 488.1739
0.69032 | 1028817.838 | 84.0113 | 202.8212 | 485.4459
0.68024 | 1030247.2818 | 80.6707 | 194.7563 | 511.5601
0.71099 | 1081104.4984 | 75.9766 | 183.4236 | 487.9859
0.67693 | 996308.2803 | 62.714 | 151.4049 | 463.6699
0.67795 | 1011470.0686 | 74.5472 | 179.9728 | 478.874
0.69147 | 1058846.6444 | 61.0346 | 147.3505 | 518.7967
0.67613 | 1009050.3688 | 71.0981 | 171.6459 | 456.1412
0.7184 | 1120391.5292 | 74.299 | 179.3736 | 484.9828
0.7055 | 1077896.2727 | 72.7874 | 175.7244 | 512.9026
0.68435 | 1044915.6994 | 69.0083 | 166.6009 | 478.4659
0.69731 | 1045582.1528 | 76.6057 | 184.9424 | 490.2472
0.66755 | 991923.1888 | 80.7147 | 194.8626 | 497.2152
0.67507 | 1017677.2074 | 78.6742 | 189.9363 | 472.0458
0.66851 | 979614.7546 | 73.2191 | 176.7665 | 494.4174
0.70703 | 1092404.4535 | 75.426 | 182.0946 | 505.3185
0.69391 | 1049113.8423 | 58.0692 | 140.1915 | 470.3195
0.70392 | 1081869.8712 | 80.2094 | 193.6426 | 487.0295
0.68458 | 1025966.183 | 73.161 | 176.6263 | 505.6955
0.70594 | 1093287.9173 | 69.9947 | 168.9822 | 484.6772
0.68307 | 1034249.1394 | 70.3066 | 169.7351 | 468.7189
0.71959 | 1116729.3085 | 87.4814 | 211.1989 | 501.4838
0.69973 | 1065812.6782 | 76.1074 | 183.7396 | 470.0756
0.67568 | 1008762.0935 | 72.8822 | 175.9532 | 499.5672
0.68946 | 1039730.0691 | 75.3512 | 181.9138 | 490.5193
0.71451 | 1093929.7505 | 72.9389 | 176.0902 | 476.9913
0.67565 | 1000592.5229 | 70.5529 | 170.3297 | 488.0041
0.67174 | 1003213.9125 | 67.4977 | 162.9538 | 495.9676
0.67666 | 1010844.2162 | 71.1693 | 171.8179 | 532.4927
0.69629 | 1055307.9612 | 77.4318 | 186.9368 | 511.8475
0.68677 | 1012882.5705 | 56.6467 | 136.7571 | 486.3766
0.67396 | 996492.3878 | 78.7554 | 190.1324 | 485.3262
0.69117 | 1043156.3119 | 68.2119 | 164.6781 | 510.9731
0.694 | 1038382.07 | 76.4415 | 184.5462 | 538.7217
0.68597 | 1036543.263 | 77.1991 | 186.3752 | 485.4264
0.72271 | 1110835.0005 | 78.5261 | 189.5788 | 449.4596
0.67539 | 1011210.3266 | 74.8102 | 180.6078 | 487.311
0.68749 | 1028213.5676 | 85.8008 | 207.1416 | 489.0198
0.68315 | 1012607.505 | 72.3596 | 174.6914 | 470.39
0.69535 | 1041810.8544 | 78.1915 | 188.7709 | 502.6695
0.69492 | 1043380.258 | 76.708 | 185.1894 | 481.5003
0.67807 | 1015296.8937 | 67.6724 | 163.3757 | 494.6765
0.69468 | 1063320.5971 | 69.1795 | 167.0141 | 476.6442
0.70333 | 1074241.1471 | 72.7641 | 175.668 | 476.7144
0.69628 | 1063997.5838 | 69.0638 | 166.7348 | 491.5237
0.67722 | 1021310.3007 | 70.0267 | 169.0594 | 492.8791
0.69692 | 1029944.5201 | 70.4057 | 169.9743 | 467.2374
0.69138 | 1059567.3451 | 76.5754 | 184.8695 | 475.3722
0.69973 | 1046839.8692 | 65.7099 | 158.6377 | 486.1747
0.68204 | 1026043.966 | 80.1296 | 193.4501 | 486.1316
0.70068 | 1065443.9908 | 72.4415 | 174.8894 | 469.5263
0.70019 | 1071772.0937 | 78.2014 | 188.7948 | 485.4648
0.69223 | 1035394.6147 | 61.7103 | 148.9819 | 501.4947
0.69194 | 1038800.0171 | 71.5224 | 172.6705 | 479.9275
0.66721 | 992617.5249 | 57.1406 | 137.9495 | 456.1959
0.69123 | 1044825.6775 | 77.2241 | 186.4355 | 507.0503
0.67621 | 1005033.0339 | 69.0142 | 166.6151 | 470.9067
0.67517 | 1017806.6017 | 86.4825 | 208.7871 | 492.291
0.69199 | 1061918.2302 | 71.2571 | 172.0299 | 497.7711
0.6704 | 994588.8319 | 74.129 | 178.9631 | 480.8755
0.69531 | 1056589.4911 | 68.6793 | 165.8064 | 501.6427
0.69541 | 1046887.792 | 75.2208 | 181.599 | 493.5309
0.6852 | 1019317.6942 | 70.8716 | 171.0993 | 475.8195
0.69541 | 1051406.4298 | 70.2049 | 169.4896 | 486.0908
0.70719 | 1090968.1689 | 67.758 | 163.5823 | 500.2708
0.68929 | 1051688.9258 | 65.0129 | 156.955 | 507.4204
0.69432 | 1047883.0961 | 64.3089 | 155.2554 | 491.6802
0.70239 | 1072571.029 | 98.9548 | 238.898 | 482.3307
0.68705 | 1040656.9673 | 72.1603 | 174.2103 | 499.0757
0.68297 | 1043322.4034 | 48.8184 | 117.8579 | 502.844
0.68615 | 1022020.0893 | 85.4505 | 206.2959 | 480.6073
0.70335 | 1079915.4058 | 56.6582 | 136.785 | 465.156
0.70242 | 1078145.313 | 73.1239 | 176.5368 | 485.5321
0.67482 | 1011737.7569 | 78.2444 | 188.8986 | 488.4173
0.69937 | 1073609.1008 | 84.9754 | 205.1488 | 489.8246
0.68695 | 1027576.0331 | 76.4222 | 184.4995 | 504.1304
0.69065 | 1047130.4648 | 72.2302 | 174.3792 | 493.9016
0.69835 | 1064927.324 | 88.9781 | 214.8121 | 537.4192
0.66723 | 993304.7424 | 57.1786 | 138.0413 | 469.0546
0.68795 | 1038061.5927 | 90.9332 | 219.5322 | 460.0659
0.69875 | 1058006.8221 | 82.2479 | 198.5639 | 494.4762
0.71083 | 1096201.338 | 86.9316 | 209.8714 | 488.7625
0.68047 | 1016548.4655 | 67.5169 | 163.0001 | 479.3999
0.70192 | 1071657.1421 | 65.0555 | 157.058 | 494.333
0.66881 | 996618.3662 | 76.1301 | 183.7942 | 503.2864
0.68444 | 1046594.5435 | 74.6479 | 180.216 | 488.6785
0.65434 | 948953.44 | 70.7973 | 170.9198 | 494.9258
0.69124 | 1050751.3744 | 85.689 | 206.8714 | 488.4641
0.7002 | 1075000.9864 | 76.7124 | 185.2001 | 488.0782
0.68729 | 1037730.8588 | 70.1103 | 169.2612 | 473.5573
0.70007 | 1058617.4863 | 72.0404 | 173.9209 | 506.4707
0.71372 | 1116217.6384 | 66.6401 | 160.8834 | 484.5648
0.67487 | 1009571.3669 | 65.7051 | 158.6262 | 488.0692
0.69134 | 1022433.2383 | 63.0171 | 152.1368 | 476.2137
0.70105 | 1066276.4359 | 81.0144 | 195.5861 | 482.7381
0.70989 | 1074398.4021 | 77.3227 | 186.6735 | 472.464
0.70339 | 1075223.9643 | 74.6233 | 180.1565 | 489.7413
0.68652 | 1040913.9608 | 70.4044 | 169.9713 | 517.6881
0.65968 | 972224.5734 | 75.2946 | 181.7773 | 481.6827
0.68716 | 1051112.8223 | 81.4767 | 196.7022 | 496.0769
0.71824 | 1091609.2719 | 82.5568 | 199.3098 | 489.1365
0.69993 | 1073955.1631 | 74.9041 | 180.8345 | 495.1095
0.69056 | 1031597.9077 | 79.5704 | 192.0999 | 533.356
0.70352 | 1062849.4909 | 73.7513 | 178.0515 | 497.2197
0.70286 | 1088813.1856 | 75.6235 | 182.5712 | 479.5111
0.68679 | 1049899.1871 | 83.0382 | 200.472 | 525.7967
0.70089 | 1073230.3632 | 72.9492 | 176.115 | 519.045
0.68413 | 1034641.8931 | 72.0491 | 173.9418 | 497.8634
0.68346 | 1033953.3337 | 70.2469 | 169.591 | 526.8074
0.68508 | 1018397.5875 | 72.5687 | 175.1964 | 467.9225
0.67987 | 1014753.0445 | 72.863 | 175.9068 | 507.5961
0.69871 | 1069679.8808 | 78.2148 | 188.8273 | 483.1038
0.67832 | 1007390.3215 | 70.7518 | 170.81 | 512.9223
0.69859 | 1049101.5234 | 65.9598 | 159.241 | 509.6504
0.71613 | 1104773.6701 | 83.7995 | 202.3099 | 469.8165
0.68035 | 1011565.2846 | 69.0394 | 166.6757 | 476.4529
0.67065 | 991059.9137 | 83.3026 | 201.1104 | 478.7146
0.69337 | 1046647.1352 | 66.3003 | 160.0631 | 481.4009
0.67881 | 1004189.5145 | 73.9169 | 178.4511 | 493.653
0.69309 | 1043454.3002 | 73.4949 | 177.4324 | 502.2885
0.69753 | 1060256.9399 | 80.9626 | 195.4611 | 491.5483
0.68838 | 1046995.2537 | 64.8085 | 156.4616 | 485.765
0.68834 | 1042123.0463 | 83.8048 | 202.3227 | 510.7402
0.68343 | 1037864.369 | 74.1348 | 178.9773 | 498.4722
0.69483 | 1050194.657 | 75.5212 | 182.3243 | 478.7473
0.68029 | 1013804.15 | 72.75 | 175.6341 | 505.2083
0.67636 | 1028910.0673 | 58.4282 | 141.0582 | 463.8692
0.69919 | 1079190.1526 | 75.3488 | 181.908 | 491.1312
0.69288 | 1061164.0909 | 69.2184 | 167.108 | 466.9636
0.67857 | 1018098.2306 | 69.6677 | 168.1928 | 465.1355
0.69256 | 1044985.7065 | 77.0871 | 186.1048 | 490.5436
0.69989 | 1065917.8815 | 74.348 | 179.4919 | 496.81
0.69178 | 1047275.36 | 62.4586 | 150.7885 | 510.7874
0.70967 | 1090936.0006 | 63.4227 | 153.1159 | 473.2389
0.66976 | 999386.149 | 65.5388 | 158.2246 | 483.8903
0.68228 | 1046066.788 | 67.3408 | 162.5751 | 468.6693
0.6981 | 1060503.6924 | 75.9047 | 183.2502 | 481.3207
0.71147 | 1091260.2784 | 67.3678 | 162.6402 | 501.4741
0.66621 | 978748.4463 | 79.8028 | 192.661 | 495.3449
0.7089 | 1097522.6822 | 68 | 164.1666 | 470.3098
0.68202 | 1028892.4102 | 62.2402 | 150.2612 | 490.4503
0.69675 | 1048189.5438 | 81.718 | 197.2848 | 496.6438
0.69076 | 1048198.173 | 68.5827 | 165.5733 | 519.0745
0.68967 | 1043855.6685 | 72.107 | 174.0817 | 485.8713
0.69608 | 1060386.325 | 87.3933 | 210.9861 | 502.462
0.69177 | 1038640.9323 | 86.1346 | 207.9474 | 516.3544
0.68888 | 1034466.9525 | 76.7196 | 185.2175 | 491.4218
0.67877 | 1021141.0514 | 59.9631 | 144.7636 | 489.0025
0.68733 | 1018846.0393 | 74.0663 | 178.812 | 515.0151
0.70518 | 1072966.4681 | 91.6943 | 221.3697 | 492.1232
0.70251 | 1067149.739 | 87.1462 | 210.3896 | 517.5216
0.67502 | 1010184.0182 | 61.7423 | 149.0592 | 508.3689
0.69402 | 1056796.466 | 72.5432 | 175.1347 | 519.5159
0.68526 | 1028806.8236 | 69.7135 | 168.3032 | 495.6633
0.67571 | 999305.4553 | 63.0169 | 152.1363 | 499.6212
0.69196 | 1035819.6188 | 76.4777 | 184.6334 | 504.9936
0.7078 | 1072441.7739 | 68.4791 | 165.3231 | 473.2935
0.70342 | 1075620.5583 | 71.0042 | 171.4193 | 448.0847
0.69475 | 1050274.5031 | 62.3703 | 150.5753 | 484.5907
0.68896 | 1055420.3546 | 74.7839 | 180.5442 | 474.3249
0.70939 | 1077602.4025 | 64.7275 | 156.266 | 495.9467
0.70186 | 1070146.4403 | 72.6389 | 175.3658 | 484.3778
0.68088 | 1024025.2447 | 64.0852 | 154.7153 | 475.9201
0.67697 | 1007533.7991 | 85.3811 | 206.1281 | 483.2503
0.6907 | 1058513.6732 | 65.3013 | 157.6513 | 492.6426
0.72957 | 1142624.453 | 76.456 | 184.5812 | 503.3167
0.67858 | 1030108.0365 | 66.6754 | 160.9686 | 525.7161
0.71455 | 1098952.0283 | 63.4519 | 153.1864 | 471.1969
0.71417 | 1084149.0992 | 66.6758 | 160.9696 | 510.022
0.69027 | 1049649.8522 | 60.6101 | 146.3258 | 491.2593
0.67456 | 997682.6641 | 63.0213 | 152.147 | 507.7017
0.68105 | 1016778.8734 | 72.8826 | 175.9542 | 486.5731
0.67882 | 1023928.7126 | 77.0627 | 186.0457 | 467.6402
0.68944 | 1049912.9866 | 62.8287 | 151.6819 | 476.2146
0.70502 | 1085391.8312 | 73.3441 | 177.0683 | 510.867
0.67291 | 1007077.765 | 74.7797 | 180.5341 | 482.7198
0.70709 | 1082083.8598 | 69.1643 | 166.9775 | 489.5913
0.6843 | 1030930.01 | 69.9548 | 168.8858 | 470.4782
0.68414 | 1039916.3009 | 62.7122 | 151.4007 | 481.5461
0.70877 | 1086025.5806 | 67.2999 | 162.4763 | 493.3548
0.71193 | 1102765.0973 | 73.7945 | 178.1557 | 501.3036
0.67282 | 1029509.7489 | 92.2113 | 222.6179 | 506.5197
0.67465 | 1005850.1843 | 92.6142 | 223.5904 | 514.1486
0.68708 | 1030652.5214 | 88.9743 | 214.8031 | 480.967
0.68715 | 1036476.5581 | 71.8593 | 173.4837 | 524.8449
0.71717 | 1111930.2522 | 63.9672 | 154.4304 | 496.8161
0.68345 | 1041263.8379 | 70.7447 | 170.7928 | 513.2013
0.69374 | 1042992.3794 | 67.9483 | 164.0417 | 499.3274
0.6858 | 1031166.1639 | 55.4848 | 133.9522 | 499.2156
0.69518 | 1041645.0102 | 61.8195 | 149.2454 | 471.4285
0.69356 | 1044322.9673 | 71.7346 | 173.1826 | 459.3511
0.69072 | 1049812.5942 | 70.5747 | 170.3823 | 473.625
0.69986 | 1063624.3091 | 66.7512 | 161.1516 | 473.777
0.68297 | 1020894.1491 | 78.0099 | 188.3325 | 497.9752
0.65359 | 953514.8125 | 67.7428 | 163.5455 | 476.2596
0.70973 | 1082475.9824 | 82.8481 | 200.0131 | 503.1343
0.70536 | 1071576.2933 | 68.6605 | 165.7611 | 495.0962
0.68536 | 1018949.5148 | 72.5782 | 175.2194 | 509.0617
0.69117 | 1054930.4518 | 67.0381 | 161.8442 | 510.3246
0.69514 | 1060417.884 | 74.469 | 179.784 | 498.607
0.71473 | 1099913.6695 | 65.5738 | 158.3092 | 436.9398
0.70138 | 1072581.0641 | 65.1621 | 157.3152 | 504.098
0.69845 | 1082642.7696 | 77.4507 | 186.9824 | 478.4895
0.70975 | 1102359.3679 | 70.1967 | 169.4698 | 474.2108
0.67673 | 1010818.7309 | 65.6052 | 158.3849 | 522.5069
0.68381 | 1043153.8367 | 85.9934 | 207.6064 | 483.3308
0.65663 | 956888.7039 | 55.9709 | 135.1258 | 490.8444
0.68909 | 1033428.001 | 74.9342 | 180.9072 | 506.5753
0.70495 | 1083283.5109 | 75.0171 | 181.1073 | 507.3441
0.69077 | 1027576.7455 | 67.2317 | 162.3117 | 481.4012
0.68262 | 1025209.1401 | 71.5502 | 172.7375 | 441.5339
0.70009 | 1059850.953 | 62.4586 | 150.7884 | 487.0942
0.67939 | 1029656.2539 | 79.0941 | 190.95 | 468.0585
0.68204 | 1028448.3528 | 66.2304 | 159.8944 | 475.8366
0.70265 | 1068774.5704 | 66.8081 | 161.2891 | 496.5253
0.68508 | 1015891.7193 | 68.8566 | 166.2346 | 512.2683
0.68548 | 1034355.5961 | 83.5273 | 201.6526 | 496.1196
0.68939 | 1042495.0213 | 80.547 | 194.4576 | 487.0823
0.69068 | 1036730.5554 | 86.1115 | 207.8915 | 481.8139
0.70666 | 1082138.1138 | 78.2354 | 188.877 | 508.5651
0.66897 | 998176.6903 | 76.2307 | 184.0373 | 504.2805
0.69562 | 1051541.3547 | 73.7951 | 178.157 | 503.2134
0.69716 | 1054694.6282 | 83.743 | 202.1735 | 479.5881
0.6828 | 1032982.5176 | 73.8541 | 178.2996 | 494.1489
0.66108 | 962143.6067 | 67.5811 | 163.1551 | 492.8835
0.67921 | 1019148.0376 | 74.9737 | 181.0026 | 514.721
0.69727 | 1049294.2311 | 73.4633 | 177.3561 | 485.5191
0.68003 | 1036103.1338 | 64.9224 | 156.7365 | 483.0678
0.6882 | 1028976.2135 | 71.2312 | 171.9673 | 488.6236
0.69516 | 1042933.8341 | 69.6642 | 168.1843 | 489.587
0.68757 | 1041424.4375 | 82.2745 | 198.6282 | 513.3379
0.70881 | 1056396.8234 | 69.205 | 167.0757 | 461.4953
0.68729 | 1029424.6689 | 79.7801 | 192.6063 | 520.1752
0.70248 | 1064948.546 | 72.2228 | 174.3613 | 475.5121
0.68531 | 1025397.3782 | 78.0867 | 188.5179 | 463.8308
0.67426 | 1010635.0443 | 63.8685 | 154.1921 | 488.3148
0.69501 | 1048505.7407 | 85.0423 | 205.3102 | 526.7974
0.68837 | 1035504.2869 | 73.4349 | 177.2876 | 493.6337
0.70913 | 1108671.322 | 70.6602 | 170.5889 | 493.4903
0.68112 | 1014485.2476 | 67.3158 | 162.5147 | 479.243
0.71757 | 1109097.9232 | 69.856 | 168.6473 | 456.4192
0.69818 | 1060950.4203 | 70.5611 | 170.3497 | 475.9357
0.72203 | 1117569.1899 | 80.7782 | 195.0159 | 498.5898
0.69606 | 1054326.0674 | 86.8966 | 209.787 | 473.6638
0.69291 | 1048762.6329 | 77.3274 | 186.6849 | 494.514
0.67022 | 987203.165 | 71.9052 | 173.5946 | 488.0846
0.70754 | 1068178.1991 | 71.9233 | 173.6381 | 491.8668
0.7009 | 1054744.7584 | 80.1547 | 193.5106 | 501.5528
0.67673 | 1007040.0602 | 80.8331 | 195.1485 | 476.0744
0.67273 | 1011598.4672 | 67.9618 | 164.0742 | 513.0306
0.70293 | 1070131.6221 | 87.3375 | 210.8514 | 473.2678
0.67088 | 995228.2931 | 70.6736 | 170.6211 | 495.1705
0.67964 | 1024832.0707 | 80.318 | 193.9049 | 499.4988
0.67621 | 1010447.4504 | 87.1908 | 210.4973 | 484.1024
0.6707 | 992440.3794 | 75.5973 | 182.508 | 498.3583
0.68302 | 1021874.2445 | 63.7502 | 153.9065 | 521.0983
0.67756 | 1003508.4889 | 74.0026 | 178.6582 | 475.5329
0.71176 | 1102673.6827 | 66.0131 | 159.3697 | 482.4341
0.70735 | 1073697.1301 | 80.8508 | 195.1911 | 501.0513
0.66967 | 999363.9975 | 67.4468 | 162.8309 | 515.7336
0.7032 | 1044568.7715 | 75.891 | 183.2171 | 481.0245
0.6805 | 1023499.4068 | 83.2553 | 200.9961 | 497.0968
0.68126 | 1023584.8206 | 65.9924 | 159.3197 | 508.7122
0.69344 | 1045646.0691 | 71.8378 | 173.4318 | 473.8317
0.70737 | 1077920.1941 | 77.8518 | 187.9508 | 476.9107
0.68433 | 1029995.7045 | 71.7047 | 173.1105 | 481.0343
0.68269 | 1009580.357 | 64.9529 | 156.8101 | 478.442
0.71442 | 1099011.2359 | 59.0846 | 142.6428 | 502.4603
0.70361 | 1077125.6511 | 72.9311 | 176.0712 | 481.2799
0.7049 | 1090329.5221 | 79.0246 | 190.7822 | 492.8433
0.6968 | 1073671.2065 | 76.991 | 185.8728 | 494.2324
0.68539 | 1019511.9386 | 63.3824 | 153.0186 | 512.7304
0.66039 | 970194.4922 | 69.9904 | 168.9719 | 485.1833
0.68761 | 1034370.4681 | 69.8507 | 168.6345 | 493.7411
0.6818 | 1025453.3007 | 67.9757 | 164.1077 | 478.5085
0.70254 | 1070479.1192 | 79.5204 | 191.9792 | 487.8767
0.6865 | 1037770.2863 | 70.1934 | 169.4618 | 492.9381
0.68096 | 1019970.8634 | 73.1201 | 176.5275 | 477.9017
0.69918 | 1058205.6297 | 85.3162 | 205.9716 | 514.3891
0.70489 | 1084436.2558 | 75.1027 | 181.3139 | 493.3172
0.69895 | 1061609.0558 | 75.892 | 183.2195 | 481.8167
0.68642 | 1032056.5383 | 62.0316 | 149.7576 | 495.4695
0.69237 | 1038225.6502 | 65.8625 | 159.0062 | 487.97
0.68687 | 1030178.0367 | 61.878 | 149.3868 | 480.8136
0.68771 | 1046319.8732 | 79.4023 | 191.6941 | 497.761
0.68431 | 1021373.284 | 79.5008 | 191.9319 | 487.7579
0.68735 | 1049655.4647 | 67.074 | 161.9309 | 494.4609
0.69394 | 1051244.1914 | 66.5371 | 160.6349 | 493.9896
0.67994 | 1017888.4312 | 69.5136 | 167.8207 | 509.6328
0.69708 | 1051542.2945 | 75.3742 | 181.9695 | 473.2887
0.68804 | 1042362.1516 | 75.697 | 182.7488 | 504.8181
0.67555 | 991565.8709 | 74.9883 | 181.0378 | 500.3335
0.66634 | 979757.4704 | 81.3913 | 196.496 | 485.6213
0.67898 | 1006653.6891 | 67.0068 | 161.7687 | 466.5032
0.71643 | 1094721.0135 | 78.0657 | 188.4672 | 479.2642
0.69291 | 1053325.5247 | 75.9564 | 183.375 | 487.3245
0.69001 | 1038604.3553 | 69.8926 | 168.7358 | 486.0112
0.6743 | 1001946.1683 | 82.2519 | 198.5736 | 481.4088
0.67723 | 1038465.4951 | 63.2639 | 152.7326 | 495.9
0.69181 | 1038841.9548 | 76.7171 | 185.2113 | 504.0258
0.68103 | 1025480.181 | 78.5846 | 189.7201 | 467.5007
0.68512 | 1030249.5683 | 83.2399 | 200.9589 | 515.6333
0.66567 | 993037.0511 | 72.8252 | 175.8156 | 492.9429
0.69319 | 1042792.8712 | 69.8523 | 168.6385 | 497.576
0.68558 | 1023794.62 | 79.5412 | 192.0294 | 471.6771
0.67633 | 1010867.2079 | 78.0031 | 188.3162 | 494.145
0.64954 | 954421.5661 | 71.7312 | 173.1744 | 497.4286
0.70647 | 1077397.4647 | 79.5898 | 192.1467 | 520.3294
0.72326 | 1103707.2 | 68.969 | 166.5059 | 484.4521
0.6915 | 1033336.516 | 74.6665 | 180.2608 | 510.7458
0.66407 | 987260.9154 | 80.7567 | 194.9638 | 509.6239
0.69447 | 1049030.4131 | 70.1406 | 169.3344 | 510.4486
0.66466 | 993800.9546 | 72.9918 | 176.2178 | 499.7908
0.70593 | 1070692.8929 | 87.0111 | 210.0633 | 466.9567
0.69275 | 1052028.8631 | 80.8893 | 195.284 | 528.9703
0.69568 | 1059631.5272 | 84.3141 | 203.5523 | 524.2677
0.66683 | 970938.8291 | 69.3667 | 167.466 | 466.158
0.69993 | 1052462.9326 | 93.4052 | 225.5002 | 487.6607
0.70237 | 1069245.2491 | 64.8675 | 156.604 | 493.7478
0.6985 | 1067445.662 | 74.6023 | 180.1059 | 507.2734
0.68681 | 1035821.0455 | 73.6249 | 177.7462 | 505.2523
0.69465 | 1050565.4297 | 80.5559 | 194.4791 | 506.3432
0.69536 | 1052033.2232 | 83.7474 | 202.1841 | 498.0487
0.68929 | 1020956.606 | 77.3004 | 186.6196 | 487.5226
0.68816 | 1028117.4872 | 74.6499 | 180.2208 | 506.4378
0.69892 | 1048240.9889 | 71.4517 | 172.4998 | 500.3382
0.69399 | 1053672.6329 | 74.1184 | 178.9377 | 501.0153
0.68669 | 1021491.4128 | 72.6181 | 175.3156 | 493.8352
0.6809 | 1028991.6759 | 55.7663 | 134.6319 | 492.7455
0.69667 | 1057750.7146 | 74.1186 | 178.9382 | 519.0615
0.72188 | 1112834.6432 | 83.7381 | 202.1616 | 482.898
0.68332 | 1033323.2028 | 77.258 | 186.5173 | 493.1641
0.69063 | 1068468.9207 | 82.405 | 198.9433 | 522.0669
0.68893 | 1026203.9137 | 70.9635 | 171.321 | 489.9761
0.67543 | 993207.0934 | 71.7849 | 173.3041 | 487.931
0.68777 | 1026566.5501 | 57.3285 | 138.4032 | 470.8832
0.7062 | 1091492.0582 | 71.7677 | 173.2626 | 478.4385
0.68854 | 1053901.8666 | 74.7216 | 180.3938 | 478.3683
0.70859 | 1072460.9228 | 68.4042 | 165.1422 | 496.699
0.68834 | 1064755.5659 | 59.4113 | 143.4315 | 484.2218
0.67667 | 1012932.7107 | 76.4775 | 184.6331 | 505.0354
0.68294 | 1018787.809 | 71.0687 | 171.575 | 498.3333
0.65144 | 950155.6476 | 69.4858 | 167.7536 | 442.2066
0.68194 | 1030142.9072 | 68.2727 | 164.8248 | 501.8206
0.7103 | 1083511.5266 | 74.4676 | 179.7808 | 459.6371
0.6904 | 1051504.8951 | 72.5588 | 175.1725 | 511.6931
0.70831 | 1086573.1946 | 74.8873 | 180.7939 | 495.4917
0.69712 | 1043634.9781 | 74.8032 | 180.5909 | 490.3338
0.69536 | 1059696.1201 | 53.8418 | 129.9856 | 510.2118
0.70671 | 1097326.6575 | 85.9502 | 207.5022 | 463.1612
0.68228 | 1024951.6782 | 59.8769 | 144.5557 | 465.4434
0.6927 | 1037268.1803 | 71.3384 | 172.2261 | 508.51
0.69785 | 1064254.6557 | 66.9829 | 161.711 | 485.8836
0.69837 | 1063873.921 | 74.3663 | 179.5362 | 457.8978
0.70113 | 1073446.1561 | 69.1956 | 167.0529 | 472.2434
0.68697 | 1018023.6247 | 80.497 | 194.337 | 518.9019
0.69308 | 1034120.6272 | 71.0829 | 171.6094 | 484.7538
0.71797 | 1105915.9102 | 76.965 | 185.81 | 499.6656
0.64971 | 948735.5211 | 65.9963 | 159.329 | 495.8316
0.66892 | 994227.7036 | 54.8474 | 132.4133 | 473.9111
0.68439 | 1035434.9331 | 72.0773 | 174.01 | 485.7898
0.6974 | 1064175.2822 | 74.4909 | 179.8369 | 503.5387
0.67159 | 1001665.9071 | 71.0556 | 171.5434 | 502.1226
0.6721 | 998192.3104 | 85.2024 | 205.6967 | 467.8595
0.66708 | 982282.6596 | 75.2302 | 181.6217 | 502.5161
0.65925 | 970177.4809 | 91.4158 | 220.6972 | 505.274
0.69847 | 1050814.7282 | 67.9946 | 164.1536 | 490.3546
0.70386 | 1066481.1594 | 80.7056 | 194.8406 | 453.9875
0.69457 | 1040191.6917 | 73.1086 | 176.4999 | 501.7379
0.67162 | 996340.6572 | 74.1243 | 178.9518 | 495.6606
0.69204 | 1034923.8729 | 90.1225 | 217.575 | 479.1674
0.70692 | 1061802.8654 | 76.7366 | 185.2584 | 501.0408
0.67137 | 981828.7016 | 80.9518 | 195.435 | 505.4773
0.70469 | 1092127.0075 | 80.8817 | 195.2657 | 477.2705
0.68402 | 1027163.1174 | 63.9866 | 154.4772 | 489.1037
0.68236 | 1030898.2418 | 77.1851 | 186.3414 | 498.836
0.69268 | 1032933.3509 | 72.7965 | 175.7463 | 460.8239
0.67841 | 1016042.6327 | 84.4247 | 203.8192 | 478.1875
0.677 | 997086.2739 | 78.448 | 189.3902 | 484.1578
0.69889 | 1067188.0684 | 62.9985 | 152.0919 | 464.3958
0.70982 | 1067741.1656 | 64.0964 | 154.7424 | 510.1856
0.6845 | 1035716.5055 | 60.3094 | 145.5998 | 500.8948
0.69001 | 1046254.7991 | 74.0211 | 178.7027 | 483.1453
0.68411 | 1026340.4073 | 61.897 | 149.4325 | 481.2873
0.68409 | 1039010.4471 | 79.265 | 191.3627 | 465.4217
0.70114 | 1076655.7821 | 74.9627 | 180.976 | 475.7848
0.70872 | 1080374.6089 | 81.1963 | 196.0251 | 525.0707
0.67656 | 996365.3713 | 66.5591 | 160.688 | 476.2103
0.68651 | 1054897.4344 | 75.0133 | 181.098 | 504.4666
0.6776 | 1003250.3939 | 63.0893 | 152.3109 | 482.0187
0.69818 | 1059876.1404 | 84.0318 | 202.8706 | 451.7588
0.6744 | 992990.8905 | 64.6709 | 156.1294 | 480.8291
0.7027 | 1080049.3202 | 63.4122 | 153.0906 | 520.4572
0.71628 | 1098391.9755 | 67.353 | 162.6046 | 482.9594
0.69103 | 1043226.319 | 80.412 | 194.1317 | 492.6379
0.70805 | 1090131.3538 | 63.5204 | 153.3519 | 525.2376
0.68389 | 1019481.7236 | 74.4848 | 179.8223 | 479.68
0.66601 | 984637.3051 | 77.1554 | 186.2697 | 511.1875
0.6993 | 1080366.4006 | 80.6267 | 194.6501 | 494.2602
0.69659 | 1059266.687 | 66.7901 | 161.2455 | 487.4167
0.69481 | 1059119.2908 | 60.7567 | 146.6797 | 472.9731
0.6977 | 1056227.2975 | 72.122 | 174.1179 | 514.89
0.6934 | 1044659.4631 | 66.1721 | 159.7537 | 475.1242
0.6784 | 1000415.1109 | 75.3404 | 181.8878 | 479.4349
0.69001 | 1041860.0027 | 70.1115 | 169.2641 | 502.7114
0.68877 | 1047179.9008 | 72.4307 | 174.8632 | 501.5468
0.70944 | 1095825.6427 | 65.818 | 158.8986 | 508.5057
0.70882 | 1083463.3314 | 76.3755 | 184.3869 | 513.6096
0.70979 | 1087683.5801 | 67.3065 | 162.4923 | 447.7454
0.69324 | 1065021.0837 | 78.4885 | 189.4881 | 510.9536
0.68516 | 1033032.2013 | 60.0955 | 145.0833 | 506.9868
0.70327 | 1081768.0094 | 91.2367 | 220.2649 | 487.9899
0.66488 | 986663.4568 | 75.8148 | 183.0332 | 515.2087
0.70966 | 1086028.2416 | 67.2166 | 162.2751 | 504.1542
0.68985 | 1035079.382 | 55.3684 | 133.671 | 488.4992
0.65654 | 954252.2002 | 67.4898 | 162.9347 | 505.4328
0.69204 | 1040319.8397 | 64.0551 | 154.6427 | 496.389
0.68267 | 1035755.2844 | 66.639 | 160.8808 | 466.7617
0.68876 | 1035765.3429 | 81.6871 | 197.21 | 533.6564
0.70146 | 1081786.9609 | 77.0581 | 186.0346 | 466.7651
0.69228 | 1044893.436 | 66.2937 | 160.0471 | 481.0339
0.72358 | 1125118.1232 | 75.9079 | 183.2579 | 538.1959
0.70055 | 1049757.9807 | 67.7244 | 163.5011 | 495.543
0.67753 | 1004960.2639 | 87.9786 | 212.3992 | 506.912
0.65997 | 979465.6009 | 57.0023 | 137.6157 | 496.3737
0.69599 | 1069404.0572 | 80.1885 | 193.5921 | 496.5187
0.66913 | 984134.6402 | 76.7568 | 185.3073 | 492.4094
0.71052 | 1071428.9242 | 60.1043 | 145.1046 | 498.4327
0.68717 | 1024661.7977 | 65.4941 | 158.1166 | 484.2894
0.68562 | 1024643.0922 | 71.1982 | 171.8876 | 487.4026
0.70232 | 1066506.1102 | 82.4865 | 199.14 | 480.122
0.70976 | 1085323.0521 | 74.6482 | 180.2168 | 500.3594
0.70015 | 1066388.2396 | 77.8132 | 187.8577 | 466.4052
0.69129 | 1033975.7914 | 79.6685 | 192.3368 | 454.7766
0.67498 | 1002130.1989 | 79.0814 | 190.9194 | 509.3311
0.66842 | 989655.1635 | 72.2782 | 174.4951 | 502.9674
0.69016 | 1036669.985 | 73.3273 | 177.0279 | 515.5226
0.68112 | 1019124.5699 | 78.087 | 188.5186 | 495.052
0.67345 | 1014702.9374 | 68.0971 | 164.4008 | 474.2021
0.69823 | 1067677.0617 | 83.0613 | 200.5276 | 484.5444
0.68035 | 1021204.356 | 55.3886 | 133.7198 | 487.8027
0.69699 | 1054532.3253 | 66.453 | 160.4317 | 480.5803
0.69966 | 1059395.1053 | 80.6506 | 194.7077 | 499.9867
0.70188 | 1069870.4064 | 70.7793 | 170.8765 | 446.7812
0.69917 | 1068324.4625 | 65.1081 | 157.1848 | 477.1002
0.68848 | 1038770.5315 | 66.2693 | 159.9883 | 511.5308
0.69338 | 1050121.1825 | 73.9333 | 178.4908 | 499.8174
0.69102 | 1051023.925 | 82.6194 | 199.4608 | 516.5762
0.69489 | 1056895.9465 | 60.6224 | 146.3554 | 489.5207
0.70619 | 1088952.5389 | 59.2259 | 142.9839 | 504.4353
0.70019 | 1057719.1862 | 71.9664 | 173.7422 | 501.3679
0.70934 | 1077730.127 | 87.1625 | 210.4288 | 515.0625
0.69227 | 1061198.9868 | 69.281 | 167.2592 | 489.1936
0.70509 | 1057352.1433 | 86.0294 | 207.6934 | 478.6472
0.68777 | 1024824.4297 | 82.4825 | 199.1304 | 519.122
0.67962 | 1011168.0976 | 67.2439 | 162.3412 | 473.915
0.69022 | 1053043.7571 | 81.2644 | 196.1896 | 505.9865
0.68463 | 1015838.2776 | 75.8052 | 183.0099 | 488.4773
0.7104 | 1074809.017 | 73.8583 | 178.3098 | 490.9489
0.65662 | 951115.0313 | 69.0071 | 166.5978 | 472.1104
0.67506 | 998895.5736 | 78.8169 | 190.2808 | 509.2628
0.70493 | 1078863.7158 | 81.5188 | 196.8038 | 472.5856
0.69206 | 1053889.7345 | 79.4735 | 191.8659 | 488.3051
0.68009 | 1010764.1831 | 94.9987 | 229.3471 | 481.8879
0.71362 | 1088756.1981 | 74.4796 | 179.8097 | 485.1901
0.72318 | 1122578.4413 | 82.6815 | 199.6109 | 493.1966
0.69312 | 1037837.3945 | 81.3348 | 196.3597 | 505.0869
0.70767 | 1085452.9979 | 67.3694 | 162.644 | 494.8303
0.69923 | 1062124.605 | 58.3414 | 140.8486 | 497.7575
0.69419 | 1053900.5615 | 72.3265 | 174.6115 | 507.4824
0.69907 | 1059030.4785 | 82.2891 | 198.6636 | 488.9783
0.71668 | 1108161.4828 | 69.8829 | 168.7121 | 483.8292
0.68901 | 1036781.9435 | 76.8768 | 185.597 | 478.145
0.66311 | 975935.687 | 51.9972 | 125.5323 | 491.8495
0.69053 | 1047098.9067 | 70.4111 | 169.9875 | 509.8318
0.66018 | 969573.2883 | 73.3832 | 177.1627 | 499.2943
0.72025 | 1122293.1539 | 66.3959 | 160.2939 | 523.9286
0.70764 | 1066522.537 | 83.4274 | 201.4116 | 492.2894
0.66244 | 986171.4704 | 63.4344 | 153.1443 | 501.3026
0.70508 | 1070131.6502 | 84.0146 | 202.8293 | 505.969
0.68999 | 1032503.3833 | 65.7251 | 158.6744 | 514.6729
0.69364 | 1043245.4045 | 92.0424 | 222.2101 | 450.4031
0.69934 | 1050817.0347 | 75.0782 | 181.2549 | 497.2465
0.67498 | 1014784.6278 | 62.0273 | 149.7473 | 488.7419
0.69125 | 1043750.5672 | 60.1839 | 145.2968 | 504.6512
0.68101 | 1018555.0257 | 61.7068 | 148.9734 | 515.2822
0.66958 | 978388.0725 | 73.0107 | 176.2634 | 483.1991
0.67814 | 1010359.0195 | 79.359 | 191.5897 | 487.6229
0.69917 | 1086266.0459 | 74.0168 | 178.6923 | 490.1208
0.65778 | 959651.2121 | 69.5044 | 167.7985 | 500.9823
0.68487 | 1023706.3053 | 74.6866 | 180.3093 | 476.7317
0.6973 | 1048334.7736 | 73.3933 | 177.187 | 491.2548
0.70258 | 1091015.5243 | 79.3355 | 191.533 | 488.5029
0.69015 | 1041817.6315 | 73.6927 | 177.9099 | 485.1008
0.70062 | 1069576.2589 | 77.9938 | 188.2937 | 464.7338
0.71331 | 1092227.5583 | 96.505 | 232.9838 | 506.3654
0.70301 | 1086599.2606 | 71.7709 | 173.2703 | 463.3375
0.71664 | 1096592.5522 | 54.062 | 130.5173 | 478.3855
0.72155 | 1112069.6958 | 65.672 | 158.5462 | 471.5624
0.69343 | 1059064.975 | 74.5575 | 179.9978 | 502.0214
0.69401 | 1047074.7792 | 70.0566 | 169.1316 | 469.0546
0.6868 | 1028775.4382 | 84.4384 | 203.8523 | 492.976
0.65847 | 957955.4333 | 72.6916 | 175.4931 | 472.8238
0.6817 | 1028034.405 | 73.6189 | 177.7318 | 520.3287
0.69799 | 1062670.9148 | 81.0944 | 195.7791 | 481.4408
0.696 | 1058712.718 | 68.7486 | 165.9737 | 484.7576
0.68873 | 1036450.4264 | 80.1565 | 193.515 | 496.8542
0.70075 | 1086501.6838 | 52.2296 | 126.0933 | 485.8313
0.68602 | 1016057.7057 | 74.0046 | 178.663 | 485.2951
0.70009 | 1051098.2491 | 84.6329 | 204.322 | 456.783
0.67924 | 1013953.1267 | 79.5119 | 191.9587 | 471.0015
0.69205 | 1038046.4144 | 81.4768 | 196.7025 | 480.1581
0.69195 | 1053458.5053 | 72.5304 | 175.1038 | 494.2974
0.67284 | 1006121.2056 | 79.6491 | 192.2899 | 478.6545
0.71231 | 1095580.3728 | 75.0843 | 181.2694 | 465.8117
0.68827 | 1029653.0069 | 74.7497 | 180.4616 | 482.3807
0.71705 | 1096566.7449 | 60.954 | 147.1559 | 491.1162
0.67236 | 986279.7663 | 64.4289 | 155.5452 | 493.3589
0.68952 | 1038271.3721 | 73.5085 | 177.4652 | 481.0304
0.6928 | 1044907.3735 | 60.1159 | 145.1327 | 511.5776
0.67871 | 1013173.2653 | 77.9061 | 188.082 | 474.6479
0.69839 | 1062586.6546 | 82.6458 | 199.5246 | 490.0615
0.70646 | 1087268.8093 | 86.757 | 209.45 | 492.8189
0.7011 | 1062808.5841 | 62.5428 | 150.9917 | 494.7441
0.68375 | 1021516.4849 | 84.2175 | 203.3189 | 538.778
0.69138 | 1055697.8117 | 60.8615 | 146.9326 | 457.5011
0.70205 | 1076255.7049 | 74.8377 | 180.6742 | 485.3647
0.67341 | 1010517.0408 | 81.5122 | 196.788 | 497.5242
0.66173 | 965383.1745 | 58.1484 | 140.3825 | 460.8121
0.69965 | 1073290.0182 | 56.6081 | 136.6641 | 512.9802
0.67892 | 1006000.9824 | 56.9124 | 137.3987 | 484.0774
0.67549 | 1030354.0124 | 74.4635 | 179.7708 | 496.9974
0.68571 | 1039021.7257 | 80.7097 | 194.8504 | 502.5494
0.66131 | 992667.4286 | 63.9614 | 154.4166 | 478.4012
0.68284 | 1034939.571 | 64.5826 | 155.9161 | 503.0336
0.70078 | 1068352.5744 | 76.1748 | 183.9022 | 485.849
0.70136 | 1051207.4822 | 67.7393 | 163.5371 | 490.0368
0.71126 | 1099112.399 | 80.6468 | 194.6986 | 468.6554
0.68639 | 1022107.4259 | 85.4275 | 206.2403 | 493.847
0.67135 | 984567.0378 | 76.5363 | 184.775 | 460.833
0.68732 | 1030321.5014 | 81.4275 | 196.5833 | 508.4513
0.68574 | 1038878.4521 | 73.9564 | 178.5465 | 482.8761
0.6773 | 1015927.3166 | 79.8957 | 192.8852 | 487.4678
0.68523 | 1017057.102 | 65.1921 | 157.3875 | 482.114
0.67699 | 998789.9961 | 58.8044 | 141.9663 | 479.2915
0.69222 | 1043150.4868 | 76.9902 | 185.8709 | 485.0509
0.68622 | 1026934.4424 | 59.6479 | 144.0028 | 498.207
0.68743 | 1050044.2247 | 84.5544 | 204.1323 | 465.2725
0.70808 | 1086643.4527 | 78.403 | 189.2815 | 506.39
0.70355 | 1081834.9615 | 88.214 | 212.9675 | 493.9732
0.69986 | 1066171.3176 | 70.7379 | 170.7764 | 441.3376
0.70479 | 1091169.1423 | 75.5798 | 182.4658 | 477.5496
0.68834 | 1038137.7852 | 78.7087 | 190.0197 | 502.4693
0.68172 | 1008631.6411 | 66.8685 | 161.4349 | 517.5824
0.69791 | 1062515.4295 | 74.3209 | 179.4264 | 502.8015
0.71963 | 1099380.1656 | 78.0498 | 188.429 | 491.3538
0.70457 | 1066483.0496 | 89.7814 | 216.7516 | 524.0405
0.69309 | 1053778.5441 | 79.3472 | 191.5612 | 472.8207
0.67437 | 1018925.6784 | 64.3812 | 155.43 | 470.7339
0.68963 | 1041516.662 | 73.7266 | 177.9917 | 479.4059
0.6892 | 1042018.7261 | 75.5457 | 182.3833 | 506.8572
0.68164 | 1028805.8976 | 64.3121 | 155.263 | 491.0068
0.67561 | 1009979.9373 | 71.2186 | 171.9369 | 500.9653
0.6689 | 988878.9464 | 66.8158 | 161.3075 | 452.0324
0.69961 | 1060786.0367 | 64.6079 | 155.9774 | 493.3638
0.69428 | 1041275.5357 | 73.2161 | 176.7592 | 484.3209
0.70161 | 1059055.5917 | 66.2988 | 160.0595 | 497.6432
0.68397 | 1029257.4938 | 75.1902 | 181.5252 | 487.1968
0.71681 | 1102332.3979 | 67.5451 | 163.0684 | 502.9392
0.66302 | 979903.0162 | 77.1315 | 186.2119 | 513.5775
0.68528 | 1022892.0075 | 65.5612 | 158.2786 | 478.4104
0.68579 | 1030292.0341 | 83.1623 | 200.7715 | 481.2479
0.70237 | 1067293.3937 | 58.3633 | 140.9015 | 500.5212
0.70059 | 1053106.5286 | 73.3811 | 177.1576 | 500.6806
0.70906 | 1087154.3566 | 71.1946 | 171.879 | 473.9201
0.70255 | 1069639.403 | 66.8716 | 161.4423 | 511.3963
0.71023 | 1100208.5561 | 75.6227 | 182.5694 | 506.2136
0.65853 | 972471.8806 | 80.2152 | 193.6566 | 495.2658
0.67242 | 1003055.7318 | 82.573 | 199.3489 | 481.4014
0.6858 | 1031299.071 | 69.8268 | 168.5767 | 508.7281
0.70067 | 1074638.5513 | 65.3529 | 157.7759 | 494.2543
0.68227 | 1017343.6302 | 77.8183 | 187.8701 | 488.471
0.68437 | 1020355.4183 | 74.4972 | 179.8522 | 506.487
0.69641 | 1064635.2927 | 67.1459 | 162.1046 | 502.7005
0.6844 | 1024795.1097 | 78.1445 | 188.6576 | 518.8691
0.69413 | 1053016.9395 | 74.0577 | 178.7912 | 511.9036
0.71807 | 1104519.5878 | 78.5842 | 189.7189 | 495.5078
0.70761 | 1085580.4472 | 71.2369 | 171.9811 | 484.6362
0.6734 | 1004104.5594 | 71.0369 | 171.4982 | 506.2774
0.69893 | 1056553.254 | 66.3956 | 160.2931 | 474.1889
0.69469 | 1072279.1012 | 80.5549 | 194.4767 | 504.7471
0.68824 | 1038460.386 | 64.2938 | 155.2189 | 476.6502
0.72334 | 1113655.3799 | 80.3496 | 193.981 | 490.6321
0.69176 | 1040608.6387 | 80.1253 | 193.4397 | 472.4237
0.68755 | 1035990.4025 | 78.6361 | 189.8444 | 480.524
0.68661 | 1014880.5843 | 54.5443 | 131.6815 | 487.748
0.70564 | 1079464.5339 | 69.1383 | 166.9147 | 466.8033
0.69973 | 1062423.4088 | 72.3343 | 174.6305 | 510.088
0.69031 | 1034635.7557 | 74.8891 | 180.7984 | 504.9195
0.68555 | 1035003.2097 | 85.1053 | 205.4623 | 496.5702
0.68682 | 1039889.6757 | 70.3041 | 169.7291 | 503.6753
0.72249 | 1117598.8797 | 84.0251 | 202.8545 | 510.2068
0.67978 | 1007441.6746 | 65.0932 | 157.1489 | 498.7338
0.70422 | 1075583.1212 | 70.9209 | 171.2182 | 485.6592
0.69149 | 1039576.174 | 59.9224 | 144.6654 | 495.2549
0.6637 | 978474.876 | 76.9403 | 185.7504 | 516.312
0.7066 | 1068633.4946 | 64.4428 | 155.5788 | 500.5931
0.68669 | 1023913.1533 | 67.3909 | 162.696 | 467.4054
0.65918 | 968693.1535 | 76.6935 | 185.1544 | 466.2359
0.68128 | 1021544.529 | 73.1991 | 176.7182 | 468.8667
0.68085 | 995965.7507 | 53.98 | 130.3193 | 487.0165
0.7038 | 1086039.0179 | 49.3137 | 119.0537 | 495.3574
0.67247 | 986147.2905 | 71.5885 | 172.8299 | 484.6072
0.71514 | 1106399.6043 | 71.8392 | 173.4352 | 485.5208
0.68259 | 1023855.3676 | 65.8812 | 159.0513 | 471.324
0.6958 | 1037705.7882 | 74.5192 | 179.9053 | 490.2749
0.69564 | 1052952.6612 | 68.4098 | 165.156 | 523.3054
0.70291 | 1054725.9159 | 74.9753 | 181.0065 | 496.7012
0.67923 | 1022236.7883 | 69.8829 | 168.7123 | 497.3416
0.70147 | 1068022.2763 | 66.8739 | 161.4479 | 470.6222
0.6981 | 1056602.318 | 64.4925 | 155.6986 | 500.0321
0.68146 | 1024515.6167 | 67.9479 | 164.0407 | 503.5697
0.67201 | 1005991.3607 | 57.4966 | 138.8091 | 474.5272
0.6842 | 1053663.6319 | 71.5083 | 172.6363 | 470.6449
0.67022 | 985902.882 | 81.4271 | 196.5825 | 512.3715
0.69356 | 1030888.9327 | 79.1718 | 191.1375 | 457.5748
0.68469 | 1043644.5423 | 70.7775 | 170.8721 | 484.7219
0.65229 | 947152.9344 | 72.5529 | 175.1582 | 488.2061
0.6834 | 1030205.6852 | 75.3226 | 181.8448 | 504.9267
0.66529 | 970983.9932 | 77.7389 | 187.6784 | 473.1039
0.65448 | 950712.6164 | 65.6525 | 158.4992 | 460.6498
0.70719 | 1081253.2712 | 69.1013 | 166.8253 | 511.597
0.69995 | 1070143.7763 | 70.9922 | 171.3903 | 487.0764
0.71281 | 1100161.6598 | 59.2801 | 143.1148 | 510.2285
0.67644 | 1021057.429 | 70.09 | 169.2122 | 490.1237
0.67639 | 1020332.5473 | 81.9417 | 197.8247 | 488.7457
0.69714 | 1062764.0506 | 84.3861 | 203.7261 | 472.5547
0.72193 | 1121010.7974 | 64.0641 | 154.6644 | 511.8683
0.68304 | 1028087.8472 | 71.7071 | 173.1163 | 506.9468
0.71322 | 1095840.5501 | 71.345 | 172.242 | 500.9453
0.69447 | 1048040.6889 | 73.6715 | 177.8588 | 487.8211
0.68029 | 1014733.6713 | 72.8165 | 175.7947 | 490.0077
0.68973 | 1035620.2738 | 55.4067 | 133.7637 | 497.1558
0.67191 | 1006538.0646 | 81.3719 | 196.4492 | 481.2787
0.68607 | 1058534.939 | 90.0271 | 217.3446 | 500.7308
0.69767 | 1070131.5515 | 78.3638 | 189.1869 | 521.9766
0.70095 | 1076176.0734 | 78.4378 | 189.3656 | 487.5123
0.68291 | 1013783.7728 | 72.4697 | 174.9573 | 491.5847
0.68634 | 1049698.349 | 71.0169 | 171.4499 | 487.219
0.70007 | 1070267.265 | 65.143 | 157.269 | 488.7345
0.68519 | 1028553.321 | 75.005 | 181.0781 | 492.9266
0.71195 | 1104934.9548 | 61.859 | 149.3409 | 492.8769
0.6865 | 1042738.3589 | 80.9052 | 195.3224 | 518.1621
0.68634 | 1032417.269 | 71.6631 | 173.01 | 488.2613
0.68731 | 1024898.4555 | 71.0399 | 171.5055 | 452.3863
0.66934 | 986314.8106 | 68.4235 | 165.1888 | 485.0164
0.69119 | 1056430.083 | 89.1828 | 215.3064 | 513.9143
0.68696 | 1031116.6263 | 65.9264 | 159.1604 | 483.5207
0.68225 | 1027526.3803 | 64.1749 | 154.9319 | 493.9785
0.68881 | 1046840.4594 | 66.7515 | 161.1525 | 480.1785
0.70697 | 1086443.7092 | 78.5121 | 189.5449 | 501.1727
0.67694 | 1014248.0718 | 91.6914 | 221.3626 | 478.0923
0.72214 | 1121422.0014 | 70.1847 | 169.4409 | 474.2742
0.70323 | 1087101.6465 | 85.5712 | 206.5871 | 477.8756
0.69191 | 1057125.055 | 69.0511 | 166.7041 | 467.4265
0.70584 | 1084832.3188 | 78.5217 | 189.5682 | 499.5352
0.68193 | 1047113.4353 | 58.9763 | 142.3815 | 470.8858
0.67044 | 1002654.5174 | 73.0065 | 176.2533 | 488.423
0.66123 | 972424.5917 | 75.1333 | 181.3879 | 484.5313
0.69174 | 1043273.309 | 73.6259 | 177.7485 | 493.8986
0.69294 | 1021972.7156 | 66.656 | 160.9219 | 492.5217
0.68908 | 1040827.5831 | 80.4544 | 194.2341 | 494.1228
0.69625 | 1041887.3294 | 87.3164 | 210.8004 | 454.5348
0.69617 | 1041669.5103 | 71.2842 | 172.0952 | 475.9642
0.66611 | 979470.4165 | 73.4725 | 177.3782 | 497.8106
0.71977 | 1114497.9631 | 82.4753 | 199.1129 | 491.972
0.67687 | 1009906.1387 | 67.4332 | 162.798 | 490.7428
0.70457 | 1071021.5654 | 82.5719 | 199.3461 | 505.8231
0.69501 | 1049286.1941 | 68.2342 | 164.7318 | 507.8751
0.70776 | 1085167.0628 | 80.0173 | 193.1789 | 497.6766
0.7059 | 1090587.4395 | 65.8314 | 158.9311 | 475.4295
0.6777 | 1001062.5669 | 66.7602 | 161.1733 | 483.5303
0.70103 | 1064606.5136 | 74.1351 | 178.978 | 500.5656
0.68947 | 1042089.5799 | 73.7846 | 178.1317 | 490.0921
0.68209 | 1025358.9447 | 81.6569 | 197.1373 | 498.0604
0.66822 | 987508.1598 | 91.8007 | 221.6264 | 493.5906
0.71283 | 1082641.869 | 62.6615 | 151.2783 | 478.0044
0.66589 | 985874.7534 | 66.9137 | 161.5438 | 472.9597
0.68562 | 1034269.2365 | 70.0472 | 169.1088 | 508.4748
0.6863 | 1035836.7769 | 73.68 | 177.8794 | 466.4355
0.692 | 1056173.3112 | 74.5084 | 179.8792 | 484.1734
0.69968 | 1056519.609 | 82.0229 | 198.0209 | 504.8695
0.71446 | 1100562.8843 | 80.3913 | 194.0817 | 529.8924
0.67352 | 995801.8084 | 54.5589 | 131.7169 | 515.8893
0.70845 | 1090458.1654 | 73.329 | 177.0319 | 491.6793
0.69452 | 1067398.6566 | 76.7929 | 185.3944 | 464.9143
0.67839 | 1000289.4008 | 70.246 | 169.5888 | 460.1065
0.69431 | 1044500.7321 | 64.1023 | 154.7567 | 493.4978
0.70631 | 1081588.709 | 72.9533 | 176.1249 | 480.2738
0.67864 | 1011891.7547 | 71.0353 | 171.4945 | 513.0344
0.70327 | 1052847.2053 | 65.7544 | 158.7452 | 525.9707
0.69754 | 1065326.8723 | 79.705 | 192.4249 | 499.6702
0.70506 | 1091779.3102 | 71.9027 | 173.5885 | 496.0177
0.6974 | 1061455.9286 | 76.0498 | 183.6006 | 474.4163
0.69868 | 1051359.0162 | 73.4596 | 177.3472 | 509.4401
0.68529 | 1023541.0464 | 81.1317 | 195.8692 | 492.8228
0.70829 | 1063684.8475 | 63.9911 | 154.4883 | 494.8866
0.67177 | 995753.1494 | 63.1607 | 152.4835 | 454.208
0.67276 | 1003849.9206 | 67.4377 | 162.8089 | 494.2311
0.67601 | 1000948.9437 | 70.54 | 170.2986 | 491.6515
0.70699 | 1091510.6132 | 69.7767 | 168.4558 | 488.1883
0.68923 | 1034604.4895 | 78.3395 | 189.1283 | 520.2767
0.6786 | 1015433.375 | 67.6292 | 163.2713 | 490.3077
0.68371 | 1015670.5977 | 61.2886 | 147.9637 | 464.5606
0.67624 | 1006672.2236 | 70.9196 | 171.2151 | 471.0401
0.67211 | 1005438.8976 | 78.0712 | 188.4805 | 485.2541
0.68053 | 1016318.3724 | 81.1219 | 195.8455 | 487.44
0.67684 | 1018564.8588 | 78.537 | 189.6052 | 539.4603
0.68076 | 1019824.2628 | 74.8525 | 180.71 | 471.3915
0.68133 | 1014415.6111 | 77.7022 | 187.5896 | 513.5941
0.67351 | 984625.9661 | 74.6895 | 180.3163 | 500.358
0.6904 | 1038096.5055 | 67.9564 | 164.0613 | 488.7467
0.69789 | 1063792.0241 | 74.412 | 179.6466 | 481.3781
0.6995 | 1086168.8052 | 81.0367 | 195.64 | 513.9745
0.71164 | 1086197.8048 | 70.874 | 171.1049 | 478.8276
0.70394 | 1078091.3378 | 78.2439 | 188.8976 | 483.9747
0.67428 | 1006729.4269 | 72.8867 | 175.9641 | 487.9873
0.67174 | 988070.3648 | 71.8061 | 173.3552 | 501.7917
0.68145 | 1016782.4248 | 61.5598 | 148.6186 | 484.6981
0.68348 | 1035187.4257 | 77.379 | 186.8095 | 491.2615
0.70804 | 1106635.6733 | 84.8993 | 204.9651 | 500.5364
0.69631 | 1041322.752 | 63.7235 | 153.8421 | 501.3162
0.6992 | 1065123.3829 | 70.7352 | 170.7699 | 460.5451
0.70534 | 1059540.967 | 59.8737 | 144.5479 | 491.1781
0.69042 | 1035120.0675 | 57.5842 | 139.0205 | 466.2163
0.68367 | 1015139.6313 | 89.4812 | 216.0267 | 456.6003
0.65749 | 965354.0132 | 71.675 | 173.0387 | 478.0435
0.68526 | 1035046.1992 | 70.1357 | 169.3226 | 493.2353
0.71711 | 1104708.7642 | 69.6237 | 168.0865 | 495.7121
0.66615 | 979014.8915 | 87.1717 | 210.451 | 469.4348
0.69193 | 1041342.183 | 68.018 | 164.2101 | 463.9544
0.67026 | 989741.1425 | 77.0645 | 186.0502 | 495.3805
0.70008 | 1091934.614 | 70.4929 | 170.185 | 485.9799
0.7095 | 1089309.4538 | 61.1949 | 147.7375 | 478.709
0.70361 | 1068368.0359 | 77.5747 | 187.282 | 494.4094
0.70626 | 1076656.7006 | 48.7172 | 117.6136 | 501.5361
0.68138 | 1001792.1267 | 75.1132 | 181.3393 | 480.8639
0.66939 | 999801.335 | 67.5041 | 162.9694 | 492.6248
0.68232 | 1022988.3505 | 78.2447 | 188.8995 | 511.502
0.68194 | 1034294.4667 | 79.1543 | 191.0953 | 496.1089
0.71246 | 1105205.4976 | 80.9575 | 195.4488 | 488.2363
0.6979 | 1068926.1997 | 56.5191 | 136.4492 | 469.5977
0.69637 | 1063861.4515 | 78.0509 | 188.4314 | 492.7151
0.71663 | 1091413.2406 | 91.78 | 221.5764 | 486.1633
0.6873 | 1047970.1744 | 68.9121 | 166.3684 | 524.9911
0.68083 | 1023060.5715 | 67.9139 | 163.9586 | 499.314
0.67893 | 1009685.4125 | 79.2138 | 191.2391 | 481.3615
0.68287 | 1017736.1193 | 63.5055 | 153.3158 | 496.936
0.69468 | 1050887.6627 | 77.286 | 186.5849 | 459.2301
0.6976 | 1059422.6932 | 79.2573 | 191.3441 | 492.9641
0.69531 | 1047460.1019 | 57.8608 | 139.6884 | 473.9329
0.69128 | 1045659.5736 | 73.8435 | 178.2739 | 477.0448
0.70409 | 1080636.5078 | 76.6886 | 185.1426 | 495.7017
0.69612 | 1058261.2355 | 70.591 | 170.4218 | 490.2884
0.69612 | 1066609.351 | 83.2294 | 200.9336 | 478.0353
0.67941 | 1004827.74 | 63.0196 | 152.1428 | 471.2303
0.67845 | 1025748.8683 | 70.2037 | 169.4867 | 492.5471
0.70561 | 1071758.1936 | 82.5073 | 199.1902 | 479.8595
0.68317 | 1011687.2318 | 65.0425 | 157.0265 | 493.8867
0.70359 | 1065244.3818 | 80.6437 | 194.6911 | 493.0786
0.6813 | 1012787.2592 | 61.3313 | 148.0668 | 483.3178
0.69585 | 1048414.7012 | 73.5518 | 177.5698 | 485.5912
0.67625 | 998811.3857 | 60.9368 | 147.1144 | 475.7992
0.71388 | 1086907.9069 | 77.7858 | 187.7916 | 467.7795
0.67801 | 1018250.0332 | 79.994 | 193.1226 | 499.766
0.67598 | 994478.6533 | 79.9133 | 192.9277 | 490.1222
0.68618 | 1054699.5125 | 91.1695 | 220.1026 | 500.2929
0.69557 | 1049514.3146 | 77.0862 | 186.1025 | 475.4008
0.68855 | 1042417.911 | 79.0091 | 190.7448 | 476.1655
0.68432 | 1020992.4229 | 72.8343 | 175.8376 | 472.8184
0.69881 | 1060233.1909 | 83.9842 | 202.7559 | 521.7416
0.70354 | 1062973.931 | 77.1904 | 186.3542 | 490.1476
0.66605 | 977986.6245 | 81.2794 | 196.2258 | 503.5119
0.69556 | 1051836.0709 | 68.3457 | 165.0011 | 482.8299
0.68965 | 1044262.596 | 73.9185 | 178.4552 | 492.5233
0.69306 | 1059685.9393 | 79.7961 | 192.6448 | 520.7289
0.68197 | 1019148.1508 | 65.6374 | 158.4628 | 502.178
0.69604 | 1066732.8655 | 67.3139 | 162.5102 | 494.6001
0.69552 | 1042839.5159 | 88.9341 | 214.7059 | 486.499
0.67763 | 1003428.3975 | 70.5461 | 170.3132 | 497.5322
0.68468 | 1004186.151 | 64.4181 | 155.5191 | 504.0904
0.69165 | 1056304.7035 | 70.9158 | 171.2059 | 488.3788
0.688 | 1041675.7223 | 72.131 | 174.1396 | 496.5388
0.70147 | 1073283.7452 | 76.451 | 184.5691 | 487.0404
0.68901 | 1040819.8062 | 85.1533 | 205.5783 | 494.0473
0.65834 | 974481.1312 | 89.2024 | 215.3537 | 465.5687
0.67656 | 1014984.4773 | 76.6455 | 185.0386 | 503.6278
0.70391 | 1071394.9925 | 77.7613 | 187.7323 | 484.407
0.71517 | 1091959.5267 | 70.8982 | 171.1634 | 511.9019
0.70232 | 1071739.2302 | 82.8913 | 200.1174 | 463.8565
0.70718 | 1083982.0407 | 67.3244 | 162.5355 | 470.5663
0.6951 | 1042835.4849 | 61.8972 | 149.433 | 529.7121
0.67237 | 985103.0964 | 85.4888 | 206.3881 | 504.1585
0.68052 | 996451.9221 | 79.5376 | 192.0207 | 506.547
0.69511 | 1042260.1256 | 69.6241 | 168.0875 | 481.7785
0.69382 | 1068204.7064 | 78.6566 | 189.8939 | 465.0993
0.69773 | 1049700.3201 | 71.6725 | 173.0326 | 533.6117
0.68396 | 1022518.9065 | 69.4194 | 167.5932 | 502.2665
0.6902 | 1032757.7144 | 67.6264 | 163.2647 | 464.0356
0.70419 | 1068984.6878 | 64.6843 | 156.1618 | 500.1531
0.67931 | 1012206.1889 | 72.7403 | 175.6106 | 470.8445
0.70408 | 1078087.9772 | 72.9472 | 176.1102 | 488.9638
0.68703 | 1039011.9816 | 73.8272 | 178.2346 | 507.7982
0.69673 | 1057124.7978 | 66.6416 | 160.8871 | 496.7762
0.68669 | 1023151.5298 | 83.9906 | 202.7711 | 470.2973
0.67743 | 1009823.2058 | 74.4828 | 179.8174 | 476.9406
0.70848 | 1076649.9209 | 84.1201 | 203.0838 | 482.1632
0.69268 | 1042410.1563 | 62.0876 | 149.8928 | 484.8395
0.68008 | 1013550.5395 | 74.4669 | 179.779 | 467.1383
0.68789 | 1027645.9135 | 81.1483 | 195.9092 | 499.5096
0.68639 | 1044493.8687 | 70.6601 | 170.5887 | 511.9529
0.69337 | 1044112.8759 | 66.1397 | 159.6753 | 495.6852
0.70197 | 1070947.4578 | 68.9516 | 166.464 | 496.0432
0.67406 | 988868.1371 | 66.303 | 160.0695 | 527.1074
0.69527 | 1037717.3173 | 71.1055 | 171.6638 | 499.6863
0.69476 | 1055469.0015 | 70.5421 | 170.3037 | 472.8898
0.71347 | 1095926.099 | 69.423 | 167.602 | 495.8348
0.68533 | 1030547.1775 | 76.8238 | 185.4691 | 515.0893
0.68775 | 1026089.1231 | 59.4668 | 143.5656 | 479.4625
0.68738 | 1028718.9429 | 69.4925 | 167.7697 | 523.8389
0.68626 | 1009919.6176 | 75.1847 | 181.5119 | 483.034
0.69352 | 1050767.7393 | 66.5469 | 160.6584 | 503.1788
0.67092 | 994385.237 | 65.0983 | 157.1612 | 473.8639
0.68607 | 1034715.116 | 66.2419 | 159.9222 | 521.0885
0.67351 | 996468.6959 | 44.5771 | 107.6186 | 509.8046
0.71424 | 1086350.5382 | 74.2505 | 179.2566 | 498.8993
0.67991 | 1017906.699 | 57.5014 | 138.8207 | 515.0354
0.69499 | 1047834.6892 | 68.1413 | 164.5077 | 496.5839
0.68664 | 1034554.6853 | 93.5937 | 225.9551 | 491.0563
0.68616 | 1020904.3642 | 83.871 | 202.4824 | 498.5638
0.71749 | 1101461.8936 | 67.427 | 162.7831 | 464.4257
0.71293 | 1084962.1256 | 70.6649 | 170.6002 | 474.2507
0.67435 | 981375.18 | 69.331 | 167.3798 | 511.9832
0.6807 | 1026302.6456 | 81.8985 | 197.7204 | 500.2657
0.67533 | 1007033.3338 | 79.427 | 191.7538 | 481.2256
0.6667 | 999514.7905 | 63.8813 | 154.223 | 464.5895
0.69272 | 1052111.1146 | 64.7174 | 156.2416 | 520.4749
0.67451 | 994538.8158 | 70.244 | 169.5839 | 489.3077
0.69494 | 1047329.9813 | 81.8641 | 197.6373 | 481.2639
0.68863 | 1019579.9682 | 80.4254 | 194.1642 | 488.4602
0.66852 | 992558.8351 | 61.2555 | 147.884 | 487.3869
0.68813 | 1014012.2857 | 73.6288 | 177.7555 | 508.199
0.70017 | 1081304.0904 | 65.8049 | 158.8672 | 496.4105
0.6971 | 1042150.1019 | 76.378 | 184.3927 | 497.9258
0.68476 | 1034324.1654 | 83.6122 | 201.8577 | 467.2594
0.68278 | 1021525.1381 | 71.2765 | 172.0767 | 494.3592
0.67599 | 999631.433 | 61.0098 | 147.2908 | 491.2699
0.68187 | 1021943.4926 | 74.8864 | 180.7917 | 451.9242
0.69267 | 1044029.2569 | 71.8061 | 173.3553 | 498.0245
0.69447 | 1052505.4906 | 80.7253 | 194.8882 | 479.9899
0.69178 | 1052510.1082 | 70.6471 | 170.5573 | 480.6271
0.69284 | 1063327.5944 | 83.3662 | 201.2637 | 490.4995
0.6953 | 1063238.5399 | 69.1122 | 166.8516 | 494.7201
0.67597 | 1007580.1943 | 63.5143 | 153.3371 | 463.7414
0.69072 | 1035799.6021 | 69.6326 | 168.108 | 510.2571
0.69638 | 1059150.0006 | 58.4162 | 141.0291 | 491.617
0.68706 | 1044933.064 | 68.7364 | 165.9443 | 476.9542
0.69724 | 1055340.03 | 77.3286 | 186.6878 | 506.6087
0.71153 | 1101970.9803 | 65.9922 | 159.3193 | 499.2242
0.67936 | 1023599.7392 | 64.2019 | 154.9972 | 488.2297
0.6733 | 1006737.6648 | 63.7126 | 153.8159 | 510.821
0.67655 | 1013167.3771 | 61.7845 | 149.161 | 523.5889
0.68426 | 1009377.8628 | 80.1292 | 193.449 | 491.4266
0.70891 | 1076197.9297 | 87.0913 | 210.257 | 479.4071
0.67437 | 994919.191 | 58.8038 | 141.965 | 498.7564
0.69848 | 1062370.6713 | 79.3769 | 191.6329 | 487.3698
0.70625 | 1085777.6836 | 80.2341 | 193.7023 | 509.1799
0.72361 | 1114386.3866 | 97.062 | 234.3283 | 499.121
0.69045 | 1030262.9992 | 69.2873 | 167.2743 | 496.6929
0.67171 | 995970.6497 | 67.0132 | 161.7842 | 502.0679
0.67313 | 1002345.7533 | 70.9405 | 171.2655 | 491.8866
0.69271 | 1038213.5075 | 81.4125 | 196.5471 | 527.96
0.67568 | 1009415.4944 | 77.966 | 188.2265 | 484.1481
0.65901 | 971862.1132 | 71.9922 | 173.8046 | 494.9073
0.6945 | 1055075.7911 | 68.6604 | 165.7608 | 516.2003
0.68065 | 1028589.9036 | 78.8665 | 190.4005 | 484.1353
0.67197 | 994701.5684 | 80.4083 | 194.1229 | 510.2292
0.71228 | 1078834.0731 | 72.1573 | 174.2031 | 485.3932
0.70755 | 1079410.2093 | 77.9401 | 188.1641 | 492.6963
0.66493 | 992693.259 | 71.124 | 171.7085 | 493.1284
0.67309 | 1010312.6563 | 73.2751 | 176.9018 | 506.4414
0.70402 | 1061433.8856 | 73.6005 | 177.6874 | 492.8927
0.7098 | 1070803.5928 | 78.7318 | 190.0754 | 516.9187
0.70693 | 1078708.0816 | 90.5072 | 218.5038 | 505.6656
0.68704 | 1034858.2726 | 71.7587 | 173.2408 | 480.3479
0.69783 | 1066225.4118 | 70.9474 | 171.2823 | 496.9584
0.69915 | 1064080.7153 | 70.6715 | 170.6162 | 500.2231
0.70231 | 1082781.5888 | 65.6943 | 158.6002 | 495.6782
0.68756 | 1041331.3529 | 68.4506 | 165.2543 | 492.6325
0.70644 | 1084227.3767 | 63.3204 | 152.869 | 487.1825
0.68749 | 1015406.4863 | 70.3646 | 169.8751 | 512.2212
0.67705 | 997201.3863 | 64.6911 | 156.178 | 504.0208
0.68183 | 1037508.2868 | 70.6563 | 170.5793 | 484.6372
0.68065 | 1003785.8208 | 78.5524 | 189.6424 | 472.8728
0.68878 | 1049743.1884 | 68.8813 | 166.2941 | 481.6189
0.70228 | 1068362.63 | 76.0129 | 183.5114 | 498.5054
0.69358 | 1055462.6752 | 84.237 | 203.366 | 474.7242
0.70974 | 1097200.3405 | 77.2433 | 186.4818 | 508.3296
0.69853 | 1062321.634 | 66.7962 | 161.2604 | 485.8878
0.69121 | 1049768.6515 | 70.5218 | 170.2546 | 511.1631
0.68601 | 1044744.4843 | 76.0948 | 183.7092 | 508.1762
0.67825 | 1004326.9286 | 72.2866 | 174.5152 | 496.3804
0.67007 | 994420.7684 | 63.2367 | 152.6669 | 486.8991
0.68501 | 1036554.2529 | 72.0894 | 174.0393 | 499.8562
0.7084 | 1073772.034 | 73.9954 | 178.6407 | 496.2757
0.6911 | 1037917.8134 | 67.8764 | 163.8681 | 508.523
0.66344 | 968436.9102 | 72.937 | 176.0855 | 489.9362
0.68865 | 1039412.4452 | 60.1602 | 145.2397 | 483.9554
0.67548 | 995705.7913 | 68.4468 | 165.2452 | 506.415
0.70861 | 1102044.4424 | 82.8386 | 199.9901 | 506.71
0.68636 | 1039576.3106 | 68.4543 | 165.2632 | 513.2217
0.67505 | 1013337.0743 | 75.0057 | 181.0797 | 485.2819
0.69688 | 1051963.9736 | 71.9151 | 173.6185 | 493.6596
0.72111 | 1109866.5883 | 71.4671 | 172.5369 | 496.2488
0.69084 | 1036378.554 | 83.0409 | 200.4784 | 467.2556
0.70572 | 1081157.2601 | 71.1365 | 171.7387 | 503.7824
0.67139 | 994254.7145 | 66.9293 | 161.5817 | 479.5498
0.70211 | 1084100.98 | 85.4708 | 206.3448 | 488.5641
0.69644 | 1071644.25 | 80.3042 | 193.8715 | 506.9953
0.69314 | 1049843.4813 | 77.3809 | 186.814 | 465.8162
0.67197 | 983869.1564 | 58.3591 | 140.8913 | 471.9065
0.70551 | 1076866.9585 | 65.0396 | 157.0194 | 490.1361
0.68432 | 1052020.8603 | 76.8147 | 185.4472 | 486.1812
0.706 | 1068454.6061 | 66.4717 | 160.4769 | 515.6599
0.70228 | 1074204.3302 | 67.183 | 162.1942 | 534.9722
0.68487 | 1021488.8601 | 77.839 | 187.9201 | 504.2586
0.68402 | 1044556.8171 | 69.0172 | 166.6223 | 510.6873
0.68836 | 1039918.3393 | 62.3286 | 150.4744 | 526.3209
0.66574 | 974630.2901 | 56.2285 | 135.7476 | 491.6424
0.70183 | 1066160.3533 | 72.3715 | 174.7203 | 521.5829
0.70234 | 1054257.1726 | 75.0027 | 181.0724 | 478.9713
0.68721 | 1036064.5356 | 81.8939 | 197.7095 | 463.7957
0.69165 | 1047488.6583 | 62.4832 | 150.8478 | 525.587
0.68983 | 1028023.1436 | 65.4551 | 158.0225 | 490.0016
0.68417 | 1031690.893 | 73.6137 | 177.7191 | 518.0797
0.70819 | 1101675.1523 | 72.2335 | 174.387 | 500.1471
0.7105 | 1081462.3829 | 74.305 | 179.388 | 491.2741
0.68544 | 1034642.6295 | 85.0891 | 205.4233 | 510.397
0.67742 | 1014868.1551 | 73.1347 | 176.5629 | 483.4714
0.69674 | 1058955.4389 | 72.408 | 174.8085 | 476.6832
0.69369 | 1051918.5139 | 91.3839 | 220.6203 | 470.5234
0.69021 | 1036657.2103 | 59.8653 | 144.5275 | 506.8645
0.69986 | 1047790.0797 | 84.3946 | 203.7466 | 477.4692
0.67743 | 1012426.3062 | 72.9582 | 176.1366 | 502.6173
0.7204 | 1126432.7782 | 76.3313 | 184.28 | 464.1346
0.69312 | 1039801.4929 | 84.5656 | 204.1595 | 457.101
0.70619 | 1075021.5007 | 74.3136 | 179.4088 | 494.4059
0.70516 | 1071346.2698 | 72.3799 | 174.7406 | 504.3907
0.66539 | 983393.5653 | 66.7953 | 161.2582 | 479.4007
0.66128 | 972018.3169 | 75.0967 | 181.2995 | 512.673
0.68415 | 1018063.9132 | 74.3537 | 179.5057 | 509.3998
0.69233 | 1046017.7282 | 66.3606 | 160.2088 | 515.049
0.68128 | 1018737.4497 | 76.3959 | 184.436 | 493.1452
0.68398 | 1028906.0758 | 76.8536 | 185.5409 | 544.9831
0.69513 | 1049938.621 | 71.9569 | 173.7193 | 489.9304
0.69609 | 1066824.9832 | 51.6273 | 124.6392 | 481.8364
0.69306 | 1051812.9224 | 72.3015 | 174.5513 | 493.9882
0.68895 | 1048546.6954 | 51.2679 | 123.7717 | 482.1009
0.7172 | 1113091.845 | 64.0313 | 154.5853 | 513.7952
0.70638 | 1086214.8692 | 73.2583 | 176.8612 | 496.9961
0.69597 | 1041695.3364 | 61.7524 | 149.0836 | 479.2056
0.68442 | 1029149.7078 | 67.9596 | 164.0689 | 488.3113
0.69328 | 1048184.3944 | 77.2433 | 186.4819 | 475.9288
0.68421 | 1037683.9458 | 70.4229 | 170.016 | 502.0383
0.6913 | 1053997.1214 | 96.093 | 231.9889 | 479.6061
0.68664 | 1044951.0859 | 84.2401 | 203.3736 | 503.2755
0.69492 | 1046882.2652 | 75.2733 | 181.7259 | 471.8322
0.70535 | 1079364.2075 | 89.2899 | 215.5649 | 462.0564
0.69692 | 1072134.7233 | 83.5644 | 201.7422 | 477.4199
0.70399 | 1069455.8588 | 70.5393 | 170.297 | 492.465
0.6784 | 1023163.4904 | 78.7106 | 190.0242 | 465.266
0.67669 | 1017699.6769 | 64.0834 | 154.711 | 464.2636
0.69105 | 1042972.9042 | 77.1067 | 186.1521 | 471.9658
0.65768 | 976304.6539 | 72.4681 | 174.9534 | 502.9543
0.67679 | 1018395.7466 | 55.527 | 134.054 | 473.8304
0.68119 | 1007744.504 | 66.862 | 161.4191 | 513.553
0.68611 | 1045542.7896 | 87.4219 | 211.0551 | 487.9904
0.67993 | 1015273.6427 | 69.3354 | 167.3904 | 500.9523
0.66799 | 989755.4415 | 65.0793 | 157.1153 | 506.0251
0.67824 | 1015159.0567 | 65.7407 | 158.712 | 521.9013
0.68053 | 1023101.9584 | 73.3912 | 177.1821 | 498.939
0.66794 | 999330.3898 | 81.2704 | 196.204 | 483.4208
0.69507 | 1057153.5081 | 79.3748 | 191.6278 | 512.7706
0.67258 | 996478.6308 | 70.5834 | 170.4034 | 488.8818
0.71089 | 1084733.6783 | 67.0195 | 161.7993 | 494.9408
0.70288 | 1068925.089 | 60.6154 | 146.3386 | 477.6227
0.69733 | 1069460.3807 | 73.0645 | 176.3934 | 490.5003
0.69008 | 1053307.9284 | 81.3012 | 196.2785 | 522.5123
0.70673 | 1079704.1954 | 79.7306 | 192.4866 | 525.9151
0.68603 | 1038748.3432 | 72.1351 | 174.1496 | 476.1052
0.67543 | 999760.8631 | 73.9591 | 178.553 | 451.1703
0.68065 | 1002015.5928 | 70.1344 | 169.3194 | 467.3088
0.68139 | 1006977.8469 | 73.842 | 178.2703 | 495.4117
0.68041 | 1027657.5926 | 73.7317 | 178.0041 | 487.1892
0.68199 | 1005705.5043 | 75.3394 | 181.8853 | 523.5597
0.694 | 1057951.751 | 74.4187 | 179.6627 | 504.0347
0.67944 | 1017213.2862 | 84.3949 | 203.7473 | 474.0517
0.71769 | 1123572.2161 | 74.5832 | 180.0598 | 503.6732
0.67034 | 993675.5594 | 54.7004 | 132.0584 | 500.3891
0.68516 | 1033636.6478 | 68.1825 | 164.6071 | 474.2986
0.66896 | 991076.8883 | 72.3233 | 174.604 | 490.4442
0.69135 | 1048019.7301 | 74.0024 | 178.6576 | 504.3937
0.6899 | 1040720.8077 | 73.6413 | 177.7858 | 508.6122
0.71334 | 1105398.6126 | 79.1685 | 191.1298 | 495.5978
0.67079 | 1002643.2996 | 67.555 | 163.0921 | 509.2717
0.67674 | 1007454.7404 | 76.0567 | 183.617 | 505.2894
0.68439 | 1038270.1974 | 77.5061 | 187.1164 | 481.7332
0.67365 | 1014105.9795 | 71.7173 | 173.1408 | 488.057
0.70114 | 1062369.1935 | 72.1853 | 174.2707 | 493.209
0.69707 | 1049606.9528 | 66.1349 | 159.6638 | 480.5091
0.70711 | 1088856.1516 | 73.3608 | 177.1086 | 520.9944
0.68349 | 1027030.0188 | 80.037 | 193.2265 | 488.5631
0.68891 | 1024351.9946 | 82.3079 | 198.7089 | 518.4021
0.69732 | 1056504.2206 | 75.704 | 182.7655 | 478.6007
0.69317 | 1034931.5424 | 67.4789 | 162.9086 | 483.8782
0.70408 | 1072157.851 | 77.7983 | 187.8216 | 464.2286
0.6744 | 1004599.9136 | 77.7405 | 187.6822 | 521.5702
0.67596 | 1004895.6333 | 69.03 | 166.6533 | 520.5368
0.68833 | 1041568.2759 | 80.5986 | 194.5823 | 477.7143
0.69399 | 1044951.0407 | 62.1222 | 149.9761 | 510.6121
0.7126 | 1092182.9913 | 54.1495 | 130.7284 | 481.3078
0.6747 | 1014608.9148 | 64.0775 | 154.6969 | 498.8535
0.70558 | 1063215.5224 | 64.2079 | 155.0117 | 486.9911
0.67673 | 1008582.981 | 80.9565 | 195.4463 | 482.7481
0.71784 | 1080288.1771 | 71.6951 | 173.0873 | 524.3179
0.70466 | 1073373.4322 | 68.8442 | 166.2046 | 467.0715
0.70404 | 1083247.5906 | 73.3008 | 176.9638 | 492.9616
0.70115 | 1069185.2238 | 72.6468 | 175.385 | 476.6437
0.65706 | 960793.9666 | 66.0874 | 159.5492 | 486.0714
0.69162 | 1044143.402 | 64.3291 | 155.3043 | 496.1621
0.67896 | 1018543.8194 | 76.6421 | 185.0303 | 490.9348
0.70508 | 1074959.7903 | 64.9633 | 156.8352 | 484.8899
0.69993 | 1063207.502 | 79.2747 | 191.386 | 471.6437
0.66303 | 979826.2459 | 64.9079 | 156.7015 | 459.6652
0.68738 | 1046307.9743 | 79.44 | 191.7852 | 529.8165
0.69171 | 1029622.0569 | 79.2856 | 191.4124 | 490.0253
0.66731 | 983336.6429 | 62.7897 | 151.5877 | 482.0276
0.7267 | 1117949.273 | 63.47 | 153.2302 | 505.9465
0.67853 | 997741.6891 | 84.3579 | 203.6581 | 507.2228
0.67093 | 994378.1837 | 70.6072 | 170.461 | 509.7599
0.68056 | 1015013.6203 | 74.522 | 179.9121 | 440.1725
0.6816 | 1022972.8876 | 55.3833 | 133.707 | 471.5785
0.69328 | 1038967.5397 | 67.7313 | 163.5177 | 508.7527
0.68983 | 1023301.9731 | 72.4159 | 174.8275 | 493.7574
0.67631 | 1008732.328 | 82.5622 | 199.3227 | 495.4947
0.6906 | 1017677.6716 | 66.6009 | 160.7888 | 492.0487
0.7071 | 1079135.9389 | 79.6466 | 192.2839 | 488.7832
0.68301 | 1014682.2796 | 75.8983 | 183.2348 | 501.5865
0.65576 | 954281.797 | 62.0082 | 149.701 | 478.4719
0.70462 | 1076000.1854 | 69.0166 | 166.6208 | 533.3399
0.70655 | 1071030.9432 | 68.5101 | 165.398 | 466.8857
0.67436 | 1006818.6579 | 71.1274 | 171.7168 | 456.5728
0.67904 | 1025628.7582 | 75.4694 | 182.1992 | 494.2731
0.68595 | 1020772.5488 | 70.8946 | 171.1548 | 500.7274
0.68304 | 1013033.9752 | 67.0308 | 161.8267 | 495.4384
0.69658 | 1057073.4475 | 77.5285 | 187.1704 | 521.8377
0.69686 | 1040920.2751 | 74.6361 | 180.1874 | 494.3979
0.70346 | 1074682.6777 | 74.5787 | 180.0489 | 514.6188
0.69514 | 1064849.6585 | 74.7807 | 180.5367 | 518.8805
0.67492 | 993185.1756 | 64.6335 | 156.039 | 473.3867
0.68449 | 1042272.4623 | 77.7933 | 187.8096 | 508.9522
0.69277 | 1043997.4855 | 71.7941 | 173.3263 | 491.7927
0.67859 | 1009987.9347 | 69.1109 | 166.8484 | 476.4885
0.67385 | 995354.2659 | 62.9409 | 151.9527 | 510.2833
0.69417 | 1059115.2886 | 76.2345 | 184.0463 | 489.4833
0.68804 | 1037094.7357 | 75.315 | 181.8265 | 490.3719
0.71797 | 1108365.5801 | 71.6829 | 173.0578 | 474.8868
0.67994 | 1008599.7902 | 80.5755 | 194.5265 | 495.6645
0.67929 | 1016770.3314 | 81.3068 | 196.2919 | 488.8908
0.68873 | 1050828.5497 | 62.9483 | 151.9706 | 508.0488
0.69527 | 1051594.227 | 77.2723 | 186.5518 | 496.1229
0.69247 | 1047951.2416 | 72.0966 | 174.0565 | 467.2402
0.68319 | 1010193.5556 | 97.1049 | 234.4319 | 503.4411
0.67754 | 1010568.476 | 65.5109 | 158.1574 | 487.5688
0.67545 | 1027122.0376 | 68.7264 | 165.9202 | 507.344
0.70041 | 1055220.1844 | 71.7744 | 173.2787 | 494.2452
0.68606 | 1034464.2292 | 68.1472 | 164.522 | 465.598
0.67916 | 1008948.6251 | 75.8972 | 183.2321 | 485.597
0.70582 | 1065777.3799 | 73.713 | 177.9589 | 480.342
0.71031 | 1088184.8227 | 67.2876 | 162.4466 | 467.4381
0.6856 | 1049710.6133 | 74.7436 | 180.447 | 472.8677
0.71283 | 1102172.1881 | 85.5895 | 206.6313 | 449.0137
0.69933 | 1066623.9787 | 81.2401 | 196.131 | 455.8224
0.69459 | 1062557.8863 | 74.6791 | 180.2912 | 474.0697
0.67583 | 1004660.9454 | 67.1859 | 162.2012 | 513.5701
0.66787 | 993025.7682 | 65.3059 | 157.6625 | 473.0967
0.67068 | 992628.3208 | 80.3945 | 194.0895 | 489.8576
0.67869 | 1005665.3434 | 75.7024 | 182.7618 | 494.5638
0.69581 | 1076937.7827 | 77.3352 | 186.7037 | 492.4968
0.70633 | 1077873.9937 | 82.8927 | 200.1207 | 499.5943
0.71924 | 1106608.9765 | 65.5595 | 158.2747 | 520.6896
0.70569 | 1073941.6023 | 77.749 | 187.7026 | 455.9628
0.68406 | 1031256.7336 | 66.2145 | 159.8558 | 493.3544
0.70622 | 1074081.9768 | 64.8062 | 156.4559 | 487.762
0.67012 | 1007300.304 | 71.6122 | 172.8871 | 494.6477
0.67694 | 993295.178 | 69.9047 | 168.7648 | 487.6269
0.66828 | 1004083.9116 | 84.6967 | 204.4759 | 501.3463
0.71071 | 1104314.9067 | 72.1501 | 174.1858 | 493.7104
0.68752 | 1031655.1565 | 79.926 | 192.9585 | 502.8671
0.68703 | 1037748.2095 | 78.8302 | 190.313 | 507.3464
0.67074 | 1003321.3799 | 76.4213 | 184.4975 | 502.453
0.713 | 1092677.0666 | 86.3887 | 208.5608 | 479.8565
0.68405 | 1041061.2381 | 72.5041 | 175.0404 | 512.5418
0.73073 | 1138568.4555 | 74.2303 | 179.2078 | 495.7326
0.71539 | 1091520.6438 | 65.0137 | 156.9569 | 480.8497
0.67758 | 1005099.5406 | 67.0419 | 161.8535 | 475.5859
0.72359 | 1124833.938 | 84.4408 | 203.8582 | 502.4228
0.66677 | 967200.0884 | 70.8136 | 170.9592 | 479.7333
0.67801 | 1025779.8829 | 70.2518 | 169.6028 | 477.3963
0.68505 | 1046990.8309 | 67.1282 | 162.0618 | 463.9784
0.69493 | 1067445.186 | 74.9858 | 181.0318 | 465.3296
0.67314 | 998769.4233 | 77.4343 | 186.9431 | 496.4384
0.71374 | 1091009.38 | 69.0855 | 166.7872 | 494.6772
0.72153 | 1127305.3247 | 81.5376 | 196.8492 | 461.3899
0.68293 | 1024067.1933 | 74.9251 | 180.8851 | 470.6166
0.68766 | 1048110.4207 | 81.1848 | 195.9973 | 462.6539
0.68715 | 1028567.1044 | 67.6518 | 163.3259 | 449.6189
0.68199 | 1040393.6518 | 76.2244 | 184.0221 | 471.5437
0.68219 | 1017817.9737 | 63.5743 | 153.4819 | 464.6683
0.68907 | 1043070.8085 | 70.2893 | 169.6934 | 485.9436
0.69709 | 1053853.6923 | 78.8973 | 190.4749 | 477.4292
0.68566 | 1044604.1625 | 76.1239 | 183.7794 | 509.4399
0.67992 | 1012007.4306 | 65.3749 | 157.8289 | 469.5567
0.67769 | 1010388.38 | 80.9869 | 195.5197 | 488.0367
0.68424 | 1028983.3837 | 59.9398 | 144.7076 | 495.8507
0.72883 | 1108547.032 | 75.9981 | 183.4756 | 465.3792
0.68897 | 1030279.9056 | 88.6977 | 214.1353 | 495.5541
0.71541 | 1097085.3674 | 74.8612 | 180.7308 | 459.8105
0.67597 | 1009362.2719 | 79.5355 | 192.0158 | 494.0686
0.6778 | 1014801.8718 | 74.8094 | 180.6058 | 452.994
0.70049 | 1075021.6991 | 81.7443 | 197.3481 | 486.8169
0.68194 | 1019745.3568 | 76.3972 | 184.4391 | 504.101
0.68671 | 1040290.5215 | 77.3949 | 186.8478 | 499.0843
0.6972 | 1069730.1111 | 69.3442 | 167.4116 | 470.6808
0.68189 | 1020752.7439 | 78.123 | 188.6057 | 503.5441
0.67221 | 984988.7342 | 66.2248 | 159.8807 | 482.7276
0.69814 | 1054794.3838 | 89.6161 | 216.3523 | 494.9878
0.68148 | 1030856.8982 | 68.3662 | 165.0507 | 473.8717
0.69422 | 1084033.8838 | 70.573 | 170.3784 | 458.5015
0.69804 | 1064719.1249 | 82.8536 | 200.0262 | 520.0416
0.68958 | 1040354.9353 | 68.1857 | 164.6149 | 495.492
0.69787 | 1071223.5055 | 71.2761 | 172.0758 | 506.8298
0.70684 | 1066015.4918 | 77.0499 | 186.015 | 466.4289
0.70146 | 1084165.1501 | 65.8577 | 158.9945 | 485.1922
0.67801 | 992599.9844 | 62.3816 | 150.6025 | 493.2075
0.68396 | 1007349.1579 | 71.899 | 173.5794 | 472.7613
0.69352 | 1038211.0192 | 67.6583 | 163.3415 | 503.0921
0.67589 | 1010999.2884 | 67.6038 | 163.21 | 500.4439
0.67701 | 1005890.0714 | 52.493 | 126.7294 | 468.9298
0.69666 | 1043212.0519 | 69.5332 | 167.8679 | 498.075
0.70158 | 1079520.2611 | 71.4481 | 172.491 | 510.7376
0.68434 | 1027823.3328 | 57.6857 | 139.2655 | 478.3304
0.6759 | 993929.9814 | 73.4765 | 177.3879 | 501.3931
0.68303 | 1012511.9074 | 74.0696 | 178.8197 | 508.4372
0.66793 | 991208.8534 | 59.1498 | 142.8001 | 491.4896
0.67834 | 1005399.7264 | 68.8227 | 166.1527 | 497.7809
0.65459 | 971460.443 | 80.6144 | 194.6204 | 514.8503
0.6805 | 1028335.8325 | 75.5061 | 182.2879 | 516.306
0.69756 | 1047689.6364 | 71.5528 | 172.7438 | 453.9168
0.68126 | 1010581.9327 | 70.6701 | 170.6128 | 504.4435
0.66883 | 998042.2349 | 65.5414 | 158.2308 | 507.6944
0.68597 | 1032587.1371 | 80.1793 | 193.57 | 502.8958
0.67531 | 1006964.9916 | 82.5401 | 199.2693 | 498.2395
0.69971 | 1050677.2076 | 83.1197 | 200.6686 | 498.8431
0.67547 | 1008221.3159 | 65.5586 | 158.2725 | 470.365
0.65256 | 948029.0057 | 67.4585 | 162.8593 | 456.2869
0.68463 | 1030638.2762 | 64.1453 | 154.8605 | 489.0416
0.6834 | 1015404.9629 | 68.9929 | 166.5637 | 468.1825
0.68851 | 1045151.4424 | 70.4871 | 170.171 | 517.4537
0.69218 | 1054075.7841 | 76.09 | 183.6975 | 535.5148
0.68727 | 1037973.506 | 66.3345 | 160.1458 | 487.7632
0.67448 | 1000269.3885 | 83.5994 | 201.8269 | 492.9649
0.7111 | 1091676.0901 | 63.3383 | 152.9123 | 479.2206
0.70746 | 1064069.3952 | 53.1391 | 128.289 | 481.5557
0.69068 | 1044579.2934 | 82.1529 | 198.3346 | 527.2009
0.67877 | 1002424.9004 | 75.4501 | 182.1526 | 488.9308
0.68536 | 1040106.491 | 70.4687 | 170.1265 | 508.7482
0.6963 | 1062025.0271 | 76.2105 | 183.9884 | 488.1469
0.67286 | 981423.4429 | 64.064 | 154.6642 | 481.0377
0.68271 | 1033773.7171 | 66.5074 | 160.563 | 480.3158
0.69076 | 1056253.3762 | 74.6476 | 180.2152 | 473.1447
0.67606 | 1022796.4181 | 70.2488 | 169.5957 | 486.5767
0.67654 | 1024786.8809 | 75.6857 | 182.7214 | 484.8985
0.68939 | 1026446.1062 | 65.3964 | 157.8808 | 507.3461
0.69573 | 1065296.0235 | 78.2276 | 188.858 | 483.1734
0.67679 | 1005206.0046 | 77.513 | 187.133 | 488.2886
0.66596 | 978233.4866 | 64.5179 | 155.7599 | 498.2768
0.68212 | 1025382.3913 | 59.9162 | 144.6505 | 483.1691
0.7089 | 1087643.5984 | 73.0935 | 176.4632 | 509.223
0.65975 | 958131.6919 | 75.7908 | 182.9752 | 462.5811
0.68427 | 1015813.231 | 75.8431 | 183.1013 | 518.411
0.70695 | 1069771.2658 | 77.3095 | 186.6417 | 503.0983
0.67913 | 1005774.3506 | 70.5549 | 170.3347 | 485.7955
0.67372 | 995621.8265 | 66.7899 | 161.2451 | 484.5298
0.68681 | 1036906.184 | 73.7014 | 177.9309 | 488.4418
0.71211 | 1099015.0945 | 67.7854 | 163.6485 | 477.7335
0.68171 | 1031033.9956 | 78.9309 | 190.5561 | 505.9712
0.70562 | 1072302.7427 | 66.7467 | 161.1409 | 507.8692
0.6767 | 1021033.2048 | 68.1931 | 164.6327 | 496.3968
0.70159 | 1058065.7183 | 73.6217 | 177.7384 | 506.0126
0.69705 | 1048867.5021 | 73.4566 | 177.3399 | 548.4945
0.68722 | 1023797.5346 | 74.4379 | 179.7089 | 523.4084
0.67571 | 1007367.2364 | 71.0245 | 171.4683 | 483.6415
0.68255 | 999839.753 | 82.5748 | 199.3531 | 467.9887
0.711 | 1107196.26 | 84.5886 | 204.215 | 490.4457
0.70405 | 1061994.0924 | 81.9358 | 197.8106 | 464.7479
0.6919 | 1048068.9103 | 75.6872 | 182.725 | 489.1113
0.69064 | 1047073.2661 | 75.7532 | 182.8844 | 521.1684
0.68434 | 1029468.6213 | 75.1656 | 181.4659 | 453.7295
0.67121 | 994174.9908 | 70.5643 | 170.3572 | 493.0962
0.68811 | 1043906.3817 | 70.444 | 170.0669 | 525.3567
0.68458 | 1018920.0162 | 74.3694 | 179.5437 | 462.9174
0.6898 | 1044679.0612 | 58.1676 | 140.4291 | 514.987
0.69952 | 1059667.4464 | 53.5197 | 129.208 | 507.2533
0.71782 | 1082506.8205 | 81.9167 | 197.7645 | 506.9824
0.69463 | 1053830.2318 | 60.4692 | 145.9856 | 492.7286
0.70128 | 1083017.5409 | 75.3907 | 182.0093 | 486.8685
0.70191 | 1080013.9573 | 76.8815 | 185.6084 | 469.1006
0.6779 | 1031264.4675 | 66.8171 | 161.3107 | 481.1176
0.68627 | 1018004.5758 | 82.1126 | 198.2373 | 477.152
0.67053 | 1008269.0341 | 66.0449 | 159.4466 | 476.628
0.71309 | 1105949.5563 | 85.8503 | 207.2611 | 495.1011
0.67842 | 1010136.4675 | 77.7061 | 187.5992 | 503.0677
0.72529 | 1127797.9029 | 70.2772 | 169.6643 | 498.9388
0.69358 | 1064840.6555 | 65.4187 | 157.9348 | 474.6672
0.69214 | 1042383.1449 | 69.9309 | 168.8281 | 506.3675
0.69706 | 1061901.7395 | 72.5754 | 175.2125 | 495.0492
0.69117 | 1036338.0806 | 67.7663 | 163.6024 | 491.0593
0.68762 | 1029729.6043 | 95.7213 | 231.0917 | 493.2506
0.72495 | 1136955.1583 | 70.8811 | 171.1221 | 478.0918
0.68123 | 1020968.4052 | 79.8282 | 192.7223 | 496.1553
0.69616 | 1055741.5844 | 64.6199 | 156.0064 | 460.2465
0.68014 | 1010494.612 | 88.1239 | 212.75 | 476.7923
0.68192 | 1007043.8634 | 64.8633 | 156.5939 | 529.4132
0.68733 | 1034529.7029 | 71.706 | 173.1137 | 517.954
0.67841 | 1003575.6408 | 73.9152 | 178.4472 | 496.3179
0.70514 | 1065411.1625 | 64.3815 | 155.4307 | 483.3331
0.69698 | 1061013.0039 | 81.0844 | 195.7551 | 499.0275
0.71661 | 1103975.7839 | 69.6262 | 168.0925 | 483.3061
0.68155 | 1018923.0303 | 72.982 | 176.1942 | 494.748
0.69292 | 1033885.4152 | 72.8383 | 175.8471 | 465.1131
0.69106 | 1043910.8818 | 75.4782 | 182.2205 | 513.6951
0.69161 | 1046466.914 | 60.3093 | 145.5995 | 491.3296
0.69536 | 1050030.5661 | 66.3242 | 160.1207 | 475.0714
0.70164 | 1079925.9802 | 65.5839 | 158.3335 | 479.7245
0.70664 | 1084565.2922 | 71.268 | 172.0563 | 479.254
0.66191 | 981430.559 | 63.1798 | 152.5296 | 458.4412
0.68197 | 1023287.2757 | 84.5827 | 204.2008 | 471.0278
0.69407 | 1042509.0073 | 73.3245 | 177.0211 | 510.4482
0.69168 | 1050727.5563 | 75.9032 | 183.2465 | 503.7254
0.68213 | 1028769.3818 | 70.0304 | 169.0683 | 498.7944
0.67335 | 998134.8267 | 66.9951 | 161.7405 | 509.8578
0.6924 | 1028847.5346 | 72.5388 | 175.1242 | 472.363
0.66005 | 963337.8827 | 65.9626 | 159.2479 | 459.3463
0.68185 | 1016746.5744 | 69.2407 | 167.1618 | 476.8504
0.69656 | 1055147.8886 | 72.1661 | 174.2244 | 489.947
0.71453 | 1078254.1936 | 73.6676 | 177.8494 | 447.6289
0.67224 | 1001875.4875 | 88.3987 | 213.4134 | 490.1936
0.70137 | 1074060.3515 | 67.2613 | 162.3831 | 480.211
0.68025 | 1010185.2106 | 72.4951 | 175.0186 | 475.7916
0.68122 | 1028552.7457 | 64.3359 | 155.3207 | 482.3433
0.70859 | 1093318.548 | 71.6451 | 172.9667 | 491.6557
0.67946 | 1032400.9264 | 79.297 | 191.44 | 478.8443
0.68315 | 1019997.5348 | 65.5792 | 158.3223 | 490.1013
0.6882 | 1055926.878 | 76.6649 | 185.0855 | 475.5484
0.69307 | 1054332.3918 | 72.4736 | 174.9667 | 459.3367
0.69004 | 1051484.0881 | 60.7356 | 146.6288 | 515.2423
0.67419 | 1001459.6833 | 83.7339 | 202.1514 | 501.1454
0.69667 | 1053549.6369 | 72.0453 | 173.9328 | 492.6328
0.71401 | 1102824.2441 | 82.2703 | 198.618 | 507.7306
0.69091 | 1035452.0203 | 71.3978 | 172.3695 | 472.0416
0.71526 | 1107106.838 | 87.253 | 210.6473 | 475.0654
0.68402 | 1029752.8875 | 68.0396 | 164.2621 | 493.1089
0.70258 | 1076594.3971 | 73.0019 | 176.2422 | 517.2274
0.68982 | 1048815.0298 | 79.3485 | 191.5643 | 473.6074
0.70022 | 1078033.2502 | 71.4887 | 172.5891 | 470.139
0.66866 | 989101.3534 | 68.687 | 165.8251 | 486.7207
0.68195 | 1028274.381 | 93.6646 | 226.1263 | 480.8641
0.70562 | 1075563.016 | 74.4116 | 179.6456 | 515.8346
0.68154 | 1039047.5736 | 77.8888 | 188.0402 | 477.918
0.70197 | 1087382.5664 | 79.1398 | 191.0604 | 450.3077
0.69255 | 1047025.5839 | 80.5281 | 194.412 | 477.8158
0.66743 | 1003529.31 | 81.6731 | 197.1764 | 490.0291
0.68063 | 1028569.9219 | 75.5086 | 182.2938 | 493.5429
0.70424 | 1071644.1239 | 68.7747 | 166.0368 | 487.9682
0.71039 | 1087369.1688 | 56.4839 | 136.3643 | 500.9641
0.68131 | 1021272.3978 | 74.8988 | 180.8216 | 489.2485
0.6605 | 986575.4432 | 79.5605 | 192.076 | 486.1343
0.69357 | 1043381.1422 | 71.6689 | 173.0241 | 472.7913
0.67028 | 1003627.2287 | 65.7658 | 158.7726 | 459.6758
0.68165 | 1026778.9616 | 68.0785 | 164.356 | 499.0011
0.68486 | 1018825.9829 | 72.6221 | 175.3252 | 501.8381
0.70712 | 1083716.2836 | 81.6317 | 197.0763 | 463.3861
0.69178 | 1048094.8145 | 66.5451 | 160.654 | 497.6219
0.69639 | 1040585.7648 | 74.6629 | 180.2521 | 498.9131
0.71773 | 1118936.4986 | 81.3608 | 196.4223 | 502.7874
0.69973 | 1063176.6816 | 85.6505 | 206.7787 | 497.0597
0.70018 | 1057671.4201 | 64.3664 | 155.3943 | 473.1563
0.70009 | 1076422.6804 | 86.6724 | 209.2456 | 490.5655
0.69131 | 1050362.8514 | 75.9179 | 183.2819 | 503.9114
0.67821 | 1014554.136 | 78.0697 | 188.477 | 494.4254
0.67859 | 1012680.7866 | 59.4813 | 143.6006 | 516.9643
0.66871 | 995697.8352 | 72.6876 | 175.4833 | 486.8712
0.6791 | 1005089.3722 | 73.9516 | 178.5348 | 492.0297
0.69121 | 1026541.1677 | 68.9606 | 166.4856 | 486.1717
0.67215 | 990937.6767 | 70.2354 | 169.5633 | 479.6203
0.70043 | 1053895.8003 | 73.4523 | 177.3296 | 467.7287
0.69101 | 1040917.5803 | 73.5374 | 177.5349 | 461.3655
0.66243 | 980535.6179 | 56.852 | 137.2529 | 520.0764
0.70171 | 1054186.4326 | 71.5712 | 172.7881 | 525.5618
0.72541 | 1120560.4908 | 67.8473 | 163.7979 | 499.5523
0.70759 | 1088893.7484 | 85.184 | 205.6523 | 503.9166
0.69172 | 1045459.3815 | 89.6463 | 216.4254 | 479.8236
0.70269 | 1074466.5545 | 69.107 | 166.8392 | 502.4675
0.71013 | 1090997.5419 | 67.4791 | 162.9089 | 481.1047
0.67925 | 1010780.8374 | 65.3593 | 157.7913 | 478.409
0.68909 | 1019977.8681 | 72.2582 | 174.4468 | 491.305
0.71813 | 1115744.0113 | 84.3954 | 203.7485 | 486.5891
0.67593 | 1003751.2853 | 72.4934 | 175.0145 | 495.5519
0.66431 | 964506.9741 | 57.87 | 139.7105 | 501.6291
0.69891 | 1064761.1601 | 58.5132 | 141.2635 | 480.9966
0.68491 | 1022747.8512 | 67.4883 | 162.9313 | 517.0758
0.70293 | 1079697.8236 | 83.4352 | 201.4304 | 469.7133
0.68885 | 1034806.948 | 75.0603 | 181.2116 | 491.5406
0.68319 | 1022683.5736 | 74.7954 | 180.5721 | 480.2742
0.6922 | 1044780.4543 | 70.0863 | 169.2032 | 474.0016
0.70395 | 1067045.5466 | 85.4469 | 206.2872 | 507.2181
0.69704 | 1058970.8492 | 70.5449 | 170.3105 | 490.4502
0.71364 | 1110470.4777 | 82.8833 | 200.0981 | 498.8731
0.69248 | 1036979.6827 | 82.8929 | 200.1212 | 498.4153
0.66742 | 985382.278 | 72.0743 | 174.0029 | 474.4184
0.70414 | 1092302.7114 | 85.8697 | 207.3078 | 498.031
0.67356 | 991783.8762 | 66.5481 | 160.6613 | 500.0056
0.69556 | 1050515.9495 | 77.1606 | 186.2822 | 479.1662
0.7157 | 1109465.2254 | 63.956 | 154.4033 | 494.167
0.67623 | 1012805.5088 | 73.1144 | 176.5138 | 507.5115
0.68665 | 1041585.109 | 66.6253 | 160.8478 | 504.7475
0.70059 | 1085642.51 | 61.7647 | 149.1131 | 483.6207
0.7025 | 1065972.1144 | 51.1149 | 123.4024 | 463.3114
0.69967 | 1058971.2223 | 70.2797 | 169.6701 | 488.8095
0.69309 | 1046503.1791 | 75.444 | 182.1378 | 487.192
0.69023 | 1033732.1366 | 71.3496 | 172.2531 | 492.9707
0.69048 | 1031761.4587 | 69.3855 | 167.5113 | 518.1231
0.6776 | 1016807.7263 | 91.8334 | 221.7054 | 518.4926
0.69165 | 1038880.6249 | 84.6704 | 204.4124 | 477.7664
0.68789 | 1027268.3452 | 63.6328 | 153.6232 | 486.9728
0.68969 | 1038737.6079 | 71.7513 | 173.223 | 488.821
0.65724 | 949414.7744 | 73.8008 | 178.171 | 480.4917
0.72992 | 1136192.0078 | 81.2367 | 196.1226 | 489.5275
0.71624 | 1100159.5972 | 71.3237 | 172.1906 | 501.8424
0.71072 | 1084332.8265 | 79.6229 | 192.2268 | 472.2327
0.70203 | 1080808.9259 | 65.601 | 158.3749 | 483.3314
0.68741 | 1044777.0806 | 70.5739 | 170.3804 | 476.9413
0.68658 | 1041279.8677 | 79.1493 | 191.0834 | 513.6973
0.6895 | 1055065.9706 | 76.4583 | 184.5866 | 466.313
0.68537 | 1022030.9061 | 79.4292 | 191.7589 | 476.6772
0.68108 | 1018960.7216 | 59.632 | 143.9644 | 486.4285
0.69231 | 1047686.9165 | 66.4681 | 160.4683 | 482.8967
0.6852 | 1009133.152 | 58.7012 | 141.7171 | 492.2378
0.69274 | 1054442.8325 | 66.8552 | 161.4027 | 501.5223
0.68944 | 1028842.5833 | 50.2693 | 121.3609 | 497.694
0.71864 | 1131259.6915 | 76.8466 | 185.5242 | 477.9258
0.68812 | 1050784.0691 | 74.5461 | 179.9702 | 491.8594
0.68092 | 1010921.0176 | 75.8489 | 183.1155 | 499.8021
0.72695 | 1130356.6027 | 74.0778 | 178.8396 | 471.9118
0.689 | 1042784.753 | 50.9826 | 123.0828 | 483.9284
0.70493 | 1080403.5158 | 78.3018 | 189.0372 | 507.1593
0.68452 | 1036633.2415 | 60.361 | 145.7244 | 493.8545
0.6928 | 1052260.9034 | 82.504 | 199.1823 | 471.6229
0.68329 | 1032498.8793 | 55.7602 | 134.6171 | 472.8811
0.7101 | 1061409.3612 | 72.9685 | 176.1616 | 504.7187
0.69371 | 1048059.0143 | 83.6296 | 201.8996 | 495.8227
0.71045 | 1077643.2985 | 75.7914 | 182.9767 | 479.2734
0.69815 | 1060335.9 | 75.8881 | 183.2101 | 495.821
0.69484 | 1031214.021 | 75.822 | 183.0506 | 457.8365
0.71141 | 1106582.7246 | 66.2801 | 160.0143 | 505.3492
0.69241 | 1046980.1896 | 77.2517 | 186.5021 | 510.3654
0.70864 | 1088027.9946 | 83.4017 | 201.3495 | 468.6997
0.69676 | 1051964.6846 | 88.0954 | 212.6811 | 512.8373
0.71298 | 1090434.4593 | 72.8615 | 175.9032 | 500.5259
0.66369 | 972742.0701 | 68.0571 | 164.3043 | 497.619
0.67935 | 1012581.6721 | 77.788 | 187.7968 | 475.5972
0.69592 | 1049673.736 | 73.6327 | 177.7652 | 510.9106
0.67019 | 995209.3772 | 67.1142 | 162.028 | 510.7238
0.70199 | 1067297.5911 | 66.7788 | 161.2184 | 492.3284
0.68271 | 996414.9938 | 64.1041 | 154.761 | 475.2166
0.68393 | 1030879.8679 | 71.8085 | 173.3612 | 506.5897
0.69986 | 1067300.0813 | 66.9821 | 161.7092 | 498.3655
0.68619 | 1019827.0763 | 54.8433 | 132.4035 | 480.4886
0.67727 | 1012023.8976 | 59.5593 | 143.7888 | 478.1927
0.69159 | 1045965.8684 | 85.2552 | 205.8243 | 482.6477
0.67357 | 1021655.7467 | 62.5784 | 151.0777 | 491.7831
0.67904 | 1021232.5486 | 62.0486 | 149.7985 | 466.3833
0.68176 | 1038404.2145 | 72.5627 | 175.1819 | 477.6285
0.71557 | 1088270.1671 | 68.7358 | 165.943 | 476.8521
0.68577 | 1036664.5565 | 88.2059 | 212.9478 | 484.8021
0.72667 | 1134414.8964 | 86.4146 | 208.6234 | 474.551
0.69018 | 1048573.45 | 72.3791 | 174.7387 | 478.3039
0.68883 | 1020910.6826 | 70.6082 | 170.4632 | 443.888
0.70731 | 1076685.9779 | 70.6829 | 170.6437 | 472.4977
0.67847 | 1015727.4112 | 57.5006 | 138.8188 | 486.5963
0.69806 | 1062652.8716 | 68.8009 | 166.1 | 487.1608
0.69884 | 1083435.765 | 66.0608 | 159.4848 | 511.8931
0.67959 | 1022394.5045 | 62.0683 | 149.846 | 489.0422
0.66802 | 995732.5728 | 74.4788 | 179.8078 | 527.4948
0.68982 | 1038758.6187 | 62.1266 | 149.9868 | 501.5946
0.67191 | 996285.2796 | 61.1748 | 147.6891 | 510.3995
0.69471 | 1072017.7206 | 75.3304 | 181.8637 | 465.8642
0.70792 | 1069049.0987 | 71.9436 | 173.6873 | 485.1619
0.67803 | 1004371.3746 | 63.1194 | 152.3836 | 496.0141
0.67866 | 1029789.3823 | 51.1143 | 123.4007 | 481.0112
0.70934 | 1089453.5185 | 67.4588 | 162.8598 | 489.3243
0.66748 | 991357.7004 | 67.1255 | 162.0552 | 493.5121
0.68557 | 1052780.5133 | 81.7945 | 197.4695 | 526.438
0.68996 | 1044685.8809 | 83.8133 | 202.3432 | 451.1012
0.66443 | 986562.1307 | 70.7376 | 170.7756 | 489.6044
0.67399 | 992912.6002 | 70.1836 | 169.4383 | 475.6158
0.67234 | 998650.9308 | 75.885 | 183.2027 | 493.0459
0.69961 | 1067333.603 | 76.229 | 184.0331 | 480.0801
0.6819 | 1038433.4799 | 66.8872 | 161.48 | 496.6784
0.68662 | 1025115.2773 | 61.597 | 148.7084 | 489.618
0.70415 | 1076202.154 | 72.812 | 175.7838 | 484.0882
0.71087 | 1086615.3758 | 76.3764 | 184.389 | 488.4631
0.69435 | 1047827.6142 | 64.3021 | 155.239 | 527.8209
0.68171 | 1013375.7562 | 67.1843 | 162.1973 | 515.8475
0.73011 | 1125785.0294 | 71.5987 | 172.8546 | 491.3419
0.69164 | 1039461.0536 | 67.9239 | 163.9827 | 473.6087
0.70163 | 1089722.0995 | 73.992 | 178.6324 | 534.9536
0.67843 | 1011591.5423 | 76.1785 | 183.9111 | 484.8213
0.68723 | 1031531.3432 | 73.2741 | 176.8994 | 510.6211
0.70453 | 1087877.6135 | 67.8212 | 163.735 | 525.6599
0.68343 | 1006087.4643 | 60.7358 | 146.6291 | 477.3515
0.68565 | 1033290.7604 | 80.271 | 193.7914 | 504.8232
0.68836 | 1057060.9536 | 74.9647 | 180.9808 | 479.4965
0.66354 | 972128.3 | 58.3952 | 140.9784 | 490.185
0.69382 | 1047283.9234 | 68.2199 | 164.6973 | 498.6605
0.70096 | 1059551.0723 | 80.5138 | 194.3775 | 501.6019
0.67728 | 1027669.8386 | 79.1878 | 191.1762 | 465.8102
0.66629 | 989397.0941 | 80.6612 | 194.7335 | 476.1429
0.72971 | 1122915.8113 | 76.8906 | 185.6302 | 469.0778
0.67061 | 1007743.9804 | 71.5906 | 172.8351 | 513.0055
0.6829 | 1020373.7013 | 71.1833 | 171.8517 | 468.2032
0.66906 | 995913.2748 | 65.3791 | 157.839 | 483.164
0.67915 | 996747.628 | 74.9813 | 181.021 | 511.1896
0.68004 | 1032864.2644 | 75.8897 | 183.2139 | 463.0587
0.67979 | 1024414.949 | 64.212 | 155.0215 | 522.9214
0.68894 | 1039329.3865 | 83.5075 | 201.6049 | 515.4395
0.67226 | 990774.8647 | 68.4347 | 165.216 | 468.9192
0.70192 | 1066289.7819 | 60.5492 | 146.1788 | 485.741
0.69394 | 1035423.0999 | 69.2836 | 167.2655 | 487.1991
0.68006 | 1014657.6914 | 81.0453 | 195.6606 | 454.6698
0.68379 | 1005658.7238 | 81.411 | 196.5436 | 485.3865
0.6742 | 1014070.3128 | 73.4267 | 177.2678 | 492.2533
0.68356 | 1025175.808 | 86.0386 | 207.7155 | 513.3477
0.67427 | 1013645.9916 | 81.6598 | 197.1442 | 468.358
0.68082 | 1023691.9451 | 80.0891 | 193.3523 | 470.3769
0.66493 | 985780.1994 | 80.5313 | 194.4198 | 496.378
0.69742 | 1072028.5151 | 71.3761 | 172.3171 | 492.8293
0.69138 | 1044754.1619 | 78.8623 | 190.3904 | 515.7174
0.68555 | 1039059.6648 | 75.7318 | 182.8327 | 525.8746
0.66568 | 980967.4037 | 76.9065 | 185.6688 | 494.6673
0.66554 | 981884.6898 | 64.7994 | 156.4395 | 493.2403
0.6699 | 992894.4619 | 77.3517 | 186.7435 | 501.9338
0.69124 | 1062690.7768 | 88.197 | 212.9264 | 492.8724
0.69987 | 1059405.2379 | 70.2885 | 169.6915 | 519.5947
0.70388 | 1088589.0727 | 93.2007 | 225.0065 | 490.1956
0.66805 | 989941.4979 | 68.8081 | 166.1174 | 483.7573
0.70365 | 1075312.2203 | 67.1217 | 162.0462 | 471.3683
0.6969 | 1048325.9527 | 64.0979 | 154.7459 | 488.6958
0.69518 | 1058887.5257 | 81.1322 | 195.8704 | 485.2086
0.69742 | 1062479.7606 | 74.37 | 179.545 | 502.5536
0.71015 | 1108921.6085 | 72.5088 | 175.0518 | 510.7867
0.69278 | 1056363.8411 | 76.1894 | 183.9375 | 474.2361
0.68599 | 1034769.6196 | 70.0426 | 169.0978 | 494.0342
0.70793 | 1078689.9323 | 72.5916 | 175.2517 | 481.1535
0.70006 | 1066535.8487 | 82.756 | 199.7905 | 494.6893
0.67029 | 996008.1559 | 75.9159 | 183.2771 | 494.3818
0.70219 | 1072542.6429 | 79.714 | 192.4467 | 478.7927
0.69186 | 1042260.9568 | 75.2723 | 181.7235 | 468.4895
0.68782 | 1042818.2999 | 92.7833 | 223.9986 | 496.2847
0.70086 | 1064925.002 | 68.6726 | 165.7903 | 504.7023
0.67488 | 1007687.0842 | 67.4835 | 162.9197 | 476.7183
0.6682 | 1000036.1128 | 61.7464 | 149.0691 | 467.8601
0.70818 | 1089807.2607 | 75.1237 | 181.3647 | 494.0808
0.70028 | 1055155.7708 | 75.2871 | 181.7592 | 493.5795
0.70172 | 1067526.5198 | 79.3947 | 191.6757 | 484.1461
0.69559 | 1049127.5359 | 70.035 | 169.0796 | 488.3717
0.68311 | 1047059.9617 | 54.1536 | 130.7384 | 524.0858
0.71098 | 1105549.8639 | 77.6962 | 187.5753 | 490.2356
0.69123 | 1039859.4453 | 96.1584 | 232.1469 | 511.3584
0.7163 | 1115474.5971 | 70.3822 | 169.9177 | 499.0347
0.70778 | 1094203.0616 | 67.9015 | 163.9288 | 492.8599
0.66933 | 978081.3003 | 66.0438 | 159.4439 | 515.8675
0.67355 | 1000308.8496 | 70.7527 | 170.8121 | 508.2476
0.72316 | 1112956.6236 | 69.5568 | 167.9251 | 508.2387
0.68614 | 1035130.3042 | 70.0518 | 169.12 | 494.8448
0.71209 | 1097216.6461 | 76.9899 | 185.87 | 470.2376
0.67559 | 1014389.0913 | 69.7203 | 168.3196 | 513.2445
0.6943 | 1056008.6921 | 72.4601 | 174.9341 | 487.2034
0.67259 | 1005574.4186 | 57.423 | 138.6314 | 514.6302
0.67709 | 1009886.0161 | 84.0768 | 202.9795 | 488.3269
0.6903 | 1044338.8474 | 72.074 | 174.0021 | 513.3087
0.69734 | 1070308.8096 | 71.2695 | 172.0597 | 493.3949
0.71292 | 1093926.5778 | 63.3068 | 152.836 | 494.704
0.68894 | 1042798.2407 | 75.6307 | 182.5886 | 511.8266
0.67652 | 1021293.9877 | 82.0031 | 197.9731 | 494.6905
0.69067 | 1035792.8798 | 57.6003 | 139.0593 | 499.3474
0.69267 | 1044601.2989 | 64.2597 | 155.1367 | 481.2738
0.68549 | 1019632.3629 | 63.3812 | 153.0158 | 501.1215
0.68765 | 1050651.6272 | 76.3427 | 184.3076 | 502.8931
0.69019 | 1031396.5223 | 84.2388 | 203.3705 | 495.7511
0.70084 | 1046552.3907 | 71.1409 | 171.7493 | 500.4039
0.68596 | 1046604.8885 | 70.8471 | 171.0401 | 483.2694
0.71071 | 1078835.2071 | 79.2197 | 191.2533 | 465.0071
0.69632 | 1065840.6207 | 79.8836 | 192.8561 | 522.8304
0.66513 | 983446.7003 | 66.8253 | 161.3305 | 467.9192
0.70801 | 1086249.5161 | 65.3745 | 157.8281 | 475.9003
0.66604 | 978826.1558 | 68.2407 | 164.7476 | 475.3274
0.68711 | 1036038.071 | 78.6901 | 189.9748 | 481.2762
0.7088 | 1098527.0186 | 79.1805 | 191.1585 | 473.9203
0.68344 | 1005965.2883 | 75.1994 | 181.5473 | 513.1167
0.71138 | 1085365.424 | 81.2665 | 196.1947 | 491.9347
0.68145 | 1012612.3531 | 67.1592 | 162.1367 | 463.0506
0.68467 | 1025927.1172 | 69.5784 | 167.977 | 493.7592
0.68737 | 1039104.7764 | 58.0625 | 140.1753 | 501.9966
0.68601 | 1030798.8621 | 81.6214 | 197.0515 | 498.5828
0.68459 | 1029497.0764 | 75.1402 | 181.4044 | 496.6951
0.6939 | 1046587.7594 | 68.1675 | 164.5708 | 492.3089
0.66517 | 975519.8316 | 58.455 | 141.1229 | 497.4858
0.67278 | 991922.2889 | 76.9441 | 185.7595 | 464.9875
0.67362 | 1016797.8198 | 78.7759 | 190.1818 | 501.8225
0.69636 | 1050489.7266 | 77.0703 | 186.0641 | 500.2963
0.69582 | 1055407.7667 | 66.62 | 160.8349 | 506.3299
0.68538 | 1029030.0566 | 83.11 | 200.6453 | 489.2837
0.69013 | 1060189.2039 | 80.1723 | 193.553 | 467.8594
0.6943 | 1045779.4837 | 76.9527 | 185.7804 | 477.1704
0.67591 | 1013774.2868 | 76.6269 | 184.9936 | 521.1243
0.66506 | 996759.987 | 82.9624 | 200.289 | 495.6095
0.68573 | 1027561.4286 | 84.4711 | 203.9312 | 502.0057
0.71957 | 1106517.2716 | 76.8352 | 185.4965 | 483.577
0.69014 | 1030078.3902 | 71.1063 | 171.6659 | 480.8583
0.71206 | 1097471.7958 | 71.5677 | 172.7798 | 489.5932
0.67966 | 1012895.554 | 79.3801 | 191.6404 | 461.8109
0.67564 | 1010948.2998 | 76.4444 | 184.5532 | 484.0582
0.66795 | 1006228.5667 | 62.1516 | 150.0473 | 464.9469
0.65224 | 954486.8056 | 62.3561 | 150.5411 | 516.1071
0.69602 | 1056705.1688 | 84.0395 | 202.8894 | 504.7605
0.67055 | 984563.9015 | 81.2772 | 196.2206 | 502.5651
0.67028 | 990483.1591 | 84.7743 | 204.6632 | 481.9737
0.68884 | 1030376.3381 | 71.2617 | 172.0409 | 492.8411
0.70929 | 1086948.9711 | 81.6254 | 197.0611 | 468.3738
0.6737 | 994018.059 | 78.5897 | 189.7322 | 484.1123
0.66914 | 984035.8166 | 68.2858 | 164.8565 | 503.3268
0.7041 | 1079054.6172 | 65.3016 | 157.652 | 516.3863
0.67981 | 1017310.1079 | 94.3162 | 227.6994 | 468.8806
0.68317 | 1039132.1144 | 58.4202 | 141.0388 | 514.7832
0.68974 | 1035822.6698 | 75.0373 | 181.1561 | 498.0687
0.70211 | 1084109.6603 | 78.8857 | 190.447 | 479.6108
0.69665 | 1054484.6191 | 77.3319 | 186.6957 | 483.3009
0.66959 | 1009233.7375 | 68.1202 | 164.4566 | 506.0378
0.68663 | 1031229.2605 | 78.3806 | 189.2274 | 500.4063
0.6911 | 1039369.7189 | 66.0556 | 159.4724 | 505.7433
0.69572 | 1051953.4788 | 80.5376 | 194.4349 | 519.5421
0.67738 | 1008733.0154 | 65.407 | 157.9063 | 472.9174
0.67133 | 993560.9319 | 75.6111 | 182.5414 | 462.8536
0.69129 | 1056551.53 | 81.4083 | 196.537 | 501.4352
0.6898 | 1039256.0562 | 66.1732 | 159.7563 | 488.6033
0.68518 | 1042894.7302 | 68.7909 | 166.076 | 496.9474
0.68201 | 1032831.2264 | 79.0335 | 190.8038 | 470.6495
0.68997 | 1051372.2216 | 62.8677 | 151.7762 | 488.5778
0.68398 | 1034280.3745 | 66.4171 | 160.3451 | 468.393
0.67897 | 1004736.9219 | 70.4985 | 170.1985 | 522.6585
0.68607 | 1040251.0231 | 85.4721 | 206.3478 | 508.6556
0.6709 | 1005071.7121 | 74.8542 | 180.7141 | 489.7151
0.68399 | 1040588.9798 | 72.4782 | 174.9778 | 475.5488
0.70325 | 1059697.2117 | 89.3784 | 215.7786 | 529.5885
0.69027 | 1053167.5738 | 72.6867 | 175.4813 | 497.9527
0.70931 | 1075162.448 | 72.2129 | 174.3373 | 458.0462
0.71207 | 1104289.8178 | 73.882 | 178.3669 | 515.2699
0.70868 | 1093700.6114 | 71.6613 | 173.0057 | 473.9946
0.69652 | 1049891.7559 | 69.9918 | 168.9753 | 475.4287
0.70325 | 1076410.537 | 78.1988 | 188.7885 | 473.2795
0.70037 | 1067041.6347 | 58.516 | 141.2702 | 510.7259
0.68246 | 1027921.9248 | 76.9508 | 185.7758 | 485.835
0.67636 | 1008557.0457 | 65.4941 | 158.1167 | 490.4332
0.70324 | 1078956.5608 | 71.2425 | 171.9947 | 499.9532
0.68114 | 1017172.144 | 63.6316 | 153.6203 | 472.012
0.68683 | 1056836.1086 | 75.1154 | 181.3446 | 477.6104
0.70663 | 1087883.1859 | 80.3463 | 193.9732 | 481.9759
0.67623 | 1003531.5346 | 68.9092 | 166.3615 | 473.2578
0.69323 | 1045216.5276 | 66.223 | 159.8765 | 470.4341
0.68723 | 1042616.4042 | 66.6358 | 160.8731 | 505.7146
0.68382 | 1055292.8332 | 82.1998 | 198.4479 | 495.1761
0.69221 | 1059859.7972 | 69.1998 | 167.0631 | 521.5562
0.68189 | 1021745.2867 | 78.1992 | 188.7896 | 476.5012
0.67387 | 986255.7737 | 67.9593 | 164.0683 | 482.1958
0.68583 | 1023764.9335 | 65.5645 | 158.2867 | 511.5972
0.69723 | 1054450.4718 | 78.9266 | 190.5458 | 487.0765
0.70808 | 1070267.6261 | 72.009 | 173.8451 | 508.683
0.68452 | 1026651.9934 | 57.6052 | 139.0712 | 490.6022
0.69182 | 1052856.0006 | 87.3073 | 210.7784 | 467.4607
0.6842 | 1023970.1303 | 63.7703 | 153.9552 | 486.707
0.64574 | 941224.2548 | 67.6817 | 163.3981 | 473.0852
0.65754 | 957244.3241 | 69.3554 | 167.4388 | 502.1525
0.68973 | 1028509.9672 | 72.7953 | 175.7433 | 480.9896
0.69603 | 1050455.9413 | 80.3879 | 194.0736 | 473.1627
0.65098 | 939040.7318 | 48.5922 | 117.3119 | 492.763
0.69713 | 1044827.9227 | 84.4859 | 203.9669 | 531.109
0.68483 | 1019575.684 | 76.0617 | 183.6291 | 479.2355
0.68197 | 1004280.232 | 84.4815 | 203.9564 | 498.8024
0.69665 | 1039886.9774 | 79.5076 | 191.9483 | 511.2924
0.69745 | 1056892.138 | 85.4227 | 206.2286 | 482.0955
0.69958 | 1060203.5173 | 75.7226 | 182.8104 | 478.5626
0.67036 | 1012274.912 | 70.1176 | 169.2788 | 495.349
0.68742 | 1026364.7999 | 71.1305 | 171.7243 | 475.2807
0.68984 | 1037192.2253 | 76.8137 | 185.4448 | 487.3145
0.67664 | 1010386.5207 | 71.1387 | 171.7441 | 463.9096
0.70897 | 1099976.5086 | 80.971 | 195.4812 | 499.9246
0.67227 | 1010740.1158 | 62.0288 | 149.7509 | 478.9354
0.69615 | 1059338.0223 | 81.0539 | 195.6815 | 503.5289
0.70664 | 1082484.381 | 67.2829 | 162.4353 | 543.3065
0.68599 | 1024937.5552 | 85.7139 | 206.9316 | 478.0957
0.68451 | 1028845.5624 | 64.045 | 154.6183 | 494.5827
0.69212 | 1059808.1571 | 65.2477 | 157.522 | 475.4229
0.69265 | 1048907.2741 | 80.6603 | 194.7313 | 498.1529
0.70905 | 1094330.7553 | 82.208 | 198.4677 | 508.5135
0.69374 | 1059630.5753 | 79.714 | 192.4465 | 488.882
0.70707 | 1088377.6473 | 88.3066 | 213.1911 | 476.3163
0.71911 | 1121234.9335 | 70.4692 | 170.1277 | 476.4173
0.69107 | 1047488.566 | 80.736 | 194.9139 | 519.89
0.70088 | 1059079.1501 | 70.1654 | 169.3943 | 506.2707
0.6953 | 1056141.2117 | 74.1522 | 179.0192 | 489.1168
0.6809 | 1021112.2709 | 71.4445 | 172.4824 | 456.8478
0.69772 | 1078467.3472 | 52.0681 | 125.7034 | 467.1916
0.68685 | 1027186.3056 | 74.7249 | 180.4018 | 489.6638
0.68313 | 1026237.3125 | 78.3998 | 189.274 | 482.6059
0.69302 | 1045734.2786 | 66.276 | 160.0043 | 488.6026
0.67139 | 1010703.1588 | 84.8592 | 204.8683 | 507.166
0.71519 | 1117420.8745 | 76.2723 | 184.1376 | 498.9117
0.6888 | 1040236.7855 | 66.3313 | 160.138 | 496.5574
0.6784 | 987661.8046 | 74.3791 | 179.5671 | 513.3434
0.67493 | 1015393.5468 | 66.0781 | 159.5267 | 487.0018
0.70341 | 1061078.9836 | 60.1256 | 145.1561 | 497.6658
0.672 | 980285.5629 | 74.5275 | 179.9254 | 485.5038
0.68108 | 1028358.5278 | 62.2937 | 150.3904 | 487.4184
0.68266 | 1027967.2843 | 62.1264 | 149.9863 | 478.4818
0.69215 | 1050834.4953 | 79.2328 | 191.2848 | 490.9529
0.70136 | 1063907.7095 | 60.4616 | 145.9673 | 522.189
0.69358 | 1035179.818 | 79.4988 | 191.9271 | 481.4799
0.68864 | 1014160.8272 | 78.4428 | 189.3777 | 482.9194
0.68343 | 1028864.4967 | 64.148 | 154.867 | 479.0346
0.70446 | 1065281.9188 | 73.8215 | 178.2209 | 505.6865
0.69366 | 1056363.0685 | 62.8296 | 151.6841 | 499.9888
0.68804 | 1025512.7783 | 74.4744 | 179.797 | 484.3702
0.68754 | 1037305.3277 | 60.135 | 145.1787 | 490.3433
0.68944 | 1030187.0196 | 85.7217 | 206.9504 | 527.0076
0.68501 | 1037225.2157 | 66.5053 | 160.5579 | 508.5965
0.68115 | 1022742.7701 | 76.7108 | 185.1963 | 490.8721
0.66672 | 984821.7967 | 62.9403 | 151.9513 | 495.4088
0.704 | 1068013.6214 | 66.6324 | 160.8648 | 492.9832
0.69132 | 1030558.6565 | 71.0185 | 171.4539 | 517.1229
0.6928 | 1052463.8493 | 75.9065 | 183.2545 | 512.4995
0.67782 | 996380.5217 | 93.9608 | 226.8415 | 496.6667
0.68855 | 1036697.5817 | 57.828 | 139.6091 | 474.513
0.68883 | 1039380.0711 | 42.5264 | 102.6677 | 527.7994
0.68629 | 1044994.5813 | 70.7036 | 170.6937 | 488.5389
0.66675 | 1002952.852 | 69.8479 | 168.6278 | 496.5
0.69105 | 1039655.8139 | 83.2792 | 201.0538 | 497.0333
0.70488 | 1073382.1004 | 84.2929 | 203.5011 | 498.3793
0.67744 | 1017789.878 | 81.6102 | 197.0244 | 509.9951
0.68221 | 1015311.0461 | 63.4154 | 153.0984 | 482.8298
0.69348 | 1068753.3382 | 73.4214 | 177.255 | 485.0298
0.69744 | 1073566.978 | 80.3335 | 193.9423 | 491.8431
0.68616 | 1024433.8522 | 69.3262 | 167.3683 | 489.3329
0.70847 | 1080785.6886 | 77.9377 | 188.1582 | 480.2208
0.68647 | 1023540.5845 | 67.3875 | 162.6877 | 508.2409
0.66894 | 986842.0176 | 68.5015 | 165.3773 | 473.2375
0.67832 | 1017653.5596 | 81.4934 | 196.7426 | 475.5723
0.66865 | 984050.3109 | 66.5144 | 160.5799 | 489.1434
0.6922 | 1033077.0846 | 87.0843 | 210.2401 | 498.4506
0.68512 | 1039625.4698 | 70.4604 | 170.1065 | 506.101
0.68026 | 1038542.7626 | 56.3365 | 136.0083 | 492.0769
0.68403 | 1020954.998 | 67.4575 | 162.8568 | 506.1034
0.68749 | 1020844.2708 | 65.2196 | 157.4541 | 502.1154
0.70447 | 1065485.9543 | 73.8341 | 178.2512 | 493.7792
0.68969 | 1039613.869 | 64.2295 | 155.0637 | 505.0056
0.67339 | 998768.9334 | 74.1099 | 178.9172 | 510.149
0.69945 | 1080118.9197 | 83.8833 | 202.5121 | 480.067
0.73201 | 1148737.7878 | 72.8685 | 175.92 | 493.6347
0.68674 | 1032719.7385 | 66.0498 | 159.4584 | 496.2744
0.68249 | 1023258.1388 | 71.4278 | 172.4419 | 523.4924
0.68448 | 1021489.12 | 79.4906 | 191.9072 | 499.9394
0.71056 | 1092528.5785 | 76.8256 | 185.4735 | 480.593
0.70632 | 1074955.2135 | 62.7896 | 151.5876 | 478.8797
0.66946 | 1004300.2844 | 73.2341 | 176.8027 | 514.6281
0.67135 | 1003458.456 | 71.2078 | 171.9109 | 478.2364
0.68607 | 1040819.5106 | 70.444 | 170.0668 | 465.5211
0.6794 | 1032874.8656 | 62.7227 | 151.4261 | 499.9542
0.70077 | 1074721.8526 | 90.9658 | 219.611 | 492.9
0.70569 | 1087322.8714 | 67.6753 | 163.3826 | 524.4787
0.71771 | 1098981.576 | 67.2547 | 162.3671 | 505.5121
0.70717 | 1086859.0465 | 65.4885 | 158.1033 | 483.0234
0.67802 | 1034185.2154 | 76.213 | 183.9946 | 488.1354
0.68886 | 1045505.2405 | 72.3062 | 174.5625 | 525.8795
0.68558 | 1036464.297 | 68.3269 | 164.9558 | 509.1842
0.68468 | 1038107.2484 | 72.2329 | 174.3856 | 502.4616
0.71165 | 1071384.0434 | 76.9145 | 185.688 | 510.5288
0.68898 | 1021307.7034 | 66.9959 | 161.7423 | 482.1124
0.67316 | 994100.6069 | 81.7456 | 197.3514 | 472.4951
0.69671 | 1084441.8204 | 72.2754 | 174.4883 | 524.9978
0.69059 | 1040441.9459 | 66.1726 | 159.7548 | 470.3323
0.67713 | 1016515.1073 | 86.1225 | 207.9181 | 480.2097
0.67768 | 997969.335 | 78.4392 | 189.369 | 507.7516
0.68311 | 1021009.7671 | 69.4027 | 167.5528 | 482.552
0.70259 | 1073186.1066 | 82.9713 | 200.3105 | 469.7015
0.71284 | 1108171.6515 | 75.8911 | 183.2173 | 477.3706
0.69517 | 1049324.1314 | 70.0905 | 169.2134 | 513.9385
0.69214 | 1044731.1178 | 81.991 | 197.9439 | 484.5953
0.677 | 998496.7904 | 76.9714 | 185.8253 | 533.3711
0.69572 | 1053231.6025 | 68.4206 | 165.1819 | 516.3308
0.70433 | 1078694.2316 | 76.5242 | 184.7459 | 524.6447
0.70394 | 1062310.998 | 70.0737 | 169.1729 | 467.7922
0.67816 | 1025960.0893 | 45.5818 | 110.0441 | 475.4446
0.68936 | 1038376.1527 | 86.4138 | 208.6214 | 505.5001
0.69294 | 1043377.173 | 73.5056 | 177.4581 | 474.5837
0.67334 | 1000418.401 | 74.238 | 179.2263 | 470.6092
0.71735 | 1122707.2131 | 78.2005 | 188.7926 | 505.1621
0.6749 | 1008973.616 | 72.9819 | 176.1938 | 502.3402
0.72934 | 1141334.6664 | 70.7263 | 170.7484 | 483.4225
0.67633 | 1010476.0235 | 87.1769 | 210.4637 | 509.7979
0.68439 | 1035445.4092 | 75.5962 | 182.5053 | 473.5409
0.68445 | 1038523.6381 | 77.5189 | 187.1472 | 475.2448
0.68726 | 1034895.5103 | 71.7385 | 173.1921 | 469.2933
0.67849 | 1008587.9579 | 67.184 | 162.1964 | 517.2862
0.70523 | 1079447.8172 | 61.008 | 147.2864 | 498.5686
0.67899 | 1022192.077 | 69.9048 | 168.7651 | 517.9865
0.67705 | 1008097.1101 | 65.3976 | 157.8838 | 490.4703
0.69381 | 1044829.1881 | 76.9368 | 185.742 | 490.5723
0.69086 | 1027490.264 | 80.788 | 195.0395 | 480.4259
0.68149 | 1046811.0426 | 71.3259 | 172.196 | 481.4093
0.71062 | 1094720.5578 | 71.5323 | 172.6943 | 487.0743
0.68308 | 1028469.7295 | 76.9219 | 185.7059 | 472.3865
0.70191 | 1076064.1462 | 65.3242 | 157.7067 | 477.7335
0.70131 | 1055506.934 | 84.8406 | 204.8233 | 514.2038
0.67475 | 1009192.9607 | 69.4494 | 167.6657 | 501.6953
0.70494 | 1073274.8099 | 62.8148 | 151.6484 | 511.5385
0.70257 | 1056564.802 | 66.0522 | 159.4641 | 486.4421
0.67697 | 1015982.4331 | 84.5994 | 204.241 | 468.0485
0.69449 | 1051961.3424 | 68.4592 | 165.275 | 473.7523
0.68746 | 1036946.2399 | 77.0616 | 186.0431 | 480.411
0.70742 | 1089864.2104 | 65.6464 | 158.4844 | 506.2684
0.70441 | 1084089.0921 | 67.5963 | 163.1918 | 499.7095
0.70852 | 1074838.0303 | 82.4045 | 198.9421 | 494.7297
0.69028 | 1041184.1379 | 64.2716 | 155.1655 | 484.4406
0.6908 | 1034445.2473 | 55.2582 | 133.4051 | 499.1858
0.68609 | 1034010.8765 | 73.5728 | 177.6205 | 485.5661
0.67315 | 1000788.7442 | 80.7588 | 194.969 | 474.7042
0.702 | 1058021.2087 | 69.9832 | 168.9543 | 464.8886
0.71052 | 1096138.4778 | 65.7362 | 158.7012 | 489.2707
0.68122 | 1010741.9762 | 88.0065 | 212.4664 | 489.8337
0.68668 | 1030998.9906 | 76.7074 | 185.188 | 476.0027
0.6774 | 1017657.2641 | 81.6048 | 197.0115 | 509.2916
0.71911 | 1125558.9321 | 81.6852 | 197.2056 | 478.0897
0.68772 | 1030204.4449 | 76.5323 | 184.7654 | 476.8869
0.69672 | 1058178.2721 | 72.3561 | 174.6831 | 484.5536
0.69948 | 1055682.4573 | 70.0806 | 169.1894 | 502.5904
0.70267 | 1062683.4473 | 78.9273 | 190.5474 | 445.5249
0.67104 | 994405.8179 | 77.3374 | 186.7091 | 509.7244
0.65799 | 965299.2181 | 84.1619 | 203.1848 | 475.2381
0.68452 | 1017283.3226 | 74.2566 | 179.2712 | 517.9665
0.6926 | 1048180.435 | 68.3993 | 165.1306 | 493.2743
0.70003 | 1061659.0598 | 74.0362 | 178.7393 | 484.5641
0.68268 | 1026839.6198 | 75.1556 | 181.4417 | 500.8617
0.70786 | 1088930.6483 | 73.2877 | 176.9321 | 483.0755
0.6837 | 1019687.6352 | 79.44 | 191.7851 | 491.2438
0.69822 | 1065974.4317 | 76.2831 | 184.1636 | 460.6469
0.72663 | 1134371.9191 | 55.1551 | 133.1563 | 506.5779
0.67437 | 978453.355 | 72.4969 | 175.0229 | 477.2344
0.67685 | 1013285.4031 | 74.8027 | 180.5897 | 479.6028
0.70863 | 1093274.7228 | 85.4014 | 206.1773 | 490.1987
0.67914 | 1002985.5717 | 70.358 | 169.8592 | 494.7394
0.69962 | 1065439.7813 | 76.0926 | 183.7038 | 462.14
0.67771 | 1005349.4722 | 72.4177 | 174.8318 | 481.1092
0.6922 | 1040111.1063 | 73.3534 | 177.0908 | 485.4288
0.68648 | 1042138.808 | 60.5087 | 146.081 | 506.8341
0.68927 | 1035726.4411 | 83.1784 | 200.8104 | 485.6076
0.69245 | 1039439.2963 | 83.0927 | 200.6035 | 493.3263
0.70034 | 1078883.5409 | 75.2034 | 181.5571 | 504.8293
0.67212 | 980274.465 | 71.1989 | 171.8893 | 498.1229
0.67929 | 1022309.3315 | 59.9856 | 144.8181 | 498.3492
0.70824 | 1082349.9009 | 76.36 | 184.3493 | 462.4866
0.69248 | 1065664.243 | 73.3144 | 176.9966 | 510.0116
0.66955 | 981877.9251 | 74.9214 | 180.8763 | 487.9293
0.7082 | 1094728.8229 | 61.6121 | 148.7447 | 480.0704
0.67875 | 1024588.0522 | 77.1206 | 186.1857 | 473.7807
0.70263 | 1079963.8583 | 69.467 | 167.7082 | 470.3125
0.69827 | 1067878.5988 | 78.1317 | 188.6267 | 482.9386
0.68495 | 1046487.5716 | 76.3396 | 184.3001 | 520.3313
0.6844 | 1024008.3285 | 65.717 | 158.6548 | 470.659
0.6881 | 1035261.9652 | 71.6761 | 173.0414 | 474.0566
0.70056 | 1086623.2376 | 82.6179 | 199.4572 | 505.2669
0.66825 | 985448.674 | 73.6836 | 177.888 | 500.8804
0.67439 | 1021245.2056 | 60.3581 | 145.7173 | 513.5217
0.67352 | 997193.2476 | 70.5348 | 170.2861 | 495.4556
0.69582 | 1039036.4003 | 69.3384 | 167.3978 | 496.1262
0.67559 | 1016662.3332 | 76.8822 | 185.6102 | 500.1307
0.6706 | 999226.9122 | 74.4526 | 179.7446 | 468.5964
0.68762 | 1025568.9467 | 52.6938 | 127.214 | 474.771
0.68706 | 1027615.7171 | 74.7334 | 180.4225 | 472.6311
0.68457 | 1020209.1137 | 77.7758 | 187.7673 | 472.5498
0.67888 | 1025653.007 | 78.8467 | 190.3528 | 487.6932
0.68358 | 1028045.4098 | 59.9429 | 144.715 | 481.8221
0.67226 | 1013253.5054 | 69.9877 | 168.9653 | 486.6929
0.68978 | 1045423.9985 | 72.2033 | 174.3141 | 502.464
0.66714 | 990399.9656 | 65.2044 | 157.4173 | 492.9134
0.72038 | 1110854.1064 | 73.4635 | 177.3565 | 475.7947
0.68044 | 1013390.6974 | 59.3615 | 143.3113 | 498.3894
0.66488 | 995655.8845 | 63.8093 | 154.0494 | 465.4045
0.70098 | 1058510.401 | 77.1477 | 186.251 | 487.8427
0.69599 | 1051966.0423 | 68.3113 | 164.918 | 529.9779
0.67144 | 985452.4953 | 74.9824 | 181.0235 | 456.0992
0.67563 | 1018129.9659 | 60.0636 | 145.0064 | 466.5078
0.67719 | 1013142.1891 | 74.7538 | 180.4716 | 484.5402
0.68683 | 1050766.5243 | 72.8839 | 175.9573 | 461.4442
0.69051 | 1051384.995 | 81.1015 | 195.7964 | 504.4287
0.69717 | 1041196.1097 | 79.5492 | 192.0488 | 514.478
0.69396 | 1049406.8327 | 70.2175 | 169.52 | 514.7849
0.69361 | 1043672.4515 | 75.1839 | 181.51 | 487.5558
0.67902 | 1020505.646 | 67.9246 | 163.9844 | 506.7046
0.69492 | 1050278.3027 | 75.5175 | 182.3153 | 476.1391
0.70385 | 1078420.2441 | 63.2133 | 152.6103 | 488.0264
0.66702 | 996578.9345 | 71.1791 | 171.8414 | 502.8736
0.69254 | 1066692.2599 | 67.6516 | 163.3254 | 503.7019
0.6977 | 1050464.9722 | 83.3428 | 201.2073 | 492.9617
0.71214 | 1086221.3299 | 72.6657 | 175.4305 | 508.3257
0.68582 | 1028071.7277 | 74.9017 | 180.8287 | 507.669
0.70104 | 1060478.4705 | 70.2419 | 169.579 | 482.989
0.70764 | 1076686.489 | 68.7659 | 166.0155 | 458.0742
0.67077 | 1009852.7742 | 66.1249 | 159.6397 | 481.2796
0.66233 | 978723.9141 | 81.7979 | 197.4775 | 482.0677
0.69855 | 1055970.6987 | 80.5182 | 194.388 | 466.9838
0.67994 | 1013598.4816 | 63.5198 | 153.3503 | 495.9963
0.67867 | 1030878.1206 | 79.2724 | 191.3806 | 513.6402
0.6804 | 1038546.5774 | 72.718 | 175.5568 | 484.9232
0.70438 | 1081697.9153 | 56.668 | 136.8086 | 478.5519
0.68812 | 1026936.0076 | 72.8534 | 175.8837 | 481.4575
0.67443 | 1016459.7025 | 64.22 | 155.0409 | 500.9528
0.65855 | 967024.5737 | 64.4955 | 155.7059 | 509.0579
0.71542 | 1110276.0109 | 79.2872 | 191.4162 | 494.8163
0.68011 | 1014109.7026 | 67.391 | 162.6963 | 491.7901
0.6881 | 1031610.8658 | 82.9888 | 200.3527 | 476.6646
0.68876 | 1047625.1791 | 77.7083 | 187.6045 | 472.9493
0.67409 | 989768.1561 | 78.2091 | 188.8134 | 497.9373
0.70144 | 1054269.0986 | 75.0997 | 181.3066 | 494.3876
0.66219 | 980048.1392 | 80.3935 | 194.087 | 486.402
0.71846 | 1112574.2722 | 71.906 | 173.5964 | 510.1788
0.69585 | 1048354.8159 | 73.547 | 177.5583 | 492.5499
0.67265 | 1007194.5925 | 73.0972 | 176.4722 | 483.5721
0.68778 | 1029328.9023 | 82.8431 | 200.0009 | 468.4916
0.69852 | 1062838.7482 | 66.8296 | 161.341 | 460.3033
0.69618 | 1056433.7804 | 75.8227 | 183.0521 | 480.257
0.68626 | 1017163.0016 | 77.3529 | 186.7463 | 507.1792
0.70051 | 1067111.1237 | 70.7344 | 170.7679 | 470.1179
0.66593 | 989025.2858 | 80.6743 | 194.7649 | 508.0373
0.68867 | 1041481.7821 | 70.2229 | 169.5332 | 498.5305
0.68815 | 1046644.6149 | 66.8039 | 161.2788 | 503.0671
0.68017 | 1028479.7418 | 75.5536 | 182.4026 | 481.6253
0.71127 | 1086701.7115 | 72.7873 | 175.724 | 479.0411
0.68455 | 1032014.9451 | 55.632 | 134.3075 | 483.7706
0.66186 | 955414.357 | 57.5361 | 138.9044 | 472.4683
0.69043 | 1041290.6414 | 66.2426 | 159.9237 | 464.9763
0.68161 | 1036430.6648 | 79.3553 | 191.5807 | 514.6558
0.68205 | 1006162.6784 | 81.6599 | 197.1443 | 487.6232
0.66457 | 991678.3109 | 72.8455 | 175.8645 | 497.6635
0.69061 | 1027093.0051 | 61.3592 | 148.1342 | 499.492
0.69055 | 1046849.9079 | 74.0053 | 178.6645 | 475.7091
0.68085 | 1001692.9107 | 79.9176 | 192.9381 | 474.9617
0.69745 | 1070464.0055 | 80.0996 | 193.3776 | 486.6094
0.66792 | 990136.589 | 66.9992 | 161.7504 | 490.1198
0.67264 | 1002030.4042 | 63.4765 | 153.2459 | 501.4906
0.67602 | 1007533.5267 | 82.4994 | 199.1712 | 507.4967
0.68647 | 1020617.7022 | 80.7599 | 194.9717 | 500.7535
0.6832 | 1012262.8094 | 59.0558 | 142.5733 | 471.8459
0.68096 | 1014318.1801 | 79.3398 | 191.5431 | 506.4407
0.68725 | 1025073.5856 | 61.5374 | 148.5644 | 474.2863
0.69577 | 1059054.5825 | 60.6691 | 146.4682 | 509.3056
0.67628 | 1016990.3732 | 78.4813 | 189.4707 | 487.6192
0.71032 | 1077391.7722 | 74.0446 | 178.7595 | 493.2553
0.70656 | 1072220.2363 | 72.2959 | 174.5377 | 463.2725
0.66457 | 990174.4945 | 82.4751 | 199.1124 | 499.9759
0.69893 | 1080303.6623 | 83.9592 | 202.6953 | 522.9784
0.67656 | 992514.9926 | 58.4724 | 141.165 | 483.8364
0.70905 | 1072643.0414 | 75.588 | 182.4856 | 507.8981
0.69091 | 1043409.5146 | 78.8149 | 190.2761 | 496.5528
0.69523 | 1034642.8804 | 72.6497 | 175.3919 | 495.4299
0.70922 | 1092179.313 | 91.342 | 220.519 | 467.863
0.68672 | 1013763.9574 | 75.4198 | 182.0795 | 501.4329
0.70494 | 1088855.3479 | 63.7264 | 153.8491 | 479.748
0.7071 | 1084023.0681 | 67.3346 | 162.56 | 474.7404
0.70604 | 1089753.4799 | 82.2131 | 198.48 | 507.882
0.71939 | 1112518.8457 | 96.0646 | 231.9203 | 490.8149
0.67506 | 1009453.5968 | 61.6943 | 148.9432 | 479.7931
0.69824 | 1054891.2949 | 62.3311 | 150.4805 | 504.3858
0.69665 | 1045208.2257 | 78.3002 | 189.0335 | 487.6258
0.69138 | 1043624.5467 | 60.1655 | 145.2523 | 491.267
0.69279 | 1057643.7609 | 60.8497 | 146.9041 | 485.1862
0.70976 | 1073977.4128 | 83.76 | 202.2146 | 492.7381
0.69677 | 1059195.4912 | 80.9702 | 195.4793 | 482.0292
0.69166 | 1043686.0158 | 81.9658 | 197.883 | 501.6877
0.70848 | 1077900.3028 | 77.7283 | 187.6528 | 480.588
0.69821 | 1056091.1837 | 70.2351 | 169.5624 | 498.1205
0.67385 | 1003261.1719 | 82.4143 | 198.9657 | 484.4048
0.69413 | 1061974.3974 | 69.1464 | 166.9343 | 503.675
0.68732 | 1021434.5861 | 67.1655 | 162.1517 | 505.1045
0.74053 | 1156401.56 | 86.4415 | 208.6883 | 507.0673
0.67902 | 1021950.5886 | 76.8913 | 185.6321 | 504.5859
0.71142 | 1101844.9871 | 85.7327 | 206.9771 | 473.93
0.67728 | 1022323.2569 | 64.3188 | 155.2793 | 512.2981
0.67289 | 1010945.0703 | 86.1901 | 208.0813 | 479.0749
0.6817 | 1024524.5412 | 75.0945 | 181.2941 | 469.9256
0.69216 | 1054313.6522 | 77.8203 | 187.8748 | 476.8238
0.70901 | 1087639.4037 | 84.9155 | 205.004 | 463.6329
0.67835 | 1025140.047 | 70.1723 | 169.411 | 504.2988
0.69372 | 1059692.0985 | 79.7199 | 192.461 | 490.6009
0.68973 | 1049786.0413 | 76.0495 | 183.5998 | 461.4489
0.69896 | 1070650.2134 | 81.589 | 196.9733 | 489.6341
0.69442 | 1040649.5423 | 82.9543 | 200.2695 | 523.011
0.68818 | 1049096.8802 | 77.8829 | 188.0258 | 483.1464
0.70839 | 1072704.5148 | 60.3566 | 145.7138 | 465.8686
0.68888 | 1022267.2457 | 65.1783 | 157.3544 | 472.2527
0.66409 | 990770.0814 | 85.5899 | 206.6323 | 493.8594
0.67543 | 1004604.3059 | 86.787 | 209.5224 | 451.4403
0.66324 | 986934.6488 | 86.8264 | 209.6175 | 495.4893
0.69316 | 1062873.9778 | 58.8941 | 142.1828 | 485.7036
0.70316 | 1072205.0154 | 70.8045 | 170.9371 | 478.9536
0.71483 | 1105405.1204 | 80.7031 | 194.8345 | 509.8756
0.69515 | 1049524.378 | 78.7932 | 190.2236 | 492.2081
0.6937 | 1044660.5003 | 62.1308 | 149.9971 | 531.6868
0.68558 | 1042143.1281 | 75.9534 | 183.3677 | 481.964
0.70227 | 1080432.2329 | 65.5554 | 158.2649 | 490.2526
0.6857 | 1033774.0545 | 73.5976 | 177.6802 | 475.3538
0.69219 | 1039579.4455 | 67.8778 | 163.8716 | 533.212
0.68212 | 1026185.0476 | 59.9629 | 144.7633 | 473.4778
0.70484 | 1065222.4647 | 62.3519 | 150.5307 | 481.6438
0.68645 | 1039268.928 | 83.8061 | 202.3257 | 512.43
0.67255 | 1008720.3523 | 78.2743 | 188.9708 | 512.615
0.68531 | 1026756.6758 | 84.4564 | 203.8959 | 475.1612
0.69327 | 1037245.9972 | 71.2782 | 172.0808 | 505.7405
0.6995 | 1064352.7148 | 72.4897 | 175.0055 | 501.332
0.69322 | 1049938.6731 | 70.3283 | 169.7875 | 469.376
0.71255 | 1086496.0046 | 76.1888 | 183.9362 | 464.5756
0.6813 | 1011481.3785 | 68.9377 | 166.4304 | 499.8967
0.69899 | 1064507.6384 | 79.4793 | 191.88 | 492.8717
0.69972 | 1076336.8859 | 81.9335 | 197.8051 | 498.0459
0.70752 | 1085250.7975 | 83.3197 | 201.1514 | 483.3517
0.69984 | 1067262.5565 | 68.9239 | 166.3971 | 502.3239
0.6968 | 1057821.3861 | 74.1105 | 178.9186 | 495.3996
0.67478 | 996789.446 | 64.882 | 156.6391 | 480.5568
0.70043 | 1064138.3302 | 58.3521 | 140.8744 | 481.0599
0.66971 | 989799.7585 | 75.5082 | 182.293 | 488.6532
0.6732 | 997602.1192 | 72.3409 | 174.6465 | 492.4882
0.67221 | 1000528.2203 | 77.6776 | 187.5303 | 476.1017
0.68603 | 1017526.6684 | 61.1928 | 147.7326 | 500.6228
0.70914 | 1100136.8551 | 96.3687 | 232.6545 | 501.5062
0.66852 | 1009941.2817 | 92.4532 | 223.2018 | 480.2367
0.70942 | 1094942.582 | 73.5306 | 177.5187 | 468.9539
0.70632 | 1074825.7065 | 82.6599 | 199.5586 | 504.0468
0.69571 | 1059043.3557 | 68.7984 | 166.094 | 494.3313
0.69899 | 1063912.7044 | 79.4346 | 191.7722 | 527.7094
0.67512 | 1014421.1142 | 81.6198 | 197.0476 | 487.8909
0.67889 | 1017504.4321 | 79.8317 | 192.7308 | 512.192
0.69333 | 1043844.0525 | 80.1929 | 193.6028 | 485.2361
0.68518 | 1019018.4957 | 70.8526 | 171.0533 | 499.9229
0.6911 | 1033447.4316 | 79.6503 | 192.2929 | 499.9805
0.69329 | 1039093.391 | 76.5725 | 184.8624 | 480.7749
0.69728 | 1052697.7569 | 64.3302 | 155.307 | 508.9467
0.70121 | 1068348.8297 | 64.9204 | 156.7316 | 488.6704
0.72977 | 1123716.9594 | 80.3605 | 194.0074 | 483.9528
0.6782 | 1012954.0395 | 74.6287 | 180.1697 | 488.439
0.69716 | 1068494.0657 | 74.8193 | 180.6298 | 517.363
0.70038 | 1061429.4173 | 83.8902 | 202.5289 | 484.3162
0.69537 | 1073099.8466 | 78.8419 | 190.3413 | 482.9742
0.68592 | 1032632.8217 | 71.7212 | 173.1502 | 509.9714
0.68843 | 1019507.2859 | 70.5521 | 170.3279 | 500.8898
0.6968 | 1062505.5024 | 76.1899 | 183.9387 | 482.2131
0.68365 | 1046106.5267 | 79.8578 | 192.7938 | 494.7526
0.69459 | 1061673.6712 | 76.3724 | 184.3793 | 489.2784
0.67819 | 1012864.4275 | 72.9082 | 176.016 | 475.8844
0.68604 | 1028312.5341 | 78.2259 | 188.854 | 472.4184
0.71685 | 1112659.3394 | 79.2987 | 191.444 | 492.4701
0.68216 | 1026920.3038 | 80.1842 | 193.5818 | 498.2852
0.68615 | 1040594.701 | 60.4475 | 145.9332 | 478.5779
0.69777 | 1074931.5715 | 80.3972 | 194.0961 | 482.3241
0.70399 | 1075192.4263 | 74.5575 | 179.9978 | 500.2947
0.68173 | 1015499.6137 | 61.4569 | 148.3702 | 488.7982
0.68031 | 1013602.4772 | 76.1187 | 183.7668 | 494.9407
0.6886 | 1035747.2119 | 78.4978 | 189.5105 | 475.3833
0.69149 | 1066569.5743 | 75.2964 | 181.7815 | 489.1169
0.70813 | 1064159.6148 | 80.0454 | 193.2466 | 485.4196
0.70069 | 1069175.1903 | 68.9635 | 166.4927 | 479.8677
0.69216 | 1056230.1298 | 67.0248 | 161.8122 | 486.7207
0.67704 | 1005744.9773 | 68.9783 | 166.5285 | 499.9445
0.71738 | 1097432.4562 | 72.8796 | 175.9469 | 510.6594
0.6946 | 1037478.4449 | 76.309 | 184.2262 | 513.2821
0.6909 | 1056949.4635 | 83.0992 | 200.6192 | 488.7393
0.68485 | 1028520.508 | 78.3778 | 189.2208 | 489.7787
0.70392 | 1077869.0779 | 67.2553 | 162.3688 | 487.3907
0.69269 | 1050028.9318 | 70.3878 | 169.9312 | 479.4648
0.67719 | 1017434.8946 | 64.0198 | 154.5575 | 487.2015
0.68652 | 1037130.2816 | 77.1813 | 186.3321 | 505.8179
0.69373 | 1061337.0888 | 69.1443 | 166.9291 | 472.062
0.70284 | 1071292.7804 | 74.4088 | 179.6387 | 491.5887
0.67401 | 996777.3501 | 78.7724 | 190.1735 | 472.9712
0.70991 | 1087221.7185 | 78.2424 | 188.894 | 516.1417
0.6768 | 987258.6147 | 65.9279 | 159.1641 | 482.6233
0.69009 | 1035147.6254 | 69.6518 | 168.1543 | 499.4077
0.69307 | 1045320.8903 | 77.0551 | 186.0275 | 490.869
0.67631 | 1007396.5122 | 72.7155 | 175.5508 | 490.7884
0.67226 | 1008735.1473 | 61.9075 | 149.458 | 487.526
0.69034 | 1038371.8871 | 81.7052 | 197.2538 | 501.8897
0.68291 | 1028652.6916 | 73.5321 | 177.5222 | 458.5392
0.70537 | 1082099.3977 | 74.8902 | 180.801 | 472.7837
0.69448 | 1048161.7391 | 66.2903 | 160.0389 | 496.0734
0.70125 | 1063159.1798 | 70.3989 | 169.9579 | 491.693
0.67342 | 996426.7788 | 68.706 | 165.871 | 497.5861
0.67359 | 1001395.5823 | 67.1897 | 162.2104 | 514.2864
0.7052 | 1081230.9827 | 76.6091 | 184.9508 | 477.7302
0.7048 | 1084344.7057 | 73.2959 | 176.952 | 486.6517
0.69025 | 1049821.2542 | 84.1904 | 203.2536 | 485.9101
0.67475 | 987164.0466 | 82.4703 | 199.1009 | 482.0132
0.69965 | 1053468.2642 | 66.1339 | 159.6614 | 484.4892
0.66143 | 966453.4367 | 60.2833 | 145.5369 | 507.0366
0.70194 | 1093354.6512 | 82.9659 | 200.2974 | 521.8584
0.70459 | 1061823.99 | 93.613 | 226.0019 | 499.2419
0.70571 | 1082039.3405 | 83.2868 | 201.0721 | 483.1106
0.69145 | 1036608.8839 | 55.3215 | 133.558 | 473.6655
0.6964 | 1057578.3511 | 68.6352 | 165.7001 | 486.558
0.67059 | 995839.5858 | 57.037 | 137.6994 | 489.3984
0.69831 | 1059630.1215 | 80.8255 | 195.13 | 491.9077
0.67446 | 1017924.109 | 75.4116 | 182.0596 | 477.4411
0.70216 | 1082653.3414 | 75.2706 | 181.7193 | 497.7047
0.69873 | 1047161.3255 | 73.1611 | 176.6264 | 483.0388
0.7118 | 1109697.7629 | 70.4596 | 170.1045 | 501.9738
0.67129 | 989326.4546 | 73.6385 | 177.779 | 465.3844
0.69775 | 1043219.4932 | 57.4247 | 138.6355 | 498.0091
0.67926 | 1013938.1319 | 76.2614 | 184.1113 | 468.1759
0.71118 | 1096632.9928 | 80.4739 | 194.2812 | 487.4609
0.70782 | 1091585.7559 | 78.7886 | 190.2125 | 497.3337
0.70407 | 1079804.2855 | 71.214 | 171.9259 | 499.7707
0.70568 | 1057292.3433 | 79.8043 | 192.6645 | 510.9237
0.67309 | 1008647.0324 | 76.5591 | 184.83 | 501.5044
0.6963 | 1059459.5282 | 64.8347 | 156.5248 | 493.5783
0.68235 | 1026294.3535 | 66.0617 | 159.487 | 478.3542
0.66493 | 982678.1271 | 73.8433 | 178.2734 | 486.6727
0.69355 | 1054023.6483 | 68.6858 | 165.8222 | 519.126
0.6988 | 1052734.2646 | 64.1925 | 154.9744 | 489.2569
0.67146 | 999825.965 | 63.4487 | 153.1786 | 472.8227
0.71824 | 1105694.4607 | 67.6159 | 163.2393 | 479.8238
0.68357 | 1010284.4197 | 72.1491 | 174.1833 | 507.2513
0.69281 | 1073983.8604 | 77.4571 | 186.998 | 487.5509
0.69134 | 1054771.0878 | 67.0118 | 161.7807 | 487.5326
0.67158 | 986940.073 | 75.0801 | 181.2595 | 524.8504
0.68032 | 1036262.1851 | 68.842 | 166.1992 | 490.6632
0.67682 | 1030325.4654 | 70.6871 | 170.6538 | 487.7295
0.69361 | 1067539.8311 | 67.6003 | 163.2015 | 479.8418
0.69043 | 1038915.1621 | 80.1492 | 193.4972 | 509.262
0.68117 | 1017899.8261 | 85.7279 | 206.9655 | 485.8298
0.68957 | 1044534.0692 | 73.9468 | 178.5235 | 457.6259
0.71779 | 1103381.952 | 59.0408 | 142.5371 | 479.5281
0.67699 | 1017819.9933 | 67.9491 | 164.0436 | 494.1505
0.68401 | 1028602.0904 | 64.0772 | 154.6961 | 470.9932
0.68149 | 1028758.219 | 81.9995 | 197.9642 | 452.6788
0.71143 | 1090086.9031 | 63.2165 | 152.618 | 482.7916
0.69024 | 1043438.6014 | 88.2076 | 212.9519 | 468.2985
0.72354 | 1113391.1841 | 76.8885 | 185.6253 | 490.2853
0.68714 | 1043498.7593 | 74.1341 | 178.9756 | 491.1658
0.69751 | 1053375.4446 | 82.0331 | 198.0454 | 499.5933
0.68331 | 1021178.019 | 57.3994 | 138.5743 | 496.2792
0.67816 | 1009416.4575 | 77.6804 | 187.5372 | 500.0738
0.70144 | 1083852.3387 | 65.8411 | 158.9545 | 509.9583
0.6753 | 995600.5627 | 80.0845 | 193.3411 | 500.6883
0.68515 | 1025758.1643 | 73.0858 | 176.4447 | 484.5128
0.68358 | 1016714.6826 | 59.2825 | 143.1207 | 480.8845
0.67439 | 1013230.0298 | 80.0268 | 193.2018 | 522.1073
0.70291 | 1074187.7558 | 58.6953 | 141.7029 | 504.31
0.67571 | 996201.9863 | 62.8206 | 151.6622 | 508.9359
0.69727 | 1065721.4692 | 81.4104 | 196.5421 | 498.4753
0.69644 | 1041943.7379 | 71.2752 | 172.0736 | 484.2056
0.68932 | 1056100.0216 | 83.2223 | 200.9165 | 526.6783
0.66324 | 969091.4428 | 73.008 | 176.2569 | 516.3236
0.70393 | 1071579.0494 | 76.0625 | 183.6312 | 497.6857
0.66276 | 981180.5044 | 66.9098 | 161.5345 | 503.2246
0.66688 | 992771.6232 | 63.4332 | 153.1412 | 466.6954
0.71116 | 1085008.8656 | 87.5265 | 211.3077 | 509.7392
0.70102 | 1063616.1619 | 62.5969 | 151.1222 | 498.0277
0.70088 | 1062873.245 | 74.0301 | 178.7244 | 500.1605
0.68048 | 1018500.5942 | 67.6458 | 163.3114 | 482.9694
0.67223 | 999134.078 | 67.1743 | 162.173 | 480.5357
0.69872 | 1064754.4568 | 76.1416 | 183.822 | 467.0521
0.7179 | 1088398.7238 | 80.7543 | 194.9581 | 508.778
0.67801 | 1020609.4382 | 78.5595 | 189.6594 | 478.4529
0.68984 | 1036817.4106 | 81.6421 | 197.1015 | 490.0302
0.70539 | 1091467.7631 | 69.932 | 168.8307 | 489.9571
0.68209 | 1023970.6114 | 71.519 | 172.6622 | 518.8853
0.69113 | 1039305.058 | 53.1286 | 128.2639 | 495.1532
0.70475 | 1080688.4326 | 80.0276 | 193.2037 | 500.4727
0.69851 | 1058245.8122 | 70.3478 | 169.8346 | 498.8777
0.66629 | 981816.6534 | 73.6277 | 177.753 | 472.8698
0.72645 | 1128312.7546 | 68.2188 | 164.6948 | 473.8458
0.70428 | 1075599.404 | 81.3481 | 196.3917 | 478.4573
0.70868 | 1092540.0048 | 82.1155 | 198.2444 | 485.3673
0.69524 | 1049518.974 | 82.0002 | 197.9661 | 485.843
0.68787 | 1034915.6435 | 59.9673 | 144.7738 | 487.8878
0.69226 | 1044481.942 | 92.3356 | 222.9178 | 509.6148
0.68907 | 1041917.1312 | 68.3387 | 164.9843 | 508.0009
0.68357 | 1035585.3651 | 70.3462 | 169.8308 | 482.3386
0.69675 | 1060253.1184 | 68.7747 | 166.0367 | 484.9502
0.65652 | 958079.0168 | 65.9552 | 159.2299 | 477.2647
0.69333 | 1037195.5731 | 73.0291 | 176.3077 | 458.4955
0.70008 | 1065308.839 | 68.7736 | 166.0342 | 487.4714
0.68054 | 1023458.6381 | 71.6467 | 172.9704 | 500.1728
0.68297 | 1031005.4675 | 77.1243 | 186.1946 | 510.5497
0.70245 | 1077177.8531 | 81.6792 | 197.1911 | 472.6848
0.68921 | 1040182.2326 | 77.1061 | 186.1506 | 489.7945
0.69752 | 1050631.9323 | 75.2609 | 181.6958 | 494.157
0.70601 | 1084577.5824 | 65.4587 | 158.0312 | 479.9522
0.68319 | 1022733.6561 | 79.7375 | 192.5034 | 512.9883
0.69198 | 1039016.7174 | 57.6708 | 139.2297 | 513.5836
0.68259 | 1023030.5803 | 67.7371 | 163.5319 | 490.8326
0.69243 | 1055436.9813 | 60.7539 | 146.673 | 497.8021
0.68506 | 1014187.5576 | 75.6348 | 182.5985 | 471.598
0.69175 | 1036180.2192 | 69.5544 | 167.9192 | 494.5256
0.69667 | 1027504.4792 | 75.3509 | 181.9132 | 494.4061
0.68501 | 1035259.1567 | 70.176 | 169.4198 | 470.5526
0.69458 | 1041242.4686 | 71.4183 | 172.4191 | 485.2941
0.65985 | 976846.6661 | 70.528 | 170.2695 | 491.3821
0.69794 | 1058402.2134 | 60.4433 | 145.923 | 506.8526
0.7166 | 1113291.9969 | 74.0133 | 178.684 | 488.3817
0.70722 | 1091135.2564 | 78.8232 | 190.296 | 508.3632
0.67576 | 1004876.0581 | 72.5932 | 175.2555 | 509.3993
0.70121 | 1081549.3248 | 63.6358 | 153.6303 | 527.7569
0.68945 | 1040492.8767 | 77.1017 | 186.1401 | 514.8981
0.70233 | 1073946.2362 | 71.0034 | 171.4174 | 520.5855
0.69634 | 1044054.7778 | 55.3282 | 133.574 | 500.0217
0.6665 | 998780.6138 | 63.854 | 154.1572 | 503.5248
0.67329 | 1007450.5353 | 73.046 | 176.3488 | 483.1567
0.66837 | 977862.6203 | 67.9362 | 164.0126 | 479.2155
0.70589 | 1076975.7967 | 77.9467 | 188.1799 | 502.1442
0.68641 | 1029224.9824 | 67.7673 | 163.6047 | 490.7518
0.70413 | 1076651.9765 | 79.7993 | 192.6526 | 464.931
0.68871 | 1045789.9085 | 68.6288 | 165.6845 | 509.1153
0.69448 | 1033943.3851 | 74.3899 | 179.5931 | 507.2464
0.68 | 1030915.3977 | 80.7524 | 194.9536 | 478.2429
0.68658 | 1041833.3838 | 72.2908 | 174.5255 | 495.5968
0.69524 | 1059621.3541 | 60.7485 | 146.6599 | 478.9335
0.70093 | 1066067.946 | 74.2481 | 179.2508 | 483.8744
0.69175 | 1049938.1052 | 94.2842 | 227.6222 | 482.3038
0.70554 | 1065741.7139 | 88.1385 | 212.7852 | 491.3534
0.68201 | 1025097.4829 | 66.0172 | 159.3795 | 506.5268
0.68557 | 1038772.3045 | 79.0753 | 190.9047 | 469.3629
0.69179 | 1048166.8516 | 70.3548 | 169.8514 | 508.7565
0.69538 | 1051980.288 | 64.4616 | 155.6241 | 500.0281
0.67873 | 1034217.4783 | 77.8478 | 187.9413 | 500.737
0.67794 | 1016045.4091 | 73.1635 | 176.6322 | 492.1143
0.68554 | 1007812.2308 | 71.7665 | 173.2598 | 497.8377
0.68277 | 1031520.0678 | 80.4715 | 194.2755 | 473.5423
0.68869 | 1037272.1309 | 69.9371 | 168.8431 | 509.2385
0.70424 | 1072604.0095 | 77.8125 | 187.8559 | 477.6857
0.69864 | 1072526.0011 | 71.2837 | 172.0941 | 496.6758
0.66384 | 993792.5689 | 57.4981 | 138.8127 | 502.2802
0.70363 | 1068749.5142 | 70.5292 | 170.2726 | 460.7425
0.7032 | 1064238.1439 | 86.8225 | 209.6081 | 481.2543
0.70537 | 1074644.2404 | 74.3743 | 179.5555 | 477.8045
0.67962 | 1008408.2957 | 61.217 | 147.7909 | 486.4418
0.67145 | 1007498.3204 | 76.6587 | 185.0704 | 494.1331
0.68236 | 1033508.779 | 88.377 | 213.3609 | 486.9221
0.71063 | 1087578.0904 | 84.7165 | 204.5237 | 475.8978
0.68648 | 1034290.2733 | 75.2818 | 181.7463 | 512.0394
0.67279 | 989088.5487 | 68.2638 | 164.8035 | 522.212
0.69753 | 1059068.2444 | 75.8646 | 183.1532 | 519.5689
0.69999 | 1055937.5333 | 80.3502 | 193.9825 | 481.0008
0.69713 | 1056503.7567 | 82.3218 | 198.7424 | 453.7323
0.6726 | 988427.874 | 78.2754 | 188.9736 | 500.1053
0.67495 | 1015354.3036 | 71.6673 | 173.0201 | 486.9619
0.69621 | 1059094.5124 | 64.8206 | 156.4907 | 476.8468
0.68061 | 1021401.9076 | 74.9851 | 181.03 | 478.327
0.6905 | 1035739.1931 | 63.9149 | 154.3043 | 486.7128
0.68159 | 1016478.2558 | 59.4418 | 143.5053 | 505.6365
0.68885 | 1039776.8002 | 73.6863 | 177.8945 | 524.8478
0.6726 | 1015590.6874 | 77.1421 | 186.2376 | 522.7763
0.69998 | 1066604.4734 | 82.7705 | 199.8258 | 481.682
0.70653 | 1097557.309 | 82.7441 | 199.762 | 542.1587
0.68997 | 1038036.7327 | 81.7218 | 197.2938 | 490.8505
0.68873 | 1038766.9939 | 78.7124 | 190.0286 | 477.379
0.69667 | 1069358.6904 | 78.4197 | 189.322 | 487.0713
0.67863 | 1019972.3428 | 53.1007 | 128.1964 | 498.3337
0.7014 | 1062486.3551 | 75.6893 | 182.7301 | 503.3411
0.68909 | 1046358.4399 | 64.7026 | 156.2058 | 475.2089
0.68494 | 1020980.6325 | 59.4136 | 143.4372 | 477.2217
0.68318 | 1014811.4192 | 63.2951 | 152.8078 | 510.7712
0.6986 | 1064149.6675 | 81.1358 | 195.8792 | 480.561
0.70908 | 1079586.9906 | 79.4581 | 191.8289 | 487.0851
0.6647 | 990727.5579 | 85.507 | 206.4322 | 493.0729
0.65729 | 964290.5336 | 87.0183 | 210.0808 | 469.1058
0.68123 | 1017629.7892 | 82.6898 | 199.6308 | 511.0864
0.6922 | 1054913.6687 | 64.9384 | 156.7752 | 468.8706
0.66884 | 995669.8019 | 80.8632 | 195.221 | 491.5852
0.70046 | 1071159.0916 | 83.0668 | 200.541 | 484.0174
0.70151 | 1066781.1078 | 82.6033 | 199.422 | 486.6983
0.67053 | 981744.6162 | 62.3877 | 150.6173 | 457.3637
0.65864 | 968457.6951 | 62.6544 | 151.2612 | 512.5886
0.69609 | 1056944.2545 | 85.5927 | 206.639 | 481.3968
0.70094 | 1063150.4046 | 66.6191 | 160.8327 | 485.9448
0.68964 | 1036935.0538 | 69.8175 | 168.5543 | 480.6212
0.70353 | 1072909.1199 | 72.6532 | 175.4004 | 494.4653
0.71203 | 1083978.4009 | 64.8696 | 156.6091 | 478.381
0.7031 | 1090619.8479 | 57.2401 | 138.1899 | 483.0523
0.68245 | 1027733.8639 | 92.1613 | 222.497 | 459.5258
0.68817 | 1041226.8851 | 82.1876 | 198.4185 | 491.4637
0.6724 | 1000993.4247 | 82.4058 | 198.9451 | 477.0586
0.66731 | 984405.7399 | 66.6717 | 160.9598 | 490.5374
0.67948 | 1021345.1452 | 66.0201 | 159.3866 | 485.9998
0.70277 | 1073569.6789 | 72.7774 | 175.7003 | 467.2352
0.71633 | 1106442.2861 | 73.5856 | 177.6513 | 499.5486
0.69461 | 1049272.0745 | 50.8855 | 122.8485 | 474.1226
0.69402 | 1056009.662 | 79.4094 | 191.7113 | 496.5449
0.69694 | 1076680.5145 | 95.9647 | 231.6794 | 500.8963
0.68409 | 1018661.2733 | 72.6922 | 175.4944 | 471.2092
0.67802 | 1020413.4322 | 71.6989 | 173.0966 | 501.4277
0.68741 | 1024970.6181 | 71.0347 | 171.493 | 476.6075
0.65748 | 966439.0023 | 60.6446 | 146.4089 | 488.0335
0.70679 | 1085464.5828 | 69.4095 | 167.5694 | 487.4004
0.69281 | 1048084.7746 | 73.8507 | 178.2913 | 492.1474
0.70647 | 1083089.6975 | 73.0382 | 176.3298 | 468.5502
0.70245 | 1068909.9901 | 74.2844 | 179.3384 | 487.7769
0.68568 | 1025666.6138 | 81.2533 | 196.1627 | 486.2859
0.69151 | 1052653.2688 | 66.8609 | 161.4166 | 487.6124
0.70412 | 1070190.2677 | 58.3763 | 140.9328 | 502.4063
0.70986 | 1082981.7151 | 58.5962 | 141.4638 | 484.1928
0.68616 | 1037445.9254 | 64.4255 | 155.5368 | 504.1408
0.69936 | 1070519.2122 | 76.4836 | 184.6477 | 507.3572
0.69216 | 1033154.7229 | 79.5061 | 191.9447 | 484.1419
0.68398 | 1037269.6175 | 83.946 | 202.6637 | 470.4374
0.71087 | 1114137.794 | 78.3114 | 189.0603 | 493.2835
0.70302 | 1057635.0306 | 71.671 | 173.0292 | 476.6512
0.6855 | 1037608.0396 | 64.4979 | 155.7117 | 538.7437
0.6928 | 1050990.5674 | 72.2721 | 174.4803 | 486.3304
0.66021 | 969237.5319 | 68.1687 | 164.5737 | 456.5879
0.64599 | 943962.6794 | 79.3755 | 191.6293 | 486.3448
0.68821 | 1039271.1779 | 73.7193 | 177.9741 | 490.2246
0.69352 | 1052210.2594 | 66.6388 | 160.8803 | 502.3545
0.68307 | 1019859.2087 | 67.4796 | 162.9101 | 497.9267
0.70549 | 1093280.1687 | 66.032 | 159.4154 | 507.2728
0.68019 | 1003337.6928 | 83.1517 | 200.7459 | 485.236
0.6845 | 1036969.4877 | 60.3826 | 145.7764 | 515.611
0.67888 | 1006798.2082 | 52.396 | 126.4953 | 500.6584
0.67612 | 1006568.4772 | 65.3885 | 157.8617 | 488.8787
0.6869 | 1024513.0177 | 79.4444 | 191.7957 | 477.4693
0.75423 | 1198179.0957 | 77.5522 | 187.2275 | 471.3229
0.68286 | 1021212.9257 | 74.724 | 180.3998 | 504.8688
0.67767 | 1020589.3183 | 69.9311 | 168.8287 | 501.6429
0.69559 | 1050856.0403 | 60.2155 | 145.373 | 504.488
0.72553 | 1115493.5025 | 83.5161 | 201.6257 | 478.5868
0.70712 | 1089767.5701 | 67.6901 | 163.4184 | 459.5793
0.69575 | 1050232.3724 | 57.9774 | 139.9697 | 501.0268
0.70108 | 1061253.2709 | 75.6355 | 182.6003 | 465.9633
0.67671 | 1008749.1995 | 61.5009 | 148.4764 | 485.1724
0.68209 | 1000832.3787 | 60.5369 | 146.149 | 492.5271
0.68369 | 1020649.3529 | 79.5167 | 191.9703 | 494.6584
0.68988 | 1053709.1428 | 74.562 | 180.0086 | 476.0096
0.69382 | 1055330.6608 | 68.744 | 165.9627 | 516.6281
0.70401 | 1076658.1105 | 71.0129 | 171.4404 | 496.6518
0.7125 | 1097367.1103 | 80.3786 | 194.0511 | 487.7134
0.68131 | 1029718.627 | 60.2407 | 145.4339 | 476.0007
0.70277 | 1067960.7844 | 64.7524 | 156.3261 | 521.7023
0.69952 | 1062452.4662 | 75.8898 | 183.2143 | 498.7441
0.70954 | 1068949.5094 | 75.2765 | 181.7335 | 459.8591
0.70461 | 1077330.9231 | 89.2155 | 215.3852 | 468.3384
0.69027 | 1046939.7087 | 70.4275 | 170.0271 | 490.2657
0.67932 | 1010166.8125 | 59.2701 | 143.0906 | 469.0721
0.71109 | 1098401.2291 | 75.4068 | 182.0482 | 477.5223
0.69061 | 1040016.592 | 75.2463 | 181.6607 | 474.217
0.67168 | 999598.3635 | 63.4135 | 153.0938 | 489.6943
0.68757 | 1036456.0272 | 73.5877 | 177.6565 | 507.6405
0.67988 | 1009318.592 | 65.2045 | 157.4176 | 511.8926
0.70282 | 1070974.1556 | 82.7737 | 199.8335 | 496.4956
0.70653 | 1082250.8018 | 81.5902 | 196.9763 | 488.4695
0.68859 | 1036770.2847 | 71.7302 | 173.172 | 484.9433
0.69016 | 1057322.4797 | 89.3912 | 215.8094 | 508.1843
0.677 | 1007695.9443 | 88.2878 | 213.1456 | 495.3895
0.68461 | 1040581.0873 | 64.7666 | 156.3605 | 497.8914
0.7007 | 1058481.1455 | 70.1433 | 169.3409 | 490.7608
0.68543 | 1033469.5414 | 75.3375 | 181.8808 | 536.441
0.68066 | 1025455.0878 | 62.1564 | 150.0589 | 480.8964
0.70407 | 1062611.1678 | 70.08 | 169.188 | 501.0196
0.69631 | 1062498.0251 | 68.9637 | 166.4932 | 460.0138
0.70829 | 1068024.1097 | 86.5056 | 208.8431 | 461.1586
0.68932 | 1033860.1404 | 71.4526 | 172.5018 | 515.0906
0.71408 | 1102747.4444 | 80.5939 | 194.571 | 497.4834
0.69044 | 1058266.1548 | 73.0204 | 176.2869 | 489.5466
0.68687 | 1038716.5556 | 64.4378 | 155.5666 | 520.0268
0.69166 | 1039413.1978 | 81.6301 | 197.0725 | 500.961
0.69354 | 1060333.7523 | 46.0642 | 111.2088 | 513.4232
0.6835 | 1010805.2248 | 58.9448 | 142.3053 | 506.0087
0.71085 | 1081783.1825 | 70.6645 | 170.5993 | 491.9224
0.67986 | 999649.9853 | 75.1202 | 181.3562 | 505.2054
0.67775 | 1014163.5188 | 69.4827 | 167.7462 | 492.1777
0.68902 | 1034009.7907 | 59.815 | 144.4061 | 491.7202
0.67303 | 999843.0275 | 68.9818 | 166.5368 | 470.1233
0.71372 | 1088874.2534 | 68.9516 | 166.464 | 505.2803
0.70964 | 1102732.7847 | 52.3459 | 126.3742 | 476.0818
0.66718 | 994315.4895 | 72.7532 | 175.6418 | 509.863
0.69577 | 1041708.2882 | 63.7965 | 154.0185 | 469.1051
0.70746 | 1075263.9821 | 68.692 | 165.8372 | 479.0013
0.6865 | 1045475.4306 | 77.8041 | 187.8357 | 495.4934
0.67202 | 991848.8571 | 77.0258 | 185.9566 | 491.6562
0.72514 | 1118401.1663 | 69.7066 | 168.2867 | 494.5921
0.67271 | 1000132.7592 | 69.0349 | 166.665 | 462.25
0.69593 | 1051753.6021 | 68.3033 | 164.8989 | 498.9171
0.70375 | 1063797.1674 | 86.7192 | 209.3586 | 471.7394
0.6568 | 974393.8841 | 65.1606 | 157.3116 | 493.5058
0.70634 | 1064906.1571 | 66.2187 | 159.8661 | 521.1187
0.6878 | 1034061.4305 | 71.6244 | 172.9165 | 492.83
0.67867 | 1020717.781 | 66.059 | 159.4806 | 474.498
0.70262 | 1085479.0306 | 65.8292 | 158.9258 | 513.075
0.69298 | 1054328.6836 | 62.7704 | 151.5412 | 503.0894
0.69081 | 1051891.9045 | 76.083 | 183.6805 | 493.9604
0.72138 | 1128390.1545 | 78.1581 | 188.6903 | 504.195
0.70126 | 1070041.1914 | 77.957 | 188.2048 | 485.5091
0.67478 | 995722.5739 | 77.0109 | 185.9208 | 474.7954
0.70435 | 1073489.7531 | 70.7696 | 170.8529 | 496.334
0.69883 | 1052955.2992 | 84.9363 | 205.0544 | 553.5308
0.68782 | 1037021.2682 | 78.6839 | 189.9598 | 481.5626
0.68559 | 1022205.767 | 79.4171 | 191.7298 | 479.1104
0.7034 | 1073577.4719 | 79.6537 | 192.3011 | 516.5799
0.69381 | 1048360.6174 | 70.1626 | 169.3875 | 503.7695
0.69876 | 1066232.8137 | 68.9638 | 166.4934 | 487.6513
0.70339 | 1073513.096 | 72.7095 | 175.5362 | 504.4062
0.69413 | 1051019.6706 | 62.4697 | 150.8151 | 511.7195
0.67699 | 996466.9145 | 70.1221 | 169.2898 | 479.8684
0.72572 | 1125814.7339 | 59.5829 | 143.8459 | 497.4381
0.71471 | 1073458.3537 | 50.5943 | 122.1453 | 504.0668
0.66427 | 982336.0817 | 83.3625 | 201.2548 | 491.972
0.68128 | 1016538.9631 | 76.2306 | 184.0369 | 479.3818
0.68507 | 1024651.4746 | 74.7341 | 180.424 | 476.9258
0.67794 | 1012695.7874 | 69.3623 | 167.4553 | 490.8406
0.69015 | 1041709.3434 | 75.4191 | 182.0779 | 488.129
0.66912 | 990708.2101 | 81.9586 | 197.8656 | 482.9139
0.70332 | 1077871.8672 | 78.2973 | 189.0263 | 489.7853
0.68335 | 1022998.8513 | 74.8011 | 180.5859 | 494.3218
0.69347 | 1039439.9208 | 71.4081 | 172.3944 | 467.227
0.6918 | 1033862.899 | 69.394 | 167.5318 | 471.0065
0.69515 | 1047425.1507 | 76.9799 | 185.8459 | 498.6835
0.67191 | 1010329.942 | 89.1894 | 215.3223 | 493.7567
0.70593 | 1081136.0905 | 72.9619 | 176.1457 | 465.409
0.69279 | 1052442.4625 | 72.3731 | 174.724 | 510.2022
0.68949 | 1035986.9496 | 71.5822 | 172.8148 | 472.3332
0.68904 | 1036511.8282 | 71.6652 | 173.015 | 506.6437
0.67006 | 984278.0877 | 88.5235 | 213.7145 | 511.1932
0.68474 | 1032648.8246 | 77.0474 | 186.009 | 478.8171
0.68448 | 1024355.8165 | 69.4912 | 167.7667 | 476.0679
0.67662 | 988477.7112 | 67.8363 | 163.7713 | 506.5605
0.7033 | 1081072.0861 | 73.2307 | 176.7945 | 511.0218
0.70483 | 1069125.4281 | 58.2595 | 140.6508 | 454.8335
0.69853 | 1054776.1649 | 73.7138 | 177.9609 | 484.8006
0.67841 | 1020535.0844 | 80.1268 | 193.4432 | 527.2541
0.68557 | 1039926.8612 | 77.4966 | 187.0933 | 518.4358
0.70917 | 1090189.7879 | 67.5206 | 163.009 | 477.562
0.66851 | 996603.3227 | 77.8012 | 187.8287 | 512.3284
0.67158 | 989843.5548 | 76.9204 | 185.7023 | 469.8825
0.67823 | 1007678.3888 | 67.1494 | 162.1129 | 471.4759
0.67786 | 1022992.6013 | 75.4064 | 182.0471 | 471.3213
0.70623 | 1063645.771 | 75.2532 | 181.6773 | 495.7663
0.6631 | 978826.4697 | 80.1835 | 193.5801 | 513.8215
0.6917 | 1043670.6435 | 64.293 | 155.2169 | 499.7924
0.71905 | 1116888.3519 | 68.2238 | 164.7068 | 487.4019
0.69705 | 1059196.9447 | 74.1801 | 179.0866 | 472.1612
0.69885 | 1030715.8359 | 58.7857 | 141.9213 | 502.5767
0.68241 | 1023802.8198 | 71.4739 | 172.5532 | 490.3375
0.68576 | 1028620.5205 | 79.896 | 192.8859 | 501.5344
0.68858 | 1039060.4924 | 64.2989 | 155.2314 | 462.644
0.6846 | 1031980.3687 | 71.8144 | 173.3752 | 475.695
0.68067 | 1035161.0118 | 91.6484 | 221.2589 | 479.307
0.70103 | 1065067.3173 | 74.1676 | 179.0564 | 501.6272
0.66692 | 982382.6642 | 71.9087 | 173.603 | 489.1264
0.67648 | 1008575.8981 | 74.4953 | 179.8476 | 520.6358
0.69016 | 1050866.0208 | 66.8772 | 161.4559 | 491.8114
0.68681 | 1018519.3419 | 63.19 | 152.5542 | 502.5988
0.69011 | 1031763.2435 | 76.382 | 184.4024 | 493.0785
0.66744 | 1001082.0339 | 71.4551 | 172.5078 | 517.4949
0.6631 | 977645.0007 | 66.6347 | 160.8704 | 506.5824
0.69531 | 1057438.686 | 58.4116 | 141.018 | 499.951
0.69633 | 1079916.2161 | 84.243 | 203.3806 | 499.1232
0.67243 | 1000117.5901 | 57.1252 | 137.9125 | 474.365
0.67731 | 1004815.9348 | 77.4234 | 186.9165 | 481.6972
0.71177 | 1073993.3648 | 70.0647 | 169.1511 | 477.7138
0.68917 | 1036901.1523 | 75.1775 | 181.4945 | 506.7218
0.68249 | 1022353.4276 | 67.7014 | 163.4458 | 524.3346
0.68091 | 1025917.5009 | 64.2007 | 154.9941 | 476.5304
0.70794 | 1095020.2196 | 75.5089 | 182.2946 | 485.1618
0.69399 | 1055990.9896 | 77.7391 | 187.6787 | 491.1859
0.67279 | 1007709.3716 | 65.7863 | 158.8223 | 478.5213
0.68745 | 1036275.0816 | 71.8149 | 173.3764 | 496.5757
0.68155 | 1005886.8506 | 70.3124 | 169.7491 | 511.69
0.70397 | 1073428.8959 | 66.9738 | 161.689 | 488.2399
0.7094 | 1087299.4782 | 76.5835 | 184.8889 | 487.1002
0.69407 | 1052036.5387 | 64.5865 | 155.9257 | 487.0122
0.68755 | 1045055.0378 | 74.2012 | 179.1375 | 502.7603
0.68199 | 1020298.1844 | 63.7483 | 153.9021 | 490.3785
0.69455 | 1051934.4052 | 73.9364 | 178.4981 | 504.4277
0.68313 | 1047816.1763 | 71.2225 | 171.9463 | 479.3847
0.71309 | 1090122.3336 | 72.8299 | 175.8269 | 482.8849
0.70413 | 1079266.4441 | 73.0224 | 176.2916 | 510.1098
0.68815 | 1030879.848 | 82.9242 | 200.1968 | 502.9245
0.67632 | 1010360.2848 | 72.928 | 176.0638 | 467.6519
0.67346 | 1010995.796 | 79.9603 | 193.0413 | 504.2968
0.71816 | 1119504.1453 | 68.4679 | 165.2962 | 487.3485
0.70729 | 1087123.6683 | 71.371 | 172.3049 | 480.4955
0.69085 | 1042807.8277 | 62.2761 | 150.3479 | 506.0017
0.70382 | 1086078.668 | 73.5147 | 177.4801 | 480.1422
0.70585 | 1088473.7309 | 78.7833 | 190.1998 | 492.0877
0.71009 | 1091607.1671 | 83.5045 | 201.5977 | 493.0075
0.65864 | 984753.8481 | 74.7056 | 180.3553 | 510.2258
0.71477 | 1096642.9014 | 76.6607 | 185.0752 | 504.7533
0.69073 | 1064639.211 | 67.6986 | 163.4389 | 456.2273
0.68769 | 1045470.8983 | 58.3906 | 140.9674 | 490.7737
0.68892 | 1026879.8241 | 79.3942 | 191.6747 | 482.079
0.71801 | 1105817.87 | 67.6454 | 163.3104 | 507.2032
0.65744 | 962955.3524 | 60.4296 | 145.8899 | 439.6427
0.67908 | 1000609.0542 | 83.0611 | 200.5272 | 500.2997
0.69211 | 1053779.2168 | 88.8407 | 214.4805 | 499.4504
0.70534 | 1079048.3058 | 76.4396 | 184.5416 | 490.7007
0.69979 | 1073741.3497 | 71.2475 | 172.0067 | 473.7841
0.6823 | 1028199.7925 | 80.2674 | 193.7826 | 510.072
0.67831 | 1009734.3018 | 65.3822 | 157.8467 | 480.1928
0.6775 | 1004365.7986 | 70.6256 | 170.5054 | 485.4536
0.68654 | 1013765.9361 | 68.5663 | 165.5336 | 501.3802
0.69369 | 1072838.2104 | 85.6097 | 206.6801 | 517.2064
0.67957 | 1020698.9383 | 84.667 | 204.4041 | 454.7756
0.70448 | 1066699.6537 | 77.3581 | 186.7591 | 476.9797
0.69364 | 1049449.3872 | 62.4209 | 150.6973 | 478.6788
0.66077 | 977896.2408 | 72.2459 | 174.4171 | 500.8349
0.68804 | 1036694.0221 | 68.0975 | 164.402 | 471.4525
0.68289 | 1010873.4812 | 81.9411 | 197.8233 | 495.1887
0.71495 | 1111794.1559 | 66.2622 | 159.9712 | 505.9919
0.69242 | 1034487.4639 | 59.5486 | 143.763 | 494.7504
0.71703 | 1118713.4572 | 79.7102 | 192.4374 | 499.0846
0.69594 | 1059777.7376 | 72.5475 | 175.1452 | 510.2513
0.69354 | 1049512.6849 | 82.2008 | 198.4502 | 478.7164
0.69829 | 1061061.5387 | 62.6915 | 151.3506 | 480.2385
0.6825 | 1019647.6251 | 79.5773 | 192.1165 | 492.7593
0.6927 | 1053606.8122 | 58.4196 | 141.0375 | 487.2514
0.68355 | 1014181.8148 | 56.9866 | 137.5777 | 498.3794
0.69162 | 1040009.2834 | 71.6385 | 172.9507 | 482.1107
0.68555 | 1043507.4362 | 66.8561 | 161.4049 | 469.7869
0.67875 | 1033306.5887 | 76.0666 | 183.641 | 490.7994
0.68715 | 1032256.099 | 57.6979 | 139.2951 | 517.8499
0.68962 | 1041383.7239 | 57.9995 | 140.0233 | 468.7346
0.68802 | 1026601.4436 | 77.8708 | 187.9968 | 479.2013
0.69505 | 1046663.7323 | 66.1411 | 159.6787 | 488.9524
0.68082 | 1000199.3976 | 75.0558 | 181.2007 | 513.8032
0.69181 | 1031018.21 | 77.777 | 187.7702 | 476.1021
0.68484 | 1024317.5022 | 59.6157 | 143.9251 | 498.8163
0.71286 | 1108430.9553 | 81.148 | 195.9086 | 482.2585
0.7259 | 1132118.0476 | 70.4874 | 170.1717 | 502.6026
0.67636 | 1019858.2245 | 71.8352 | 173.4255 | 489.821
0.7141 | 1096833.1165 | 63.3695 | 152.9875 | 486.5061
0.67162 | 1005930.4675 | 81.3578 | 196.4151 | 522.6093
0.66513 | 1002740.1895 | 75.3289 | 181.86 | 493.2084
0.69583 | 1059247.1335 | 74.313 | 179.4076 | 504.6189
0.70019 | 1047427.2565 | 87.2858 | 210.7265 | 492.7657
0.69663 | 1038404.0921 | 69.2155 | 167.1009 | 493.4384
0.70123 | 1054916.7462 | 83.2743 | 201.0419 | 481.0771
0.68581 | 1023138.4528 | 77.8582 | 187.9663 | 475.6582
0.71263 | 1107046.6308 | 87.5703 | 211.4134 | 481.8188
0.69946 | 1060548.6133 | 83.9311 | 202.6276 | 498.4161
0.70026 | 1078136.7367 | 94.2218 | 227.4716 | 468.3801
0.68769 | 1029363.9131 | 63.7813 | 153.9817 | 485.5321
0.70973 | 1116338.7743 | 83.7808 | 202.2646 | 504.9744
0.65877 | 976777.6687 | 78.9772 | 190.6678 | 500.3139
0.70106 | 1063921.2831 | 77.5322 | 187.1794 | 507.7682
0.68666 | 1024164.3082 | 67.4097 | 162.7414 | 513.0761
0.69955 | 1079376.2047 | 90.0313 | 217.3549 | 510.7662
0.69367 | 1039409.6064 | 65.814 | 158.8891 | 502.8325
0.67815 | 1006114.4657 | 72.4265 | 174.853 | 503.5021
0.72127 | 1108856.1026 | 75.0496 | 181.1856 | 466.0086
0.69638 | 1059129.8675 | 64.8071 | 156.4582 | 499.7927
0.68393 | 1037452.9282 | 79.1645 | 191.1201 | 482.651
0.69438 | 1047463.1752 | 62.2363 | 150.2518 | 489.3703
0.6699 | 1000062.8858 | 76.2687 | 184.129 | 473.3903
0.69191 | 1048575.7114 | 82.3202 | 198.7385 | 465.4248
0.69971 | 1054092.6634 | 69.9513 | 168.8773 | 493.1342
0.66118 | 964947.5635 | 74.5612 | 180.0068 | 479.9166
0.68987 | 1026248.2977 | 63.3869 | 153.0296 | 513.0344
0.68287 | 1011890.7643 | 80.4916 | 194.3239 | 499.1166
0.66185 | 987585.2381 | 85.6032 | 206.6643 | 521.2462
0.68157 | 1021015.4299 | 63.8319 | 154.1037 | 509.2069
0.6925 | 1046787.3959 | 77.2267 | 186.4418 | 485.134
0.6838 | 1026336.9488 | 73.271 | 176.8918 | 487.9545
0.6787 | 1004285.543 | 72.2362 | 174.3935 | 530.3385
0.66988 | 987197.0614 | 75.2902 | 181.7667 | 490.8882
0.69921 | 1067704.7193 | 74.5445 | 179.9664 | 458.1024
0.68866 | 1021164.4501 | 74.0916 | 178.873 | 494.9176
0.70267 | 1051695.1813 | 67.6443 | 163.3078 | 490.1747
0.70105 | 1062596.487 | 70.3813 | 169.9156 | 484.4402
0.6878 | 1027991.9963 | 81.1865 | 196.0015 | 485.2445
0.6938 | 1063178.1421 | 63.2227 | 152.6331 | 475.9829
0.66399 | 968376.0464 | 72.8719 | 175.9284 | 459.2671
0.67459 | 1006229.4785 | 65.5148 | 158.1667 | 482.6848
0.71677 | 1100846.697 | 83.4269 | 201.4104 | 459.0166
0.7324 | 1152760.3186 | 78.6444 | 189.8645 | 475.7762
0.70457 | 1079049.2434 | 72.9618 | 176.1454 | 471.3467
0.65053 | 957888.0134 | 66.5495 | 160.6648 | 518.2003
0.70342 | 1064883.3823 | 72.1214 | 174.1164 | 516.3936
0.67467 | 993616.6801 | 84.4893 | 203.9753 | 500.9096
0.66856 | 1012986.7499 | 73.9664 | 178.5706 | 479.3588
0.67711 | 1019918.1525 | 71.7602 | 173.2445 | 456.1099
0.67652 | 1002259.7896 | 77.3169 | 186.6596 | 481.962
0.68838 | 1027819.0683 | 69.3313 | 167.3805 | 472.6185
0.67493 | 993523.3768 | 70.1287 | 169.3056 | 466.6625
0.6769 | 1018911.0337 | 69.8953 | 168.7421 | 479.1694
0.69869 | 1055810.7604 | 64.3898 | 155.4507 | 459.7817
0.69464 | 1044002.6177 | 76.7848 | 185.3749 | 482.1781
0.69013 | 1032553.9424 | 81.2715 | 196.2066 | 525.0069
0.70009 | 1071737.1631 | 84.7405 | 204.5816 | 459.3389
0.68199 | 1018215.9582 | 81.1001 | 195.793 | 479.9869
0.70975 | 1102715.0227 | 74.0181 | 178.6956 | 497.8014
0.69023 | 1061482.8397 | 69.5051 | 167.8001 | 477.82
0.69122 | 1045321.2189 | 75.5636 | 182.4266 | 491.4461
0.69511 | 1056176.3317 | 87.1673 | 210.4406 | 505.5156
0.64377 | 924747.8765 | 74.9666 | 180.9855 | 498.6392
0.6719 | 996842.3235 | 72.4263 | 174.8524 | 486.7263
0.67885 | 995346.4411 | 74.9081 | 180.8442 | 487.2597
0.66219 | 975276.7873 | 76.8634 | 185.5646 | 506.6682
0.69095 | 1052065.043 | 85.8313 | 207.2151 | 500.7447
0.69095 | 1032404.0521 | 77.9783 | 188.2562 | 486.0257
0.71908 | 1133720.9848 | 75.1113 | 181.3347 | 484.5486
0.69459 | 1046586.3461 | 68.0989 | 164.4052 | 491.5215
0.69312 | 1063029.9559 | 74.87 | 180.7521 | 500.129
0.6948 | 1059290.3179 | 70.793 | 170.9095 | 493.0894
0.69799 | 1048879.0641 | 84.7093 | 204.5063 | 472.8264
0.67424 | 1011999.4838 | 73.2721 | 176.8944 | 498.7499
0.69941 | 1064273.9849 | 74.2842 | 179.338 | 484.8351
0.68715 | 1038073.1717 | 70.148 | 169.3524 | 489.2797
0.66363 | 972826.8018 | 73.2468 | 176.8334 | 459.6832
0.69487 | 1051491.7303 | 73.8717 | 178.342 | 458.1814
0.70797 | 1069307.9786 | 64.3582 | 155.3744 | 507.974
0.68641 | 1036999.3139 | 82.0643 | 198.1206 | 502.7634
0.68989 | 1040085.9598 | 70.005 | 169.0069 | 510.6915
0.70169 | 1072917.9304 | 72.8444 | 175.8619 | 479.1971
0.68605 | 1035608.4414 | 66.3012 | 160.0652 | 504.0604
0.68414 | 1030060.691 | 73.5003 | 177.4454 | 487.3547
0.67001 | 996594.0402 | 75.9921 | 183.4611 | 479.2158
0.68792 | 1019295.105 | 80.4858 | 194.31 | 507.0097
0.676 | 1003386.2116 | 59.1612 | 142.8277 | 477.7071
0.69221 | 1034336.3265 | 87.1886 | 210.4918 | 502.5504
0.66785 | 963868.7257 | 73.7347 | 178.0114 | 506.2277
0.66811 | 990578.3001 | 67.0091 | 161.7743 | 483.1975
0.71097 | 1114716.0432 | 68.8645 | 166.2536 | 505.742
0.68918 | 1016625.7893 | 66.6689 | 160.953 | 497.1185
0.69416 | 1056598.7651 | 82.6812 | 199.61 | 504.4223
0.66733 | 981101.1791 | 66.4456 | 160.4138 | 469.0502
0.72342 | 1141254.0232 | 75.1565 | 181.4439 | 486.2294
0.68905 | 1017471.1668 | 75.4405 | 182.1294 | 507.6731
0.68373 | 1029539.9613 | 81.7932 | 197.4663 | 496.1461
0.69663 | 1049627.1337 | 76.9769 | 185.8386 | 491.2889
0.69924 | 1061039.0918 | 79.1916 | 191.1854 | 470.858
0.68262 | 1029979.9293 | 66.2725 | 159.9961 | 484.7317
0.71966 | 1118959.077 | 84.4587 | 203.9014 | 500.0723
0.68249 | 1023822.638 | 67.7987 | 163.6806 | 519.4764
0.69142 | 1058264.0201 | 65.218 | 157.4501 | 475.0543
0.67222 | 993451.3636 | 87.6584 | 211.6262 | 481.9827
0.68874 | 1045214.7299 | 66.6547 | 160.9187 | 482.7366
0.67429 | 994423.2784 | 62.8406 | 151.7105 | 480.6151
0.70373 | 1078309.6958 | 76.5621 | 184.8372 | 500.4905
0.71122 | 1075263.4716 | 70.2018 | 169.4822 | 501.1312
0.70011 | 1070526.0076 | 65.1548 | 157.2975 | 503.8064
0.67031 | 985659.4574 | 73.4729 | 177.3792 | 500.7175
0.69105 | 1045085.9949 | 78.9257 | 190.5434 | 496.0095
0.70093 | 1061382.6453 | 80.6556 | 194.7198 | 486.9429
0.68842 | 1025666.709 | 63.4843 | 153.2647 | 467.6639
0.70113 | 1067001.2449 | 74.292 | 179.3569 | 515.0378
0.68949 | 1025300.9579 | 79.2069 | 191.2224 | 491.714
0.69521 | 1041622.3318 | 63.8423 | 154.129 | 506.2914
0.69713 | 1066603.8274 | 67.2003 | 162.2359 | 465.5548
0.68611 | 1047187.2967 | 48.7753 | 117.7541 | 499.9661
0.68804 | 1045334.1115 | 74.1674 | 179.056 | 509.4238
0.67877 | 1012860.1238 | 65.5405 | 158.2287 | 508.0867
0.70867 | 1087116.0108 | 80.0578 | 193.2767 | 502.7245
0.69885 | 1057659.2633 | 86.8233 | 209.61 | 461.9073
0.69164 | 1031228.8516 | 65.4878 | 158.1015 | 472.2469
0.70495 | 1068707.832 | 77.4513 | 186.9839 | 516.9155
0.69547 | 1047293.2754 | 66.1409 | 159.6782 | 498.6104
0.69924 | 1065553.1171 | 56.2327 | 135.7578 | 475.3036
0.69599 | 1056054.4724 | 79.188 | 191.1768 | 461.8678
0.70789 | 1077920.2243 | 72.5438 | 175.1361 | 506.3184
0.6637 | 969667.5077 | 83.8158 | 202.3491 | 504.0842
0.7188 | 1117637.754 | 75.9038 | 183.248 | 501.564
0.69901 | 1066034.4171 | 64.9835 | 156.8839 | 479.887
0.68945 | 1044873.9421 | 72.1997 | 174.3055 | 457.0847
0.69178 | 1043244.132 | 66.2374 | 159.9113 | 486.8464
0.68 | 995515.1854 | 79.5237 | 191.9873 | 507.6035
0.71089 | 1085561.0818 | 78.0156 | 188.3464 | 486.3885
0.71519 | 1099986.0915 | 71.4169 | 172.4157 | 496.6993
0.69511 | 1063288.1098 | 74.6744 | 180.2799 | 514.9721
0.69158 | 1042880.5265 | 77.0413 | 185.9942 | 474.187
0.69805 | 1050996.5484 | 80.1962 | 193.6108 | 492.8393
0.67409 | 1012621.9306 | 61.9769 | 149.6254 | 468.272
0.70225 | 1087922.9723 | 57.1676 | 138.0147 | 492.9544
0.70132 | 1062116.0088 | 64.5312 | 155.7922 | 491.2118
0.68161 | 1025330.0158 | 71.6642 | 173.0127 | 521.8355
0.69171 | 1041944.2624 | 71.762 | 173.2489 | 462.6652
0.67527 | 998587.7617 | 63.0122 | 152.1248 | 509.3935
0.68597 | 1047071.8723 | 72.7193 | 175.56 | 504.6454
0.71 | 1086958.9179 | 72.9344 | 176.0792 | 488.6557
0.67413 | 1006375.8151 | 61.5914 | 148.6948 | 503.6026
0.68102 | 1006254.6308 | 81.7901 | 197.4587 | 482.6928
0.68607 | 1028438.833 | 74.9008 | 180.8266 | 482.8905
0.65656 | 965699.3701 | 70.072 | 169.1688 | 505.6666
0.68926 | 1041841.3068 | 73.7895 | 178.1435 | 457.2907
0.68703 | 1040211.8603 | 70.3043 | 169.7296 | 483.2733
0.69793 | 1074310.044 | 83.6136 | 201.861 | 504.3782
0.70791 | 1088273.6674 | 76.8133 | 185.4438 | 508.6487
0.70123 | 1069686.0914 | 74.4675 | 179.7805 | 477.7633
0.6808 | 1002122.88 | 81.4804 | 196.711 | 481.6551
0.69618 | 1042597.516 | 81.3492 | 196.3943 | 509.8648
0.69774 | 1046451.4951 | 79.8852 | 192.8599 | 480.4083
0.68868 | 1033607.7026 | 65.9204 | 159.1459 | 510.3204
0.676 | 1017572.5493 | 64.1405 | 154.8489 | 500.2932
0.6871 | 1042236.0082 | 77.4959 | 187.0916 | 471.1792
0.67322 | 990831.7326 | 70.1163 | 169.2758 | 479.3707
0.67009 | 991228.0494 | 75.5736 | 182.4509 | 480.2451
0.7027 | 1075916.9543 | 79.9071 | 192.9128 | 507.1942
0.69472 | 1039271.8154 | 67.6111 | 163.2277 | 474.271
0.70062 | 1083910.8559 | 69.9209 | 168.8039 | 508.6636
0.68175 | 1034374.9282 | 74.0676 | 178.8149 | 458.6655
0.69786 | 1052940.7735 | 68.1918 | 164.6295 | 507.884
0.71086 | 1103535.8692 | 79.3107 | 191.473 | 491.0894
0.68517 | 1025040.3808 | 78.076 | 188.4922 | 483.3648
0.6977 | 1071387.9885 | 74.9633 | 180.9775 | 494.3012
0.6927 | 1046126.0046 | 78.8162 | 190.2791 | 474.6607
0.69935 | 1060506.2612 | 74.0278 | 178.7188 | 482.5308
0.68604 | 1017726.8827 | 75.7898 | 182.9727 | 484.3451
0.68595 | 1039939.813 | 66.5884 | 160.7586 | 476.6945
0.68798 | 1045996.4855 | 79.3469 | 191.5605 | 471.6773
0.66881 | 996671.0903 | 76.134 | 183.8038 | 521.144
0.69739 | 1059207.3679 | 68.6436 | 165.7203 | 504.3177
0.68624 | 1027122.4151 | 73.0665 | 176.3982 | 475.3703
0.69841 | 1056625.5864 | 82.1808 | 198.402 | 495.6534
0.65584 | 954406.2336 | 65.7706 | 158.7843 | 486.8888
0.67737 | 1001472.2526 | 77.1592 | 186.2788 | 523.0863
0.66741 | 981553.1711 | 84.3718 | 203.6916 | 491.3756
0.69416 | 1033831.8193 | 80.9001 | 195.3101 | 503.7835
0.67068 | 1005264.0862 | 71.4072 | 172.3923 | 493.76
0.70317 | 1085828.6065 | 77.1579 | 186.2757 | 500.1167
0.70686 | 1089871.9284 | 94.3579 | 227.8002 | 496.2633
0.69002 | 1053896.6505 | 76.3153 | 184.2414 | 489.5196
0.68368 | 1034003.9355 | 64.4445 | 155.5827 | 469.2316
0.67347 | 997204.3228 | 59.0181 | 142.4823 | 483.0756
0.69327 | 1038514.9459 | 79.7898 | 192.6296 | 486.1705
0.68854 | 1055849.6637 | 73.0556 | 176.3717 | 519.1059
0.68216 | 1011632.3258 | 67.0243 | 161.8111 | 492.0721
0.66628 | 985727.6766 | 68.6975 | 165.8503 | 515.3016
0.67689 | 1012894.5506 | 78.0941 | 188.5358 | 469.1863
0.68948 | 1020745.4398 | 70.5296 | 170.2734 | 510.5043
0.68066 | 996799.198 | 79.5486 | 192.0472 | 497.774
0.67543 | 1005630.4063 | 77.7024 | 187.5901 | 498.8467
0.67774 | 1011900.0038 | 87.119 | 210.3239 | 471.0841
0.67528 | 1004486.7724 | 82.3412 | 198.7894 | 473.6533
0.68471 | 1031570.293 | 68.0912 | 164.3866 | 504.2192
0.66238 | 978072.4188 | 75.4388 | 182.1253 | 461.1542
0.71001 | 1085939.6582 | 56.4392 | 136.2564 | 488.2315
0.67796 | 1022766.5178 | 92.3226 | 222.8865 | 470.0582
0.69356 | 1071766.9593 | 69.8408 | 168.6107 | 479.4
0.68612 | 1035270.7293 | 77.0872 | 186.1049 | 494.6677
0.70429 | 1088112.7558 | 83.9232 | 202.6084 | 485.5691
0.69811 | 1054119.4891 | 64.3399 | 155.3302 | 511.3841
0.6903 | 1055053.051 | 83.0221 | 200.4332 | 492.9715
0.68574 | 1021937.2044 | 52.6512 | 127.1111 | 508.7753
0.69321 | 1062714.2101 | 73.0345 | 176.321 | 494.2327
0.69148 | 1050073.9279 | 74.1335 | 178.9741 | 498.7157
0.68809 | 1004102.5746 | 67.76 | 163.587 | 510.4546
0.67464 | 994943.6856 | 73.6888 | 177.9005 | 518.3384
0.69047 | 1056998.8252 | 84.7394 | 204.579 | 497.3033
0.66849 | 992115.5513 | 68.9138 | 166.3727 | 500.7789
0.67695 | 1017587.8258 | 75.109 | 181.3293 | 486.4589
0.66583 | 995266.7824 | 81.196 | 196.0245 | 503.8045
0.69031 | 1034701.1084 | 78.2249 | 188.8517 | 508.3034
0.69296 | 1049623.5706 | 58.1767 | 140.451 | 491.5813
0.69333 | 1039611.0094 | 84.5242 | 204.0594 | 539.407
0.68804 | 1021234.8123 | 59.1601 | 142.825 | 479.9748
0.68814 | 1037727.8596 | 75.3504 | 181.9121 | 497.6742
0.68576 | 1038468.4871 | 75.6653 | 182.6722 | 489.8975
0.67367 | 1021030.1363 | 80.7295 | 194.8983 | 501.8611
0.69622 | 1059327.1913 | 64.8332 | 156.5212 | 493.4332
0.68628 | 1033253.3175 | 66.1287 | 159.6489 | 465.5265
0.71105 | 1082399.8976 | 79.4436 | 191.7938 | 475.6378
0.69409 | 1036101.2532 | 63.607 | 153.5609 | 476.7669
0.69626 | 1050424.146 | 71.8738 | 173.5187 | 452.9497
0.69073 | 1053725.3616 | 79.6141 | 192.2055 | 483.5376
0.70249 | 1081050.5523 | 78.6211 | 189.808 | 446.7659
0.69904 | 1059353.3722 | 66.5612 | 160.693 | 517.0171
0.68404 | 1023483.2072 | 69.4768 | 167.7319 | 476.0462
0.70338 | 1071058.4615 | 88.8502 | 214.5033 | 520.2392
0.71391 | 1105415.933 | 73.7669 | 178.089 | 481.0689
0.73182 | 1157231.0029 | 63.0283 | 152.1639 | 500.2686
0.67679 | 1023028.2569 | 66.3925 | 160.2857 | 498.5169
0.65742 | 968937.5008 | 68.4364 | 165.2202 | 527.228
0.69242 | 1068180.6008 | 101.31 | 244.5839 | 469.7603
0.71211 | 1077922.1865 | 70.2873 | 169.6886 | 472.7509
0.66273 | 978936.9628 | 56.7341 | 136.9682 | 482.6292
0.72041 | 1104405.4561 | 74.8381 | 180.6752 | 474.4101
0.6845 | 1019049.1595 | 61.4214 | 148.2844 | 496.2945
0.69439 | 1071193.0017 | 69.7203 | 168.3196 | 482.8145
0.65999 | 968170.8188 | 89.7745 | 216.7347 | 487.4579
0.68932 | 1047922.1038 | 77.6676 | 187.5062 | 478.6236
0.70399 | 1054395.2376 | 59.6974 | 144.1224 | 494.5252
0.72964 | 1141372.3241 | 74.5244 | 179.9177 | 464.286
0.67749 | 1009969.6559 | 65.4766 | 158.0745 | 494.8713
0.67602 | 1014632.5958 | 63.9538 | 154.398 | 492.2144
0.67015 | 993403.5394 | 74.0682 | 178.8166 | 459.5683
0.68601 | 1030631.1384 | 73.3414 | 177.0617 | 491.068
0.69781 | 1045381.6875 | 78.1829 | 188.7503 | 516.861
0.67462 | 999185.7649 | 81.9863 | 197.9325 | 494.047
0.69747 | 1054478.6511 | 68.3297 | 164.9625 | 479.1546
0.70897 | 1088072.8656 | 86.5131 | 208.8611 | 491.1036
0.68453 | 1020655.8463 | 72.7881 | 175.7261 | 475.5348
0.67642 | 1003749.4035 | 67.0668 | 161.9135 | 511.3142
0.70753 | 1099049.4091 | 84.3783 | 203.7073 | 489.8651
0.68457 | 1035596.6614 | 58.1025 | 140.2719 | 493.4514
0.68008 | 1007425.4591 | 74.0168 | 178.6923 | 487.8682
0.69415 | 1061436.3404 | 79.802 | 192.6592 | 469.6191
0.66441 | 978215.361 | 66.5416 | 160.6457 | 518.2558
0.6938 | 1066891.8697 | 83.53 | 201.6592 | 462.8907
0.69762 | 1063343.679 | 68.889 | 166.3128 | 483.4157
0.69499 | 1056101.7108 | 82.5442 | 199.2794 | 474.1212
0.70415 | 1072957.1705 | 62.8663 | 151.7726 | 523.6589
0.69669 | 1064235.3587 | 65.0905 | 157.1422 | 516.2929
0.68501 | 1036867.9395 | 70.2848 | 169.6825 | 490.5692
0.70197 | 1069081.654 | 82.7273 | 199.7214 | 473.3472
0.6787 | 1020834.798 | 64.0908 | 154.729 | 467.1407
0.70932 | 1093014.8921 | 83.7028 | 202.0764 | 489.9835
0.69861 | 1071559.4055 | 69.3227 | 167.3598 | 501.5889
0.67133 | 971005.2813 | 68.9068 | 166.3558 | 478.1503
0.69822 | 1064222.917 | 70.7747 | 170.8652 | 495.1139
0.68717 | 1042861.4981 | 64.6665 | 156.1187 | 505.1084
0.69013 | 1036224.4724 | 86.1388 | 207.9576 | 506.1299
0.70756 | 1073522.0153 | 72.2816 | 174.5031 | 496.3681
0.70181 | 1065171.0936 | 77.5409 | 187.2003 | 486.8226
0.68484 | 1026597.4323 | 74.9007 | 180.8263 | 500.9365
0.70177 | 1060731.801 | 73.7877 | 178.1392 | 479.1889
0.70124 | 1063737.5884 | 75.7961 | 182.9879 | 471.4809
0.69836 | 1068597.1659 | 84.7008 | 204.4858 | 490.0878
0.67411 | 990696.1968 | 75.0823 | 181.2648 | 484.5481
0.67239 | 1003601.1823 | 72.864 | 175.9092 | 492.9951
0.68299 | 1037928.2794 | 77.64 | 187.4395 | 489.9125
0.691 | 1047335.3609 | 75.733 | 182.8356 | 498.8475
0.68788 | 1036136.8025 | 71.7597 | 173.2432 | 496.2923
0.68324 | 1024512.089 | 74.9244 | 180.8835 | 500.055
0.68513 | 1022472.8826 | 71.098 | 171.6457 | 490.8237
0.70064 | 1061162.1933 | 79.0423 | 190.825 | 513.3499
0.692 | 1055788.31 | 89.0236 | 214.922 | 486.3262
0.69439 | 1058317.5881 | 87.4355 | 211.0879 | 512.5406
0.68728 | 1029972.1903 | 79.8236 | 192.7112 | 471.7778
0.66969 | 997315.826 | 74.4104 | 179.6426 | 510.5273
0.68995 | 1040062.488 | 78.6713 | 189.9294 | 483.8939
0.69273 | 1033182.3639 | 65.5081 | 158.1505 | 470.4357
0.70005 | 1070809.9515 | 74.672 | 180.2741 | 501.5253
0.71438 | 1095208.0998 | 69.2886 | 167.2775 | 528.2985
0.68709 | 1061674.614 | 52.0504 | 125.6607 | 494.7212
0.67737 | 1019759.4704 | 87.8432 | 212.0723 | 471.3928
0.70005 | 1069833.572 | 83.0123 | 200.4094 | 483.382
0.71574 | 1102881.1987 | 75.2227 | 181.6036 | 487.9933
0.68667 | 1019257.0865 | 67.0861 | 161.9603 | 483.6607
0.68295 | 1027864.1638 | 76.8909 | 185.6311 | 467.2306
0.70059 | 1059459.3667 | 72.0444 | 173.9305 | 504.9609
0.7008 | 1064754.4821 | 80.9276 | 195.3766 | 482.0198
0.69275 | 1052785.9123 | 99.8031 | 240.946 | 459.2308
0.69273 | 1063359.0488 | 71.2782 | 172.0808 | 480.3751
0.66718 | 986571.6626 | 66.8312 | 161.3447 | 484.0172
0.69608 | 1051086.5116 | 68.2453 | 164.7588 | 481.6002
0.68397 | 1030547.2427 | 62.1632 | 150.0753 | 495.6362
0.69367 | 1046380.4865 | 71.8648 | 173.4969 | 461.6053
0.68464 | 1028236.5941 | 55.4204 | 133.7968 | 500.4719
0.69746 | 1060835.548 | 64.8102 | 156.4657 | 492.3898
0.70039 | 1059773.5658 | 68.3864 | 165.0994 | 496.1171
0.69143 | 1039043.1167 | 64.0329 | 154.589 | 469.6455
0.69709 | 1069028.111 | 80.0337 | 193.2184 | 489.4274
0.69006 | 1058911.018 | 63.3105 | 152.8451 | 490.4565
0.68332 | 1016592.5092 | 80.8124 | 195.0983 | 490.2904
0.6757 | 998782.6263 | 64.9231 | 156.7383 | 515.3718
0.66212 | 974069.9854 | 60.6955 | 146.532 | 509.5091
0.68212 | 1021598.0872 | 52.9131 | 127.7436 | 490.0787
0.68668 | 1030492.4116 | 76.6691 | 185.0955 | 497.643
0.70183 | 1074124.8501 | 69.1701 | 166.9915 | 510.0047
0.67886 | 1002162.8654 | 72.0666 | 173.9841 | 501.8108
0.66693 | 989261.4594 | 63.2045 | 152.5892 | 482.4429
0.69716 | 1054555.9514 | 91.1569 | 220.0722 | 484.0123
0.66166 | 979359.7392 | 63.0699 | 152.2642 | 477.1318
0.67469 | 1000584.3065 | 70.6522 | 170.5694 | 492.0506
0.68106 | 1021949.449 | 81.508 | 196.7777 | 498.7917
0.69188 | 1046739.7431 | 68.3758 | 165.0737 | 468.5483
0.67004 | 1006141.2739 | 67.866 | 163.843 | 478.5594
0.69698 | 1045801.5699 | 79.9227 | 192.9504 | 472.6836
0.66269 | 988385.5347 | 67.4081 | 162.7374 | 527.1302
0.67921 | 1030930.1303 | 51.1298 | 123.4383 | 488.0622
0.68772 | 1026645.2841 | 71.119 | 171.6964 | 522.2356
0.71213 | 1083755.5237 | 53.7669 | 129.8048 | 509.1029
0.69099 | 1038985.6888 | 83.2323 | 200.9404 | 492.5932
0.69818 | 1060504.753 | 77.6023 | 187.3484 | 485.8519
0.67937 | 1009390.1662 | 74.2381 | 179.2266 | 501.9471
0.70118 | 1064410.6085 | 66.6748 | 160.9672 | 483.2075
0.67719 | 1008431.9361 | 76.0788 | 183.6706 | 479.7091
0.68153 | 1034508.8956 | 77.5495 | 187.221 | 483.2642
0.64946 | 946944.687 | 58.1153 | 140.3027 | 476.688
0.68552 | 1018604.9645 | 80.7132 | 194.8589 | 476.8124
0.68445 | 1027673.6069 | 75.0228 | 181.121 | 495.978
0.67847 | 1008610.2075 | 79.1825 | 191.1635 | 488.0569
0.68526 | 1034864.3487 | 68.2533 | 164.7781 | 475.8242
0.65952 | 979725.6593 | 63.298 | 152.8149 | 515.4271
0.6703 | 998740.8282 | 76.1224 | 183.7758 | 491.5037
0.69375 | 1057410.8356 | 81.1864 | 196.0012 | 496.6299
0.70121 | 1067165.963 | 72.5044 | 175.041 | 471.5015
0.66389 | 981226.6953 | 54.5396 | 131.6703 | 481.5501
0.68559 | 1033318.3291 | 80.2803 | 193.8138 | 536.4335
0.69714 | 1038127.3743 | 65.4054 | 157.9026 | 525.0769
0.70586 | 1087163.7125 | 71.5176 | 172.6588 | 516.4539
0.70303 | 1072759.789 | 76.2437 | 184.0685 | 490.8654
0.70917 | 1082269.4722 | 62.9627 | 152.0054 | 499.9923
0.68722 | 1042835.8304 | 62.6067 | 151.1461 | 492.7306
0.67623 | 999929.3999 | 56.7932 | 137.1108 | 464.3192
0.68946 | 1026069.9935 | 82.3803 | 198.8836 | 497.0352
0.68629 | 1040785.1838 | 70.4188 | 170.006 | 508.2857
0.687 | 1046042.7992 | 77.79 | 187.8016 | 495.3146
0.67749 | 1012055.8692 | 90.0236 | 217.3361 | 515.5381
0.69475 | 1047733.5392 | 77.0465 | 186.0067 | 490.1204
0.71416 | 1089154.6846 | 74.4501 | 179.7383 | 505.8228
0.67184 | 1009831.7178 | 73.3763 | 177.146 | 467.9609
0.70482 | 1079029.6988 | 71.0878 | 171.621 | 497.7359
0.71493 | 1094282.925 | 76.4796 | 184.638 | 461.3087
0.68757 | 1029971.9546 | 85.9369 | 207.4701 | 471.4099
0.71205 | 1114258.447 | 83.3523 | 201.2303 | 471.2542
0.69354 | 1039239.4414 | 71.3875 | 172.3447 | 504.8768
0.68819 | 1035656.4887 | 69.8792 | 168.7032 | 486.1511
0.68482 | 1024824.3584 | 82.837 | 199.9863 | 507.7902
0.68275 | 1031774.4775 | 64.3935 | 155.4597 | 498.1017
0.71041 | 1085644.7312 | 83.0117 | 200.4079 | 520.7061
0.67964 | 1010014.6644 | 67.1649 | 162.1504 | 471.1345
0.69221 | 1041992.4945 | 76.9055 | 185.6663 | 494.3418
0.69686 | 1053433.7106 | 80.5191 | 194.3903 | 484.9536
0.70623 | 1086535.0048 | 75.105 | 181.3194 | 500.7798
0.67555 | 1003387.0891 | 65.2369 | 157.4958 | 480.6081
0.67796 | 1006426.8683 | 80.637 | 194.6749 | 490.4784
0.69631 | 1060823.2548 | 70.7422 | 170.7867 | 474.122
0.68412 | 1025202.7873 | 69.5854 | 167.994 | 464.5664
0.68393 | 1013728.7235 | 72.3577 | 174.687 | 470.5334
0.69129 | 1053505.8177 | 72.6031 | 175.2793 | 495.8861
0.67414 | 1009271.2126 | 76.4869 | 184.6556 | 500.8119
0.67666 | 1004308.6204 | 63.2433 | 152.6828 | 501.6408
0.71473 | 1101037.5143 | 73.3899 | 177.179 | 508.9574
0.67871 | 1020310.4491 | 83.2151 | 200.899 | 499.3547
0.69988 | 1056402.2055 | 57.9732 | 139.9597 | 468.9492
0.70354 | 1069741.7521 | 72.4384 | 174.8819 | 492.1332
0.7183 | 1107440.2257 | 78.7673 | 190.1611 | 475.4616
0.68935 | 1038408.8782 | 83.3838 | 201.3063 | 483.0908
0.69632 | 1050177.9149 | 75.3579 | 181.9301 | 487.8028
0.69937 | 1059126.5497 | 80.6641 | 194.7404 | 485.3817
0.67996 | 1015936.9139 | 65.6241 | 158.4306 | 512.2802
0.64777 | 934468.7291 | 66.9852 | 161.7166 | 469.733
0.68452 | 1024346.1799 | 84.3556 | 203.6525 | 460.7987
0.69004 | 1056125.5879 | 84.7224 | 204.5379 | 508.8317
0.68079 | 1001697.9446 | 73.5191 | 177.4908 | 475.4153
0.70457 | 1076369.8414 | 74.5783 | 180.0479 | 503.0358
0.70129 | 1054681.2342 | 71.6472 | 172.9716 | 488.9292
0.6934 | 1031535.0568 | 59.2947 | 143.1501 | 470.5688
0.68594 | 1037758.5366 | 83.7461 | 202.181 | 493.3085
0.67806 | 1009883.9234 | 79.331 | 191.5219 | 485.8583
0.66806 | 989595.6868 | 70.5698 | 170.3706 | 484.3942
0.68006 | 1009334.9352 | 67.0782 | 161.9411 | 510.2509
0.66882 | 988425.9454 | 68.6231 | 165.6708 | 503.6663
0.66439 | 989886.9492 | 63.486 | 153.2689 | 465.6278
0.68503 | 1032380.8917 | 78.6514 | 189.8813 | 471.3336
0.69457 | 1044549.5726 | 73.415 | 177.2395 | 481.2698
0.67998 | 1021072.8817 | 67.867 | 163.8454 | 479.7675
0.71297 | 1098311.27 | 76.9716 | 185.8259 | 492.7262
0.6737 | 1014032.2937 | 89.2777 | 215.5353 | 466.5991
0.68389 | 1011498.5447 | 63.0222 | 152.1491 | 480.8027
0.70562 | 1076972.8511 | 72.7124 | 175.5433 | 467.0478
0.67214 | 991050.7793 | 71.9799 | 173.7748 | 507.6568
0.67683 | 1002415.6754 | 65.0507 | 157.0462 | 510.8683
0.68496 | 1019514.2697 | 72.6612 | 175.4197 | 479.2414
0.69565 | 1044238.3794 | 75.0041 | 181.0759 | 509.5741
0.67177 | 989363.2171 | 62.7551 | 151.5041 | 485.4078
0.71642 | 1091541.09 | 64.9217 | 156.7349 | 483.4222
0.69223 | 1042155.71 | 81.7789 | 197.4316 | 523.6176
0.68139 | 1017294.6954 | 69.3244 | 167.3638 | 495.6969
0.67471 | 999383.9556 | 74.0106 | 178.6774 | 518.4763
0.67818 | 1002690.2644 | 73.8755 | 178.3513 | 485.6567
0.71345 | 1092078.0472 | 83.1471 | 200.7348 | 520.4788
0.70926 | 1077305.7995 | 79.2695 | 191.3736 | 478.3445
0.68473 | 1025295.6353 | 81.3372 | 196.3653 | 494.3387
0.6863 | 1036720.6133 | 75.4792 | 182.2229 | 485.7861
0.71205 | 1102549.6411 | 73.7678 | 178.0912 | 504.0181
0.70845 | 1081835.8193 | 72.7498 | 175.6335 | 505.3815
0.71478 | 1110460.5923 | 72.1383 | 174.1574 | 482.7818
0.71724 | 1097907.8646 | 72.9257 | 176.0582 | 484.0684
0.70605 | 1080199.2469 | 65.1909 | 157.3849 | 536.6846
0.68969 | 1047999.1265 | 72.3905 | 174.7661 | 517.8051
0.6885 | 1046467.0344 | 72.4101 | 174.8134 | 494.1433
0.6871 | 1024770.1369 | 57.2834 | 138.2944 | 481.6399
0.69485 | 1067879.2234 | 69.4588 | 167.6884 | 471.5277
0.67056 | 1000276.9253 | 77.8493 | 187.9448 | 470.6888
0.67525 | 1008580.54 | 74.6316 | 180.1765 | 500.7076
0.68494 | 1028797.9901 | 80.0047 | 193.1484 | 476.0268
0.66768 | 981442.9399 | 71.758 | 173.2392 | 494.7096
0.69118 | 1027652.3212 | 80.763 | 194.9792 | 496.4611
0.7048 | 1096472.154 | 82.8646 | 200.0528 | 483.0697
0.68177 | 1018444.0199 | 61.6308 | 148.79 | 487.0869
0.69246 | 1034151.442 | 88.582 | 213.8558 | 494.8696
0.69989 | 1051658.5879 | 55.4483 | 133.864 | 462.8004
0.69197 | 1052465.7147 | 75.9975 | 183.4741 | 490.1824
0.67904 | 1028584.7458 | 80.683 | 194.7861 | 499.3162
0.69258 | 1023840.9488 | 68.6433 | 165.7197 | 500.6122
0.6653 | 972271.1558 | 73.0208 | 176.2879 | 498.3333
0.6872 | 1034933.2081 | 69.9305 | 168.8271 | 472.0626
0.68626 | 1032058.4948 | 59.9428 | 144.7146 | 479.8433
0.69254 | 1034683.8218 | 52.7849 | 127.4339 | 490.6459
0.65281 | 955167.7231 | 60.3662 | 145.7369 | 497.9909
0.68388 | 1043209.9613 | 62.9348 | 151.9381 | 473.8606
0.7024 | 1080813.7671 | 76.8854 | 185.6177 | 488.7018
0.6964 | 1060981.7191 | 74.3738 | 179.5542 | 488.7434
0.69037 | 1019738.4654 | 70.3697 | 169.8875 | 476.3041
0.70547 | 1095893.4335 | 79.3631 | 191.5995 | 472.0764
0.68766 | 1054343.8542 | 56.5784 | 136.5922 | 491.4785
0.67229 | 994952.5527 | 66.8872 | 161.4799 | 510.0002
0.68943 | 1055431.122 | 79.8934 | 192.8797 | 461.4421
0.66176 | 968614.9136 | 81.0226 | 195.6058 | 483.614
0.6658 | 990377.3096 | 70.8656 | 171.0846 | 478.158
0.67287 | 999344.5843 | 61.2748 | 147.9305 | 483.776
0.70221 | 1061382.8502 | 80.5095 | 194.3672 | 506.247
0.70273 | 1092074.9998 | 86.024 | 207.6803 | 470.937
0.69281 | 1037038.8772 | 71.3114 | 172.1609 | 484.0723
0.6772 | 1004665.8077 | 67.0501 | 161.8732 | 489.042
0.69305 | 1068753.9987 | 61.4655 | 148.391 | 493.9832
0.70791 | 1088572.3263 | 76.8345 | 185.4949 | 490.3636
0.70018 | 1059265.2641 | 83.7429 | 202.1732 | 492.0497
0.71033 | 1092988.397 | 78.6118 | 189.7856 | 499.6414
0.70038 | 1072912.1576 | 71.133 | 171.7301 | 498.1825
0.69128 | 1041180.3739 | 76.9488 | 185.771 | 500.3484
0.69791 | 1074924.7826 | 82.038 | 198.0573 | 469.5528
0.68143 | 1013216.8822 | 72.5862 | 175.2387 | 490.5484
0.71446 | 1083969.2163 | 80.8129 | 195.0995 | 500.9357
0.67155 | 1003520.6983 | 71.1906 | 171.8692 | 488.7464
0.67561 | 1004056.2325 | 69.0079 | 166.5998 | 466.6491
0.67575 | 1015058.8679 | 75.0552 | 181.1992 | 489.2661
0.69173 | 1041862.5459 | 71.7551 | 173.2321 | 474.8992
0.7023 | 1076177.488 | 90.8907 | 219.4295 | 516.7059
0.68311 | 1019933.6545 | 72.8881 | 175.9675 | 493.1585
0.6944 | 1049924.3035 | 72.0316 | 173.8997 | 476.0968
0.68258 | 1009611.5932 | 70.4653 | 170.1184 | 465.6319
0.70848 | 1076863.246 | 74.2007 | 179.1362 | 511.1134
0.69094 | 1046348.1812 | 70.3194 | 169.766 | 495.0758
0.70224 | 1080810.3865 | 80.3226 | 193.916 | 515.8586
0.66849 | 991399.3617 | 68.8639 | 166.2522 | 466.2689
0.67675 | 999539.3475 | 68.5817 | 165.5708 | 502.0001
0.70084 | 1073340.6516 | 76.5234 | 184.7437 | 503.7508
0.69279 | 1039401.1039 | 59.7995 | 144.3688 | 488.2288
0.66931 | 992374.4838 | 75.7491 | 182.8746 | 484.6505
0.68187 | 1014376.8341 | 72.6221 | 175.3254 | 494.7149
0.6671 | 982810.6499 | 70.1871 | 169.4467 | 512.6706
0.67726 | 1003848.6582 | 77.3547 | 186.7507 | 492.9608
0.68584 | 1035860.8936 | 78.8232 | 190.296 | 487.9121
0.66361 | 971762.5154 | 79.543 | 192.0337 | 511.8423
0.68527 | 1034766.2397 | 70.116 | 169.275 | 489.9986
0.71609 | 1085238.4263 | 70.3708 | 169.8901 | 493.8541
0.70104 | 1078312.3336 | 78.5841 | 189.7189 | 496.9838
0.66869 | 979763.9288 | 73.2108 | 176.7466 | 486.2955
0.70256 | 1087862.6067 | 77.3695 | 186.7864 | 503.4021
0.69874 | 1059338.9221 | 74.0099 | 178.6758 | 470.0025
0.66615 | 993656.5876 | 77.8465 | 187.9381 | 502.4271
0.68155 | 1003945.0153 | 64.699 | 156.1971 | 509.319
0.70433 | 1088286.5147 | 75.4288 | 182.1013 | 483.1265
0.68749 | 1027311.2551 | 63.6726 | 153.7193 | 487.2509
0.69145 | 1055322.262 | 72.7117 | 175.5415 | 495.6971
0.66875 | 988220.0402 | 75.4953 | 182.2617 | 479.1963
0.70019 | 1063926.7854 | 74.1769 | 179.0788 | 481.8163
0.66764 | 997613.2639 | 69.384 | 167.5078 | 460.2733
0.68852 | 1032218.6122 | 82.9871 | 200.3485 | 479.7244
0.67035 | 1007122.5529 | 78.4069 | 189.2911 | 517.3836
0.71102 | 1091827.4883 | 74.9626 | 180.9757 | 491.6472
0.67359 | 1000844.281 | 85.2409 | 205.7898 | 502.7972
0.69608 | 1058046.2591 | 72.414 | 174.8229 | 478.4946
0.68773 | 1024612.3649 | 79.3563 | 191.5829 | 464.464
0.68745 | 1038092.4572 | 66.3249 | 160.1224 | 507.1383
0.69885 | 1055631.3225 | 68.2694 | 164.8169 | 513.0083
0.70006 | 1056807.4848 | 70.0971 | 169.2294 | 466.0531
0.68068 | 1008561.3581 | 61.1307 | 147.5825 | 498.8589
0.68816 | 1032791.9981 | 73.2655 | 176.8785 | 517.5454
0.69579 | 1068916.0296 | 67.4756 | 162.9006 | 482.4721
0.69346 | 1043206.6756 | 75.1668 | 181.4687 | 504.2439
0.69955 | 1068334.2277 | 67.0764 | 161.9369 | 498.9632
0.6748 | 1014961.5594 | 64.09 | 154.7269 | 454.79
0.69273 | 1045630.8032 | 88.0744 | 212.6305 | 513.9329
0.68745 | 1036183.5114 | 77.006 | 185.909 | 478.3177
0.69646 | 1071525.5037 | 86.7276 | 209.3789 | 492.8389
0.68794 | 1040915.6372 | 73.8651 | 178.3261 | 479.2536
0.68272 | 1026600.8829 | 57.7541 | 139.4307 | 492.1815
0.70393 | 1076044.8224 | 76.3793 | 184.3959 | 467.8561
0.65438 | 967268.5278 | 73.8576 | 178.308 | 511.0388
0.68858 | 1033923.8677 | 55.4083 | 133.7674 | 481.9425
0.72289 | 1119528.4332 | 84.1241 | 203.0935 | 491.9037
0.69173 | 1042268.906 | 75.2866 | 181.7578 | 523.1028
0.69146 | 1050632.5728 | 79.2968 | 191.4394 | 490.7745
0.70308 | 1076644.5194 | 74.7549 | 180.4744 | 497.5792
0.69377 | 1041943.7675 | 75.042 | 181.1675 | 503.671
0.66473 | 985179.1879 | 66.9832 | 161.7116 | 498.6222
0.69935 | 1069565.4496 | 89.2385 | 215.4407 | 495.4513
0.71135 | 1088329.5935 | 72.8882 | 175.9677 | 507.9372
0.67049 | 995761.557 | 63.2819 | 152.776 | 509.4734
0.68322 | 1033379.3406 | 60.2859 | 145.5431 | 478.562
0.682 | 1018499.4867 | 74.6199 | 180.1484 | 484.801
0.68145 | 1003921.3624 | 87.3823 | 210.9596 | 493.7255
0.69941 | 1063696.0291 | 84.1862 | 203.2435 | 494.1655
0.68304 | 1020531.3057 | 76.3328 | 184.2837 | 474.0739
0.72074 | 1099470.1542 | 91.9308 | 221.9405 | 476.4386
0.69204 | 1041497.9081 | 78.5418 | 189.6166 | 471.9069
0.70566 | 1069836.5468 | 70.3982 | 169.9562 | 488.4721
0.68425 | 1038967.6656 | 75.8691 | 183.1641 | 487.5602
0.68421 | 1038998.6403 | 62.6508 | 151.2523 | 453.9774
0.69374 | 1054642.9024 | 66.7717 | 161.2012 | 471.9464
0.67155 | 993702.8153 | 70.4948 | 170.1895 | 473.1407
0.69043 | 1038470.8619 | 66.0625 | 159.4891 | 476.3667
0.67902 | 1022061.0359 | 86.3506 | 208.4689 | 479.8198
0.6855 | 1027536.2697 | 93.1138 | 224.7966 | 513.6381
0.71116 | 1094018.4006 | 71.4328 | 172.4541 | 482.4723
0.69177 | 1047673.9781 | 77.3742 | 186.7977 | 477.5062
0.68373 | 1040248.8333 | 72.4817 | 174.9863 | 469.2093
0.70404 | 1060886.2416 | 64.2083 | 155.0125 | 517.2846
0.70771 | 1079408.0512 | 70.8219 | 170.9791 | 486.7538
0.67836 | 999975.6901 | 66.6227 | 160.8415 | 525.0397
0.70501 | 1082953.2647 | 71.3269 | 172.1984 | 499.7259
0.68227 | 1007210.7054 | 64.8409 | 156.5397 | 509.0704
0.71724 | 1102091.4662 | 65.4744 | 158.0693 | 505.5997
0.68023 | 1017941.3621 | 65.728 | 158.6815 | 497.2247
0.6982 | 1068837.0566 | 84.7395 | 204.5792 | 495.3442
0.68831 | 1035718.6174 | 62.0808 | 149.8762 | 505.0554
0.7149 | 1096304.6949 | 71.2068 | 171.9085 | 473.0235
0.66385 | 973267.2353 | 71.5705 | 172.7864 | 470.7368
0.70433 | 1073346.5585 | 70.7618 | 170.8341 | 496.1104
0.68562 | 1038415.9047 | 73.9371 | 178.4999 | 487.2144
0.683 | 1017141.0907 | 59.3578 | 143.3025 | 513.4097
0.66999 | 998769.996 | 76.1603 | 183.8671 | 502.8513
0.69635 | 1055691.1821 | 75.7501 | 182.8769 | 510.8945
0.67844 | 1022211.8013 | 76.9772 | 185.8394 | 487.7725
0.71108 | 1095034.1917 | 76.9454 | 185.7627 | 500.2467
0.69178 | 1039967.4786 | 71.6188 | 172.9031 | 492.8209
0.69485 | 1063413.9824 | 71.0637 | 171.5631 | 507.819
0.6765 | 1027073.673 | 66.6832 | 160.9874 | 483.4715
0.69736 | 1060814.9929 | 74.26 | 179.2794 | 481.4976
0.68294 | 1031812.5654 | 58.0285 | 140.0932 | 514.9111
0.67077 | 994101.5312 | 68.8164 | 166.1374 | 489.6965
0.70725 | 1077471.9483 | 77.8331 | 187.9056 | 507.2959
0.68462 | 1038198.7311 | 62.5647 | 151.0446 | 487.4757
0.71363 | 1089064.8462 | 60.827 | 146.8493 | 491.2646
0.68283 | 1010969.319 | 75.6406 | 182.6126 | 483.7176
0.70239 | 1083456.3578 | 73.4872 | 177.4137 | 495.4853
0.69368 | 1047340.015 | 80.4203 | 194.1518 | 505.8627
0.6738 | 990962.1223 | 82.9045 | 200.1493 | 487.5566
0.70296 | 1089844.9652 | 88.9424 | 214.726 | 508.244
0.6846 | 1033918.6387 | 71.9493 | 173.7009 | 487.0903
0.69441 | 1050998.2103 | 68.4045 | 165.143 | 520.7401
0.69874 | 1066077.5103 | 76.2339 | 184.0449 | 497.0302
0.67769 | 1012132.336 | 74.6247 | 180.1599 | 473.8337
0.67327 | 1005496.5368 | 74.6222 | 180.1538 | 497.9362
0.70496 | 1077285.9777 | 70.9583 | 171.3085 | 509.6048
0.69391 | 1058087.6246 | 60.7767 | 146.7278 | 507.8662
0.68297 | 1024625.9046 | 65.8936 | 159.0813 | 493.131
0.68088 | 1000885.8005 | 62.6373 | 151.2199 | 482.9989
0.69969 | 1060302.7744 | 64.5711 | 155.8884 | 487.1226
0.6847 | 1032957.7656 | 85.0425 | 205.3107 | 487.8817
0.69346 | 1037034.2513 | 55.184 | 133.2259 | 480.6061
0.68373 | 1026518.2398 | 75.0167 | 181.1062 | 466.0326
0.71142 | 1092217.6618 | 88.0756 | 212.6332 | 487.9187
0.67996 | 1031492.5228 | 70.4405 | 170.0584 | 502.6681
0.70287 | 1081230.0456 | 75.0964 | 181.2986 | 477.9523
0.7086 | 1075334.8293 | 77.5309 | 187.1761 | 500.0757
0.70514 | 1085202.6293 | 71.4614 | 172.5231 | 459.9628
0.68881 | 1034999.2714 | 73.3519 | 177.0871 | 489.0784
0.68327 | 1024538.2439 | 79.8688 | 192.8203 | 471.2843
0.68905 | 1035702.1596 | 78.4433 | 189.3788 | 500.8336
0.69942 | 1081377.6951 | 82.3524 | 198.8162 | 492.9544
0.66614 | 996340.424 | 90.1356 | 217.6067 | 494.5367
0.69926 | 1057605.9778 | 70.2305 | 169.5514 | 492.3528
0.69253 | 1057710.0483 | 79.708 | 192.4322 | 531.1513
0.69127 | 1038955.9585 | 75.0973 | 181.3008 | 490.0513
0.68673 | 1025827.8553 | 76.3161 | 184.2433 | 488.3716
0.70857 | 1089194.3943 | 80.2227 | 193.6748 | 484.5542
0.69703 | 1058509.3896 | 79.2531 | 191.3339 | 485.9131
0.73093 | 1136154.8884 | 86.0428 | 207.7258 | 471.3854
0.69464 | 1038049.3069 | 69.3895 | 167.521 | 489.497
0.71065 | 1093614.1585 | 78.6206 | 189.807 | 503.1273
0.68248 | 1013857.3067 | 67.1398 | 162.0898 | 510.0353
0.69409 | 1048929.1762 | 75.5103 | 182.298 | 498.3865
0.7085 | 1099635.2768 | 80.9995 | 195.5502 | 470.7045
0.69279 | 1046311.1067 | 80.4448 | 194.2109 | 507.2662
0.68891 | 1023448.2553 | 70.7752 | 170.8663 | 502.5522
0.67413 | 1011235.283 | 73.2288 | 176.7899 | 486.1954
0.68649 | 1036598.1909 | 70.115 | 169.2726 | 514.0447
0.69188 | 1054027.4438 | 66.9116 | 161.5389 | 492.6741
0.70143 | 1069440.9669 | 68.9078 | 166.3582 | 480.2797
0.68947 | 1039279.5624 | 83.439 | 201.4395 | 494.0484
0.669 | 1008979.5284 | 70.032 | 169.0722 | 520.6617
0.68996 | 1045384.1239 | 88.4073 | 213.4341 | 491.5816
0.69811 | 1071490.972 | 69.3681 | 167.4693 | 475.5524
0.68816 | 1066775.2553 | 80.9018 | 195.3141 | 485.9028
0.70728 | 1106258.1416 | 78.1521 | 188.6758 | 465.1267
0.67913 | 1003463.1452 | 75.4883 | 182.2449 | 502.229
0.66058 | 960081.0177 | 74.2523 | 179.2609 | 471.648
0.69782 | 1048056.0779 | 76.7313 | 185.2458 | 491.3596
0.69194 | 1045333.1279 | 83.6255 | 201.8898 | 505.062
0.67156 | 1017748.1911 | 75.7237 | 182.8132 | 498.0871
0.70479 | 1051327.3703 | 59.4561 | 143.5397 | 478.0196
0.68227 | 1016517.9359 | 67.337 | 162.5659 | 503.4724
0.6649 | 978532.5999 | 75.1881 | 181.52 | 487.6916
0.70332 | 1064589.0345 | 73.8928 | 178.393 | 500.5626
0.68186 | 1018079.0494 | 72.8884 | 175.9681 | 491.5541
0.67012 | 989717.1247 | 73.7964 | 178.1603 | 529.9357
0.67404 | 988651.6928 | 68.1079 | 164.4269 | 502.803
0.71206 | 1105214.8114 | 81.0035 | 195.5598 | 505.6348
0.69701 | 1056026.4502 | 79.0696 | 190.8909 | 511.6096
0.70714 | 1073498.9438 | 77.5577 | 187.2408 | 475.5726
0.68098 | 1018280.3379 | 76.3952 | 184.4344 | 488.1582
0.70245 | 1080181.8988 | 85.1203 | 205.4985 | 489.1553
0.67383 | 999759.8259 | 74.1352 | 178.9781 | 473.1612
0.67863 | 1003354.7245 | 86.2695 | 208.2729 | 493.254
0.68128 | 1020735.3102 | 63.8415 | 154.1269 | 462.4521
0.69507 | 1073779.7278 | 77.1907 | 186.3548 | 486.6708
0.68815 | 1030727.5474 | 78.1687 | 188.7159 | 489.364
0.69191 | 1037434.4437 | 73.1952 | 176.7088 | 508.0618
0.70739 | 1080415.8237 | 69.0287 | 166.65 | 522.4217
0.69737 | 1051700.4894 | 73.6207 | 177.736 | 540.297
0.67459 | 995598.9989 | 68.5306 | 165.4476 | 471.6602
0.70629 | 1070353.3152 | 83.8881 | 202.5237 | 481.6007
0.70494 | 1088899.9404 | 71.7258 | 173.1614 | 504.9671
0.68423 | 1033081.9677 | 66.3159 | 160.1007 | 486.38
0.67658 | 1004483.9254 | 74.1821 | 179.0913 | 492.3229
0.67119 | 988558.529 | 70.1669 | 169.3979 | 485.1562
0.70415 | 1080370.1966 | 63.3005 | 152.8209 | 503.1455
0.68733 | 1035648.6651 | 75.2876 | 181.7604 | 518.3814
0.67542 | 1005669.3144 | 72.687 | 175.4819 | 489.0499
0.68161 | 1032867.3451 | 66.5566 | 160.6819 | 500.7209
0.68465 | 1032290.7246 | 75.3368 | 181.8792 | 529.6532
0.69919 | 1101111.9968 | 78.6892 | 189.9726 | 494.2524
0.71389 | 1096058.3684 | 80.1263 | 193.4419 | 509.7224
0.69179 | 1065936.3135 | 71.547 | 172.7298 | 503.7059
0.69693 | 1062706.9476 | 68.9157 | 166.3772 | 451.3802
0.70061 | 1057512.9738 | 86.5935 | 209.0552 | 498.9493
0.68382 | 1018772.2375 | 74.4409 | 179.7162 | 514.972
0.6851 | 1034679.523 | 58.006 | 140.0389 | 510.0937
0.66703 | 972704.1339 | 77.6735 | 187.5204 | 501.1643
0.6922 | 1045406.4693 | 66.3343 | 160.1451 | 512.3312
0.68257 | 1046859.982 | 81.6924 | 197.2228 | 488.3958
0.6993 | 1039977.9554 | 75.9787 | 183.4288 | 498.3798
0.69286 | 1046993.7424 | 88.1735 | 212.8696 | 521.3532
0.69515 | 1046955.9345 | 87.8796 | 212.16 | 507.797
0.70226 | 1093365.8913 | 70.3661 | 169.8788 | 509.8068
0.68585 | 1028770.4798 | 78.2822 | 188.99 | 490.4584
0.66299 | 966435.9617 | 67.6863 | 163.4092 | 475.0958
0.68485 | 1016355.1552 | 74.1526 | 179.0202 | 498.2594
0.68248 | 1011580.6764 | 74.0608 | 178.7985 | 466.9613
0.6778 | 1028907.7687 | 60.5049 | 146.0718 | 484.4303
0.6946 | 1031516.6975 | 50.025 | 120.771 | 487.4325
0.6879 | 1040303.181 | 60.277 | 145.5215 | 516.3609
0.67323 | 1012928.0386 | 64.1108 | 154.7773 | 501.9529
0.66913 | 988018.8292 | 66.7343 | 161.1109 | 510.3632
0.7022 | 1060753.1307 | 83.6191 | 201.8742 | 510.6663
0.69598 | 1060631.5112 | 74.3946 | 179.6044 | 465.0735
0.71152 | 1086136.3877 | 81.3088 | 196.2967 | 522.2873
0.68373 | 1019229.0852 | 80.9738 | 195.488 | 510.6982
0.68687 | 1025421.1574 | 77.9118 | 188.0958 | 502.4205
0.68789 | 1028366.1748 | 71.2205 | 171.9414 | 488.1704
0.71405 | 1111648.2014 | 84.5659 | 204.1602 | 490.9844
0.70024 | 1064317.9905 | 79.3226 | 191.5016 | 486.6814
0.69792 | 1066215.5383 | 79.7282 | 192.481 | 499.3488
0.6836 | 1030105.6562 | 83.413 | 201.3769 | 491.3737
0.69761 | 1059608.094 | 75.8947 | 183.2261 | 453.1779
0.67966 | 1029729.4365 | 73.961 | 178.5577 | 500.0913
0.68828 | 1027136.4425 | 82.6072 | 199.4315 | 508.5501
0.68647 | 1030584.4084 | 71.5215 | 172.6682 | 513.7887
0.72364 | 1107651.6409 | 81.5301 | 196.831 | 500.0472
0.69063 | 1042705.6989 | 60.1779 | 145.2824 | 498.462
0.69139 | 1052971.2906 | 48.6703 | 117.5005 | 475.8462
0.7069 | 1069933.2126 | 73.8873 | 178.3796 | 492.2164
0.70312 | 1068029.65 | 62.6696 | 151.2978 | 491.3032
0.6928 | 1040335.0657 | 89.0683 | 215.03 | 501.9633
0.72788 | 1111020.2402 | 51.4173 | 124.1324 | 501.724
0.69119 | 1033172.4623 | 78.0095 | 188.3316 | 490.2885
0.71651 | 1114616.399 | 70.307 | 169.736 | 506.0951
0.66875 | 987422.3897 | 86.1537 | 207.9934 | 490.0211
0.68523 | 1031680.7132 | 69.9107 | 168.7794 | 474.7351
0.72022 | 1126941.2587 | 60.0981 | 145.0898 | 493.1277
0.66764 | 1002339.4937 | 57.663 | 139.2107 | 516.5206
0.6991 | 1052826.6413 | 73.5172 | 177.4861 | 510.5389
0.66967 | 995083.285 | 89.5475 | 216.1869 | 506.0404
0.67942 | 1008576.6328 | 88.0499 | 212.5713 | 498.6529
0.68891 | 1028219.1315 | 69.3048 | 167.3166 | 508.6219
0.6908 | 1039629.1726 | 73.4687 | 177.3692 | 507.5917
0.65631 | 962777.1891 | 60.5226 | 146.1145 | 489.788
0.72047 | 1118411.0443 | 75.7803 | 182.9498 | 480.9178
0.68594 | 1026338.3425 | 74.7617 | 180.4908 | 517.9486
0.70343 | 1081384.8275 | 69.4791 | 167.7374 | 494.6943
0.70166 | 1076031.5944 | 76.6254 | 184.9901 | 486.1712
0.70604 | 1076167.5577 | 84.3731 | 203.6947 | 473.0907
0.67517 | 1008617.4879 | 45.0091 | 108.6617 | 478.3562
0.68939 | 1039156.5864 | 68.1257 | 164.4701 | 502.3296
0.6907 | 1035160.685 | 65.8265 | 158.9191 | 494.7481
0.70962 | 1096298.9664 | 65.8297 | 158.927 | 514.895
0.66706 | 989614.1486 | 65.1603 | 157.3109 | 513.0546
0.68958 | 1028430.1175 | 74.5185 | 179.9037 | 499.5216
0.67628 | 1011280.6922 | 61.6947 | 148.9442 | 486.0134
0.67636 | 1016642.4568 | 78.4446 | 189.3821 | 490.9353
0.67098 | 983608.373 | 81.1462 | 195.9043 | 489.7555
0.70756 | 1086814.6773 | 86.5853 | 209.0354 | 489.9136
0.70589 | 1070239.5738 | 77.4595 | 187.0038 | 500.3124
0.67568 | 996044.8324 | 64.7473 | 156.3139 | 477.6678
0.69449 | 1036571.4952 | 61.5792 | 148.6653 | 485.8361
0.68556 | 1041137.8364 | 68.6368 | 165.704 | 457.9151
0.67016 | 1001300.5908 | 85.7162 | 206.9373 | 503.2017
0.68773 | 1045918.4533 | 87.248 | 210.6353 | 479.7497
0.68499 | 1024291.416 | 69.4352 | 167.6315 | 526.8413
0.68198 | 1029290.8766 | 70.0821 | 169.1932 | 460.9844
0.7116 | 1085017.4124 | 74.4341 | 179.6998 | 503.1737
0.6952 | 1042246.9133 | 67.7575 | 163.5812 | 509.942
0.70844 | 1083372.0654 | 90.7056 | 218.9827 | 472.1378
0.67336 | 1008596.5441 | 65.789 | 158.8287 | 516.0079
0.68477 | 1029612.3784 | 69.8175 | 168.5544 | 505.8448
0.66265 | 968522.0013 | 69.6314 | 168.105 | 506.7907
0.68803 | 1030444.5263 | 63.8165 | 154.0666 | 476.512
0.70456 | 1089518.8273 | 71.8046 | 173.3516 | 535.6304
0.69206 | 1052853.5517 | 82.6378 | 199.5052 | 492.3149
0.68969 | 1052183.9128 | 65.0063 | 156.939 | 486.1434
0.7005 | 1078882.4375 | 76.9565 | 185.7895 | 504.8242
0.68233 | 1005087.126 | 70.176 | 169.4198 | 502.4424
0.69412 | 1061220.9243 | 67.1515 | 162.118 | 491.7116
0.66531 | 986320.4437 | 52.3767 | 126.4485 | 488.4211
0.69805 | 1051803.0117 | 68.0995 | 164.4068 | 493.8107
0.68306 | 1019629.8346 | 59.4981 | 143.6411 | 476.8915
0.69576 | 1053260.0053 | 70.2935 | 169.7035 | 515.5656
0.67258 | 1014409.2769 | 80.3355 | 193.9469 | 516.8386
0.67552 | 1005040.6906 | 77.6455 | 187.4528 | 517.0144
0.70052 | 1074025.5624 | 87.9569 | 212.3466 | 494.5034
0.66416 | 986083.2055 | 70.7328 | 170.764 | 489.9176
0.68822 | 1031853.1126 | 76.599 | 184.9263 | 486.1417
0.68173 | 1031234.6088 | 77.281 | 186.5728 | 480.8562
0.6868 | 1031973.9817 | 65.9968 | 159.3304 | 495.6153
0.69209 | 1071424.8469 | 71.8848 | 173.5452 | 501.7497
0.70512 | 1069397.2855 | 74.0375 | 178.7424 | 501.207
0.68802 | 1034094.7332 | 80.0567 | 193.2739 | 500.545
0.6727 | 1007742.241 | 71.368 | 172.2976 | 476.4458
0.70909 | 1081693.3279 | 68.9442 | 166.4459 | 515.1117
0.69985 | 1073474.0149 | 93.8698 | 226.6217 | 490.0351
0.68482 | 1024520.5374 | 63.747 | 153.8989 | 501.6195
0.68988 | 1029944.7395 | 85.6474 | 206.771 | 503.5076
0.70391 | 1063249.839 | 80.4563 | 194.2387 | 478.7234
0.68732 | 1035612.6568 | 81.8456 | 197.5928 | 511.6293
0.70243 | 1076457.0878 | 81.6269 | 197.0647 | 472.4941
0.67797 | 1001466.7829 | 64.8796 | 156.6333 | 492.1525
0.69287 | 1053134.3188 | 75.9465 | 183.3511 | 493.3003
0.69369 | 1037065.8798 | 78.0213 | 188.36 | 511.9533
0.69727 | 1054644.0193 | 78.9364 | 190.5692 | 489.2826
0.69249 | 1045983.6183 | 68.2664 | 164.8097 | 478.7373
0.66878 | 986455.6768 | 64.7852 | 156.4054 | 488.1779
0.7139 | 1102458.4634 | 71.707 | 173.1159 | 496.6228
0.66621 | 980543.9836 | 71.8502 | 173.4616 | 491.3335
0.69676 | 1054984.1016 | 75.6557 | 182.6491 | 491.565
0.68249 | 1019321.2119 | 79.5523 | 192.0563 | 512.8068
0.71319 | 1092934.3645 | 79.9766 | 193.0807 | 474.9746
0.68915 | 1043964.8808 | 77.3934 | 186.8442 | 503.9837
0.69877 | 1060150.3726 | 68.5688 | 165.5397 | 482.6436
0.6913 | 1036786.1682 | 73.2144 | 176.7551 | 511.55
0.68972 | 1044740.2728 | 70.3355 | 169.805 | 510.1601
0.70564 | 1083495.1403 | 69.3968 | 167.5387 | 475.3516
0.70822 | 1076362.8024 | 53.6954 | 129.6322 | 510.8343
0.6952 | 1067341.9882 | 71.2908 | 172.1111 | 487.5838
0.71547 | 1100854.2576 | 69.5397 | 167.8837 | 506.4845
0.68685 | 1044235.9917 | 75.965 | 183.3958 | 507.994
0.6642 | 973357.1319 | 58.4101 | 141.0145 | 490.385
0.68252 | 1021394.5111 | 67.6359 | 163.2874 | 452.6091
0.69039 | 1044568.4921 | 60.3059 | 145.5914 | 490.3229
0.71239 | 1097179.0457 | 78.6847 | 189.9617 | 465.394
0.68587 | 1022525.6958 | 74.4918 | 179.8391 | 475.8826
0.6882 | 1020125.8523 | 68.8294 | 166.169 | 502.4227
0.68469 | 1040475.2319 | 64.7517 | 156.3246 | 499.2302
0.69557 | 1045501.4034 | 86.2304 | 208.1786 | 490.8657
0.70882 | 1092450.6541 | 77.0096 | 185.9175 | 497.3539
0.67706 | 1008225.8383 | 76.0785 | 183.6698 | 488.3245
0.70174 | 1051699.7883 | 73.1627 | 176.6303 | 464.288
0.68995 | 1034254.7866 | 79.8458 | 192.7648 | 500.3587
0.69405 | 1066883.9645 | 85.0901 | 205.4257 | 513.5464
0.72024 | 1100810.6435 | 70.9696 | 171.3358 | 497.458
0.6876 | 1031296.3766 | 76.6265 | 184.9928 | 473.6008
0.69626 | 1059348.66 | 68.7645 | 166.0121 | 517.3799
0.6889 | 1028961.4146 | 76.3084 | 184.2249 | 496.5506
0.69792 | 1061369.2613 | 72.45 | 174.9097 | 469.8914
0.68607 | 1035726.5866 | 53.3365 | 128.7657 | 504.3145
0.69696 | 1061837.5187 | 87.4031 | 211.0097 | 489.4914
0.67647 | 1007722.8177 | 80.919 | 195.3557 | 507.9938
0.69185 | 1043333.649 | 68.1569 | 164.5453 | 467.0229
0.70109 | 1059722.1953 | 95.2653 | 229.9908 | 507.5857
0.69265 | 1051974.8772 | 70.5225 | 170.2564 | 502.9627
0.70249 | 1071712.3117 | 72.6797 | 175.4644 | 509.3916
0.68744 | 1051616.6041 | 83.096 | 200.6116 | 459.2475
0.67777 | 1017105.746 | 69.6817 | 168.2265 | 481.4928
0.67852 | 1007288.4268 | 70.7246 | 170.7443 | 502.986
0.67035 | 993647.7982 | 75.7292 | 182.8265 | 476.1467
0.69379 | 1038802.2972 | 63.8004 | 154.0278 | 496.3442
0.70039 | 1049930.2827 | 71.4166 | 172.415 | 511.1291
0.69485 | 1047572.3413 | 70.0056 | 169.0084 | 494.1914
0.69926 | 1050791.8754 | 76.7729 | 185.3463 | 513.1217
0.67835 | 1017477.8343 | 63.9126 | 154.2987 | 501.4116
0.66206 | 986864.3199 | 67.3687 | 162.6425 | 485.9739
0.68471 | 1037475.9227 | 77.4105 | 186.8854 | 481.7533
0.6873 | 1030151.0897 | 63.8664 | 154.1871 | 500.199
0.69184 | 1051751.0335 | 72.4241 | 174.8472 | 492.6464
0.63194 | 909296.127 | 73.5122 | 177.4742 | 489.2163
0.68496 | 1032968.3343 | 78.7032 | 190.0064 | 480.9452
0.6834 | 1023622.8817 | 61.7972 | 149.1917 | 498.2255
0.67969 | 1013812.7369 | 81.0223 | 195.6052 | 483.2193
0.70264 | 1077599.6368 | 80.038 | 193.2287 | 497.4561
0.68161 | 1018660.2108 | 69.3957 | 167.536 | 493.3748
0.68276 | 1031970.7591 | 78.8812 | 190.4359 | 491.2973
0.69173 | 1041006.1417 | 78.5404 | 189.6133 | 492.0218
0.70045 | 1064785.0739 | 75.9562 | 183.3744 | 481.821
0.66918 | 987512.9758 | 73.736 | 178.0144 | 520.5935
0.68565 | 1027015.07 | 78.1718 | 188.7233 | 485.0337
0.69799 | 1052516.1085 | 97.7162 | 235.9079 | 501.2601
0.68979 | 1035517.6785 | 75.0091 | 181.088 | 479.2265
0.67248 | 996448.3062 | 78.9255 | 190.5429 | 497.9532
0.69087 | 1041431.2008 | 78.6693 | 189.9245 | 460.3058
0.68045 | 1031185.5312 | 82.3191 | 198.7358 | 509.7673
0.70363 | 1078086.4194 | 87.899 | 212.2071 | 522.8662
0.70848 | 1099980.2864 | 68.1931 | 164.6327 | 487.9178
0.69547 | 1040308.4863 | 65.7001 | 158.614 | 490.5363
0.70546 | 1088103.667 | 77.068 | 186.0586 | 498.1811
0.71259 | 1100982.6984 | 73.6071 | 177.7033 | 510.2141
0.69631 | 1054930.5774 | 75.7005 | 182.7571 | 490.5376
0.70194 | 1061697.6043 | 80.5643 | 194.4995 | 486.5254
0.69497 | 1059007.3746 | 85.8984 | 207.3771 | 532.4943
0.69266 | 1051702.272 | 72.3355 | 174.6333 | 486.8877
0.69142 | 1050411.5251 | 79.2853 | 191.4116 | 466.8973
0.68548 | 1024557.8114 | 65.6485 | 158.4895 | 464.9705
0.70047 | 1076754.1435 | 75.0413 | 181.1657 | 493.589
0.69978 | 1070493.5395 | 56.4504 | 136.2833 | 490.6394
0.69061 | 1038066.9724 | 73.378 | 177.1501 | 503.1926
0.67998 | 1028852.2806 | 73.8634 | 178.3219 | 479.8218
0.68071 | 1018997.5502 | 78.1245 | 188.6092 | 499.2607
0.67202 | 981091.1224 | 74.5858 | 180.0661 | 462.0084
0.70679 | 1078754.1733 | 70.8712 | 171.0981 | 515.6438
0.66764 | 994131.352 | 50.1591 | 121.0949 | 459.0202
0.71541 | 1100969.5326 | 65.575 | 158.312 | 493.3796
0.7086 | 1087587.0381 | 89.5566 | 216.2088 | 513.8649
0.67458 | 1011665.4777 | 84.5384 | 204.0939 | 485.8324
0.6872 | 1029924.7716 | 69.5923 | 168.0106 | 482.1857
0.68916 | 1045421.3237 | 68.5591 | 165.5163 | 459.4826
0.6609 | 958579.6409 | 69.0985 | 166.8186 | 522.7814
0.66724 | 986790.5559 | 75.5564 | 182.4092 | 490.1413
0.67787 | 1012484.1037 | 69.355 | 167.4378 | 495.3227
0.66951 | 993145.7956 | 59.1253 | 142.7412 | 491.0547
0.69759 | 1064451.0979 | 72.6947 | 175.5005 | 491.6328
0.67418 | 1001646.9691 | 67.1478 | 162.1091 | 505.3914
0.71967 | 1115174.3781 | 71.9527 | 173.7092 | 465.3163
0.6885 | 1036792.5426 | 66.1406 | 159.6776 | 485.332
0.69297 | 1061242.7534 | 78.2409 | 188.8902 | 516.0077
0.72342 | 1127016.4926 | 76.052 | 183.6057 | 491.128
0.6824 | 1041888.4924 | 69.0042 | 166.5909 | 510.6708
0.6887 | 1041972.4013 | 73.8585 | 178.3101 | 474.7042
0.68774 | 1036206.2094 | 73.5524 | 177.5711 | 486.436
0.66857 | 977779.6271 | 77.8994 | 188.0658 | 482.1245
0.69699 | 1065904.4546 | 87.7337 | 211.8079 | 508.6307
0.68696 | 1030669.3013 | 63.9302 | 154.3411 | 481.8323
0.68817 | 1045138.1013 | 84.0681 | 202.9584 | 506.1422
0.68499 | 1021712.816 | 76.2041 | 183.9729 | 497.2836
0.68657 | 1031919.686 | 76.7875 | 185.3815 | 480.2985
0.6841 | 1023197.8552 | 69.4511 | 167.6697 | 477.3514
0.7049 | 1085261.7033 | 71.4896 | 172.5912 | 510.9126
0.67821 | 1024856.689 | 50.9033 | 122.8914 | 506.326
0.70858 | 1087229.5298 | 91.0103 | 219.7183 | 485.14
0.69409 | 1047407.5408 | 75.401 | 182.0342 | 478.5189
0.7147 | 1092813.0688 | 72.8448 | 175.863 | 483.0619
0.68737 | 1029649.4135 | 73.1258 | 176.5414 | 472.5071
0.69732 | 1073401.8548 | 59.1227 | 142.7348 | 472.1311
0.65752 | 960377.6569 | 64.1524 | 154.8775 | 512.482
0.70991 | 1095300.5389 | 77.0919 | 186.1164 | 490.4216
0.72346 | 1112230.2769 | 75.0508 | 181.1886 | 515.8116
0.68785 | 1022490.5361 | 79.1781 | 191.1528 | 491.8019
0.69809 | 1046932.007 | 76.6192 | 184.975 | 482.4658
0.68582 | 1041223.4082 | 75.8593 | 183.1406 | 467.213
0.69841 | 1061945.1969 | 80.9894 | 195.5257 | 485.5658
0.70337 | 1067320.6829 | 60.4828 | 146.0184 | 485.6118
0.69922 | 1071888.9123 | 78.3189 | 189.0786 | 502.1484
0.71073 | 1090094.2451 | 78.3588 | 189.175 | 481.9861
0.68152 | 1028378.821 | 68.1978 | 164.644 | 475.4364
0.68382 | 1024847.2882 | 65.8263 | 158.9188 | 488.4342
0.69512 | 1062831.0353 | 67.1565 | 162.13 | 485.5463
0.7044 | 1068888.0657 | 66.6491 | 160.9052 | 469.6916
0.67013 | 994948.3689 | 72.4796 | 174.9813 | 488.4196
0.68258 | 1028086.9366 | 64.1791 | 154.942 | 491.0165
0.69512 | 1039989.5108 | 76.4361 | 184.533 | 498.7761
0.6863 | 1031534.7894 | 62.0115 | 149.7091 | 510.1835
0.70496 | 1067975.8058 | 58.186 | 140.4735 | 479.6081
0.69422 | 1038503.9878 | 69.4626 | 167.6975 | 505.0194
0.69249 | 1044685.6137 | 66.2601 | 159.9659 | 482.3002
0.69028 | 1045538.3226 | 91.2784 | 220.3655 | 473.8194
0.68309 | 1017024.4704 | 69.1344 | 166.9052 | 500.7926
0.67831 | 1009798.4117 | 65.3869 | 157.8579 | 493.4908
0.66727 | 979357.7985 | 64.4647 | 155.6315 | 491.0526
0.71021 | 1100505.6757 | 71.9517 | 173.7067 | 487.2013
0.68049 | 1028566.8549 | 75.5239 | 182.3308 | 480.8515
0.67895 | 1035467.3832 | 79.593 | 192.1545 | 486.7866
0.68945 | 1049645.5658 | 70.6931 | 170.6682 | 472.0213
0.69573 | 1060154.2561 | 76.1383 | 183.8141 | 494.6226
0.67804 | 1018951.9825 | 73.3619 | 177.1114 | 502.409
0.68798 | 1005524.2742 | 79.3917 | 191.6686 | 483.406
0.69232 | 1041270.3033 | 75.1505 | 181.4293 | 501.9433
0.69585 | 1052906.2691 | 68.3869 | 165.1005 | 526.1481
0.67567 | 1007187.7044 | 76.1564 | 183.8577 | 486.6183
0.70104 | 1054444.538 | 80.1159 | 193.4169 | 464.3061
0.67042 | 986489.3268 | 70.1005 | 169.2376 | 481.5431
0.68559 | 1050398.7985 | 92.3305 | 222.9056 | 501.8689
0.69925 | 1061538.4685 | 82.4639 | 199.0854 | 513.5829
0.69101 | 1028190.4811 | 65.3543 | 157.7793 | 474.2972
0.71562 | 1101051.754 | 81.953 | 197.852 | 537.8425
0.69564 | 1056435.8677 | 75.8815 | 183.1942 | 490.6096
0.7023 | 1076260.0319 | 74.8111 | 180.61 | 473.5244
0.72074 | 1118400.0736 | 75.7512 | 182.8795 | 494.8377
0.68382 | 1028632.4445 | 81.7102 | 197.2658 | 506.5681
0.69388 | 1044352.8743 | 75.2034 | 181.5571 | 505.0041
0.68897 | 1050233.0988 | 74.415 | 179.6538 | 492.3225
0.6954 | 1064062.0687 | 60.9888 | 147.24 | 530.1643
0.67881 | 997048.793 | 69.9749 | 168.9344 | 489.371
0.69429 | 1042327.4209 | 83.1029 | 200.6282 | 497.4018
0.70876 | 1096348.9296 | 79.0278 | 190.7899 | 497.4204
0.68243 | 1024562.7625 | 59.8405 | 144.4678 | 459.2785
0.70447 | 1079567.4299 | 84.828 | 204.793 | 472.0826
0.70807 | 1087411.6009 | 81.8004 | 197.4835 | 487.3834
0.69607 | 1066475.7441 | 74.7949 | 180.5708 | 518.4456
0.69418 | 1053710.8657 | 56.013 | 135.2272 | 521.1454
0.68671 | 1032007.0185 | 75.0911 | 181.2861 | 483.3893
0.68743 | 1037423.6903 | 60.1517 | 145.219 | 483.622
0.70123 | 1073228.9888 | 71.0671 | 171.5712 | 496.0739
0.67669 | 1011525.3773 | 74.6901 | 180.3179 | 502.4324
0.69422 | 1053246.454 | 83.9823 | 202.7513 | 483.0132
0.69518 | 1052830.488 | 70.3231 | 169.7751 | 519.5402
0.70258 | 1086389.4246 | 83.9932 | 202.7776 | 494.7769
0.66513 | 973093.2862 | 82.4704 | 199.1011 | 494.2138
0.67484 | 1016807.6295 | 55.6006 | 134.2316 | 468.2911
0.70945 | 1068720.6237 | 68.083 | 164.3669 | 485.6834
0.66891 | 988568.8717 | 75.5041 | 182.2829 | 474.2158
0.69691 | 1058789.2172 | 79.2882 | 191.4187 | 475.064
0.68661 | 1060982.0211 | 73.6164 | 177.7256 | 508.5908
0.69627 | 1059914.7682 | 77.7719 | 187.7579 | 504.9736
0.66842 | 999231.4387 | 74.695 | 180.3296 | 448.6894
0.67577 | 1001639.9953 | 77.3542 | 186.7496 | 482.4881
0.69802 | 1040116.3049 | 77.7661 | 187.744 | 495.0312
0.6907 | 1043024.7869 | 62.303 | 150.4128 | 532.478
0.67452 | 1019170.5075 | 64.3824 | 155.4329 | 500.8207
0.67693 | 1018405.1865 | 64.1053 | 154.7639 | 503.8693
0.6864 | 1044219.7972 | 85.7569 | 207.0354 | 487.097
0.6864 | 1042849.1519 | 84.1003 | 203.0362 | 510.1693
0.67752 | 1013535.4094 | 79.6816 | 192.3684 | 487.0913
0.70845 | 1098593.8466 | 66.0761 | 159.5217 | 486.5749
0.68253 | 1028953.112 | 75.3267 | 181.8547 | 502.2558
0.68403 | 1030644.1621 | 66.1786 | 159.7693 | 459.8868
0.67311 | 994651.8554 | 68.6155 | 165.6525 | 497.6992
0.68541 | 1022151.3259 | 67.4004 | 162.7189 | 473.3125
0.68104 | 1020977.174 | 79.8513 | 192.778 | 509.2223
0.69835 | 1062886.6584 | 64.8531 | 156.5692 | 500.2055
0.69881 | 1059439.7726 | 77.4549 | 186.9926 | 499.0022
0.68406 | 1028667.0236 | 89.1946 | 215.3349 | 522.9008
0.71602 | 1110747.3642 | 79.2538 | 191.3357 | 489.067
0.69517 | 1049933.265 | 62.3124 | 150.4354 | 484.3506
0.7058 | 1080404.1873 | 84.734 | 204.5661 | 479.0137
0.69349 | 1062097.4493 | 78.2445 | 188.899 | 480.5579
0.71743 | 1100775.258 | 73.097 | 176.4717 | 503.6858
0.70357 | 1077740.6617 | 78.2599 | 188.936 | 492.5523
0.67218 | 1001318.4581 | 80.9182 | 195.3539 | 504.9478
0.70644 | 1078210.2809 | 74.508 | 179.8782 | 479.3677
0.6997 | 1070360.1158 | 74.6782 | 180.2891 | 493.5054
0.70006 | 1052534.3026 | 71.627 | 172.9229 | 446.1286
0.67814 | 1023232.8019 | 62.2528 | 150.2915 | 485.484
0.69191 | 1046909.8964 | 80.5936 | 194.5702 | 509.3548
0.71709 | 1088888.3603 | 64.7036 | 156.2082 | 482.4316
0.71619 | 1106958.1194 | 73.6339 | 177.768 | 482.3067
0.69252 | 1052567.5967 | 72.4093 | 174.8114 | 499.7161
0.68974 | 1049804.6023 | 68.7892 | 166.0719 | 505.4991
0.67856 | 1009459.6332 | 70.8728 | 171.1021 | 523.8106
0.70941 | 1071003.3198 | 75.4346 | 182.1153 | 495.459
0.67258 | 989080.6361 | 68.2849 | 164.8544 | 504.4693
0.71059 | 1094374.8043 | 67.6438 | 163.3065 | 479.6305
0.71047 | 1094044.8708 | 83.6466 | 201.9407 | 517.7103
0.69186 | 1054546.0242 | 88.9378 | 214.7147 | 464.1996
0.67512 | 1014200.8854 | 71.5679 | 172.7803 | 500.8087
0.68927 | 1030971.7662 | 69.4535 | 167.6757 | 494.9352
0.7059 | 1075850.6302 | 60.747 | 146.6562 | 498.9312
0.66312 | 972252.8147 | 56.3137 | 135.9533 | 503.8725
0.68327 | 1038566.0707 | 72.4134 | 174.8215 | 475.9372
0.68829 | 1029573.0374 | 84.3215 | 203.5702 | 480.9729
0.68081 | 1014032.4642 | 77.7323 | 187.6625 | 520.5283
0.68429 | 1031557.0142 | 69.9985 | 168.9914 | 481.1378
0.69616 | 1058252.8359 | 88.6991 | 214.1385 | 484.206
0.66946 | 984319.4629 | 70.0472 | 169.1089 | 475.8469
0.67321 | 993825.2197 | 66.7201 | 161.0765 | 529.2456
0.66734 | 977106.8275 | 69.754 | 168.4011 | 486.1187
0.69102 | 1058101.6209 | 67.2542 | 162.366 | 495.7807
0.66618 | 989396.3551 | 82.2113 | 198.4757 | 499.955
0.70479 | 1079096.7576 | 71.0951 | 171.6388 | 450.3114
0.69213 | 1038111.0211 | 76.6277 | 184.9955 | 471.0538
0.68223 | 1014304.6032 | 70.8295 | 170.9976 | 485.0375
0.71269 | 1082178.0109 | 84.0528 | 202.9213 | 486.452
0.66762 | 980889.2136 | 75.0619 | 181.2156 | 487.8499
0.69954 | 1054935.017 | 73.6183 | 177.7302 | 504.949
0.67468 | 990205.2096 | 62.5385 | 150.9812 | 488.6461
0.7077 | 1082472.1986 | 84.6688 | 204.4086 | 476.5863
0.6882 | 1042765.8115 | 73.968 | 178.5746 | 526.7449
0.69487 | 1068269.205 | 69.4819 | 167.7441 | 516.1408
0.70202 | 1061372.229 | 82.1247 | 198.2666 | 486.2231
0.70756 | 1077631.9083 | 60.7052 | 146.5552 | 475.9089
0.70411 | 1084019.6484 | 71.4882 | 172.5877 | 508.5098
0.70425 | 1075594.1618 | 74.5583 | 179.9997 | 500.9978
0.70765 | 1073891.8585 | 70.4663 | 170.1206 | 505.9517
0.67008 | 1006268.049 | 73.3093 | 176.9843 | 495.6453
0.67628 | 993283.5184 | 66.381 | 160.2578 | 478.4327
0.70294 | 1093296.22 | 66.2732 | 159.9976 | 487.5612
0.69719 | 1045743.2909 | 57.6098 | 139.0824 | 520.7443
0.67461 | 1003313.3357 | 63.3726 | 152.9949 | 480.3839
0.68807 | 1015021.2448 | 84.6289 | 204.3121 | 467.122
0.69165 | 1054254.9833 | 76.162 | 183.8713 | 494.2047
0.69525 | 1066196.6382 | 74.8634 | 180.7363 | 496.6436
0.70871 | 1087136.4832 | 73.079 | 176.4284 | 482.1305
0.67028 | 981574.6247 | 74.817 | 180.6242 | 506.8395
0.68433 | 1032353.7288 | 64.2801 | 155.1859 | 491.0039
0.68917 | 1036503.8341 | 67.9736 | 164.1027 | 476.4028
0.70153 | 1071963.003 | 78.0658 | 188.4676 | 459.0381
0.69523 | 1040265.9811 | 73.045 | 176.3462 | 503.6864
0.67705 | 1000323.4055 | 77.1066 | 186.1518 | 502.466
0.71145 | 1089273.8516 | 84.7514 | 204.608 | 469.6161
0.70307 | 1071361.2634 | 81.1669 | 195.9543 | 492.7087
0.70603 | 1085912.0203 | 67.5541 | 163.09 | 516.577
0.67043 | 993374.768 | 61.1305 | 147.5821 | 498.3461
0.71286 | 1089671.8001 | 67.139 | 162.0879 | 480.5345
0.69217 | 1027659.0332 | 79.0812 | 190.919 | 495.6664
0.71193 | 1080005.7846 | 83.9732 | 202.7292 | 494.0279
0.68211 | 1041799.2203 | 86.0953 | 207.8524 | 493.7073
0.70524 | 1087569.2471 | 69.697 | 168.2635 | 495.927
0.70142 | 1056827.7897 | 81.8433 | 197.5872 | 497.7389
0.6624 | 992252.9193 | 73.1266 | 176.5433 | 505.0952
0.68301 | 1026931.6031 | 80.0855 | 193.3436 | 482.2455
0.68929 | 1037778.4064 | 73.4979 | 177.4396 | 521.3475
0.68313 | 1030204.1817 | 64.2592 | 155.1355 | 482.329
0.66832 | 982476.0022 | 60.6513 | 146.4252 | 498.2706
0.68822 | 1035659.8205 | 86.3309 | 208.4212 | 489.7412
0.66903 | 1011033.9796 | 71.9941 | 173.8092 | 476.8675
0.67293 | 1005910.1416 | 69.4108 | 167.5724 | 466.4275
0.6737 | 1012182.4702 | 69.7637 | 168.4244 | 486.3912
0.66135 | 955247.162 | 67.0694 | 161.9198 | 483.0105
0.68861 | 1013142.6792 | 84.4059 | 203.7738 | 478.8071
0.67507 | 1008932.9548 | 59.5705 | 143.8158 | 467.0502
0.70504 | 1052738.2612 | 63.6247 | 153.6036 | 479.6128
0.68932 | 1041402.9786 | 80.4704 | 194.2726 | 487.3541
0.69237 | 1040957.2703 | 73.3948 | 177.1907 | 484.9848
0.70253 | 1072049.4474 | 65.0225 | 156.9783 | 506.583
0.69874 | 1070887.478 | 58.8647 | 142.112 | 495.3037
0.69571 | 1038694.9368 | 65.576 | 158.3145 | 519.9734
0.68613 | 1035216.4278 | 68.1904 | 164.6261 | 487.859
0.665 | 987475.9513 | 54.7958 | 132.2889 | 493.7917
0.69835 | 1076002.2913 | 67.6737 | 163.3787 | 478.0729
0.7093 | 1090047.7171 | 80.2028 | 193.6268 | 511.7898
0.67528 | 1017908.0703 | 81.8811 | 197.6785 | 490.7562
0.66043 | 979350.3133 | 82.0854 | 198.1717 | 500.6201
0.69003 | 1046605.9425 | 72.2592 | 174.4492 | 510.8568
0.6886 | 1030718.4924 | 69.504 | 167.7976 | 496.9288
0.68633 | 1048990.2129 | 102.9815 | 248.6195 | 487.8154
0.6751 | 1023111.7444 | 79.0908 | 190.942 | 504.9903
0.70613 | 1077895.6145 | 74.519 | 179.9048 | 473.5136
0.68191 | 1027769.3656 | 68.1185 | 164.4526 | 488.7074
0.68913 | 1046473.6182 | 70.512 | 170.231 | 495.1005
0.6826 | 1009519.6997 | 72.1977 | 174.3006 | 499.5115
0.67634 | 1016969.7228 | 73.4029 | 177.2103 | 503.3291
0.69077 | 1025393.1977 | 54.7772 | 132.2439 | 505.1626
0.69239 | 1051188.4273 | 87.0971 | 210.2711 | 503.6135
0.69933 | 1073475.0529 | 76.6984 | 185.1664 | 469.4001
0.67926 | 1015449.979 | 74.6965 | 180.3332 | 492.5418
0.68896 | 1040212.3506 | 66.3146 | 160.0977 | 504.0224
0.69342 | 1037205.0395 | 78.0628 | 188.4604 | 474.1289
0.69557 | 1055676.1787 | 79.2064 | 191.2212 | 494.079
0.66833 | 993545.4043 | 61.3335 | 148.0721 | 493.0766
0.67332 | 986675.2151 | 73.2201 | 176.7691 | 505.1602
0.68891 | 1058225.583 | 78.478 | 189.4626 | 492.1826
0.70238 | 1064663.2992 | 80.7387 | 194.9204 | 495.5341
0.68789 | 1055081.3772 | 67.3677 | 162.64 | 492.1818
0.69761 | 1065546.2394 | 82.9696 | 200.3062 | 517.1187
0.68968 | 1021870.7342 | 70.5875 | 170.4134 | 499.08
0.69152 | 1039472.4025 | 62.017 | 149.7223 | 490.8007
0.67555 | 1021216.9306 | 62.3681 | 150.5699 | 472.0408
0.69555 | 1051568.9995 | 66.4038 | 160.3129 | 490.3781
0.65736 | 954203.3511 | 65.6048 | 158.3841 | 459.3499
0.70481 | 1092072.1726 | 75.6407 | 182.6129 | 473.2222
0.68966 | 1042827.2908 | 78.9136 | 190.5142 | 482.1648
0.66304 | 961806.9331 | 63.7134 | 153.8178 | 500.1632
0.69597 | 1044028.8353 | 91.8036 | 221.6335 | 510.5921
0.68662 | 1027278.8338 | 63.7511 | 153.9088 | 502.2537
0.70423 | 1084130.4642 | 76.9211 | 185.7041 | 491.009
0.69312 | 1036789.1641 | 84.3209 | 203.5688 | 478.8464
0.70598 | 1083374.4393 | 81.7385 | 197.3342 | 486.6642
0.71282 | 1101073.2173 | 78.916 | 190.52 | 495.2066
0.66947 | 996129.2182 | 76.0179 | 183.5235 | 473.8384
0.69773 | 1063083.4603 | 64.9228 | 156.7375 | 492.6978
0.70706 | 1085416.9931 | 54.2359 | 130.937 | 498.2943
0.69727 | 1049157.7486 | 71.6836 | 173.0595 | 490.5283
0.68658 | 1013818.0618 | 60.921 | 147.0762 | 525.9194
0.70958 | 1079837.9993 | 70.6638 | 170.5976 | 485.8027
0.69316 | 1061514.2677 | 67.2625 | 162.3861 | 499.984
0.679 | 1014535.4085 | 82.7088 | 199.6767 | 458.8328
0.68897 | 1038268.5197 | 81.8595 | 197.6262 | 458.2394
0.67582 | 989020.7485 | 77.9499 | 188.1878 | 478.4812
0.68046 | 1012424.0337 | 76.0138 | 183.5136 | 495.723
0.68135 | 1024153.2494 | 64.049 | 154.6281 | 497.2731
0.67522 | 1018932.718 | 60.1476 | 145.2093 | 500.5325
0.70627 | 1092107.1435 | 78.9993 | 190.7212 | 528.6421
0.69193 | 1041392.6062 | 76.8919 | 185.6336 | 474.109
0.66003 | 973225.3708 | 68.468 | 165.2965 | 481.5922
0.6929 | 1057505.8822 | 74.5054 | 179.8719 | 480.8024
0.69161 | 1040840.7661 | 73.4682 | 177.3679 | 470.1889
0.6974 | 1062545.9359 | 68.8591 | 166.2407 | 496.9134
0.69793 | 1062436.295 | 70.685 | 170.6486 | 476.622
0.67716 | 1010028.7849 | 71.0589 | 171.5513 | 492.67
0.66265 | 973459.1069 | 73.4026 | 177.2096 | 490.7695
0.6685 | 996565.5879 | 85.5222 | 206.4689 | 487.8091
0.66343 | 993093.0266 | 73.0753 | 176.4193 | 492.9734
0.71402 | 1104533.727 | 80.7311 | 194.9021 | 488.2231
0.691 | 1044537.5635 | 80.5169 | 194.3851 | 473.7242
0.69686 | 1052927.1527 | 64.3831 | 155.4344 | 499.0408
0.705 | 1079780.258 | 78.2488 | 188.9094 | 494.7062
0.67818 | 1006943.4141 | 72.4824 | 174.988 | 497.5978
0.69632 | 1050092.9983 | 75.3518 | 181.9154 | 497.5635
0.70445 | 1075087.5014 | 58.6161 | 141.5118 | 473.0877
0.68148 | 1031561.8729 | 83.7912 | 202.2898 | 489.3021
0.70669 | 1082806.4134 | 87.9018 | 212.2137 | 484.6697
0.67672 | 997036.9807 | 49.6308 | 119.8193 | 468.548
0.67347 | 995862.9254 | 64.9474 | 156.7968 | 496.3511
0.68048 | 1015197.1158 | 65.5262 | 158.1943 | 517.3646
0.67891 | 1028446.4641 | 79.0571 | 190.8608 | 489.6661
0.68865 | 1047138.7745 | 75.9771 | 183.425 | 511.1682
0.68812 | 1028785.8389 | 65.6669 | 158.5339 | 501.6531
0.69237 | 1044905.2184 | 80.3858 | 194.0685 | 501.351
0.70843 | 1092144.1946 | 71.585 | 172.8215 | 501.3074
0.67837 | 1019588.9896 | 71.6035 | 172.8661 | 447.4554
0.66747 | 993808.4528 | 70.9332 | 171.248 | 502.359
0.69933 | 1060631.5204 | 74.0383 | 178.7443 | 496.9025
0.67706 | 1012360.8769 | 72.9935 | 176.222 | 488.1431
0.69245 | 1038729.1032 | 67.7966 | 163.6756 | 518.8394
0.69555 | 1060578.0853 | 72.6427 | 175.375 | 469.9722
0.7244 | 1104530.8013 | 84.4024 | 203.7654 | 501.1376
0.6816 | 1015026.6523 | 72.6973 | 175.5068 | 479.5881
0.67887 | 1028026.5097 | 83.8249 | 202.3713 | 471.0892
0.70884 | 1071996.8475 | 73.828 | 178.2365 | 492.2985
0.69555 | 1048288.3841 | 78.6549 | 189.8896 | 490.0763
0.68859 | 1042591.0184 | 68.4301 | 165.205 | 499.867
0.68507 | 1029080.3713 | 80.0119 | 193.1658 | 509.6653
0.67328 | 1005609.9057 | 67.5035 | 162.968 | 489.7378
0.6905 | 1048326.0557 | 56.0241 | 135.2542 | 462.6402
0.70861 | 1098694.7626 | 66.0677 | 159.5014 | 485.5023
0.69281 | 1036504.9303 | 55.2077 | 133.2832 | 508.6159
0.69266 | 1057693.4115 | 87.6016 | 211.4889 | 458.2572
0.69938 | 1068612.5843 | 105.4927 | 254.6819 | 505.4291
0.67751 | 1010095.8714 | 77.8071 | 187.8429 | 510.8272
0.70211 | 1070857.5514 | 62.9255 | 151.9156 | 500.7826
0.71113 | 1089525.3943 | 84.8095 | 204.7481 | 483.1409
0.70552 | 1079676.4249 | 79.8652 | 192.8115 | 501.6216
0.70747 | 1085924.1166 | 71.2741 | 172.071 | 465.226
0.67083 | 1004376.0209 | 73.0895 | 176.4536 | 503.1734
0.70928 | 1083064.108 | 78.0129 | 188.3397 | 495.2327
0.676 | 1012662.8643 | 55.279 | 133.4554 | 489.2021
0.69978 | 1060815.9058 | 74.004 | 178.6614 | 491.8493
0.69375 | 1032812.5864 | 77.6951 | 187.5725 | 476.8233
0.67945 | 1024826.0378 | 66.2485 | 159.9379 | 478.1507
0.69922 | 1050379.8917 | 71.5661 | 172.7758 | 484.7958
0.6697 | 992596.7298 | 72.355 | 174.6803 | 478.5446
0.68357 | 1012541.0542 | 77.3037 | 186.6277 | 468.3487
0.7072 | 1088565.3969 | 89.8145 | 216.8313 | 496.5414
0.67865 | 1002220.2691 | 80.2187 | 193.6651 | 484.9629
0.69564 | 1039851.8982 | 74.6899 | 180.3174 | 484.2244
0.70035 | 1064498.9603 | 75.9467 | 183.3515 | 454.2286
0.68882 | 1018108.6209 | 73.8522 | 178.295 | 508.7964
0.67201 | 994729.919 | 72.2606 | 174.4524 | 497.8944
0.67 | 997720.819 | 59.3547 | 143.295 | 504.2994
0.6724 | 1015946.3055 | 77.1924 | 186.3588 | 486.1318
0.70036 | 1068394.4934 | 79.6125 | 192.2015 | 499.7995
0.70402 | 1075308.2448 | 74.5631 | 180.0111 | 512.4209
0.67665 | 1018245.6197 | 81.7418 | 197.3421 | 476.792
0.69852 | 1050689.147 | 76.847 | 185.5251 | 509.4454
0.69111 | 1035607.6109 | 69.58 | 167.981 | 515.2685
0.69596 | 1062492.8409 | 56.3355 | 136.0059 | 490.0895
0.68068 | 1025570.4006 | 60.0534 | 144.9818 | 472.2297
0.69163 | 1054798.3947 | 70.8163 | 170.9657 | 478.8188
0.69034 | 1039106.5851 | 71.7089 | 173.1206 | 488.6118
0.71128 | 1073986.8271 | 66.3192 | 160.1088 | 476.8645
0.71132 | 1097091.5404 | 67.7426 | 163.5452 | 491.2972
0.70725 | 1068262.0625 | 68.2651 | 164.8065 | 525.3527
0.69484 | 1055830.808 | 79.302 | 191.4521 | 487.829
0.68405 | 1025837.7365 | 78.2641 | 188.9463 | 510.9648
0.69211 | 1041090.1367 | 69.8476 | 168.627 | 501.2984
0.68144 | 1017947.0209 | 65.6118 | 158.4008 | 511.9564
0.6879 | 1045658.1832 | 62.7137 | 151.4042 | 489.0348
0.67928 | 1012532.5952 | 61.4978 | 148.4687 | 478.2911
0.69527 | 1050975.6705 | 72.0134 | 173.8558 | 488.5831
0.71537 | 1116501.9451 | 62.2082 | 150.1839 | 493.9941
0.69695 | 1066797.8693 | 71.0749 | 171.5901 | 459.2944
0.68528 | 1042421.8914 | 87.2671 | 210.6815 | 465.0221
0.67986 | 1024254.234 | 73.546 | 177.5557 | 481.2604
0.70172 | 1057920.6025 | 80.3021 | 193.8665 | 498.1758
0.69307 | 1030336.6992 | 63.3461 | 152.9311 | 503.0991
0.69969 | 1053761.119 | 68.0665 | 164.3272 | 501.9208
0.69988 | 1071688.647 | 87.8457 | 212.0782 | 490.8617
0.68283 | 1021548.6717 | 69.4677 | 167.7098 | 484.3287
0.6868 | 1027834.2095 | 76.4583 | 184.5866 | 487.651
0.65587 | 959507.9402 | 86.7744 | 209.4919 | 480.8989
0.69556 | 1063467.3945 | 74.6383 | 180.1928 | 475.9595
0.67476 | 995944.3593 | 68.5369 | 165.4627 | 491.2847
0.71338 | 1087911.1042 | 74.447 | 179.731 | 486.7962
0.69932 | 1063895.9321 | 60.6373 | 146.3914 | 495.878
0.71409 | 1097795.8909 | 65.5063 | 158.1461 | 487.2706
0.69174 | 1048465.3062 | 53.5495 | 129.28 | 498.8673
0.68014 | 1020147.7018 | 83.0266 | 200.444 | 491.5456
0.66221 | 982107.1443 | 75.7694 | 182.9234 | 495.3856
0.68937 | 1030144.7521 | 67.5374 | 163.0496 | 492.4384
0.70332 | 1060520.8235 | 66.2293 | 159.8916 | 498.604
0.68668 | 1035819.835 | 77.0654 | 186.0522 | 487.9285
0.67665 | 1010216.1269 | 84.1594 | 203.1788 | 484.8036
0.67224 | 1006745.405 | 67.6851 | 163.4062 | 547.4227
0.70574 | 1075565.9034 | 70.7667 | 170.846 | 473.6907
0.69188 | 1038920.758 | 81.5657 | 196.9169 | 471.4251
0.6923 | 1049653.1427 | 83.9272 | 202.6181 | 489.2331
0.69147 | 1044647.1918 | 83.6282 | 201.8963 | 477.9587
0.70159 | 1075171.7509 | 69.2609 | 167.2106 | 466.7773
0.68584 | 1023371.1427 | 63.5807 | 153.4974 | 496.3329
0.70452 | 1089852.586 | 73.6977 | 177.9219 | 505.3798
0.66512 | 981494.7778 | 66.6936 | 161.0126 | 474.9454
0.69187 | 1055763.0469 | 84.469 | 203.9262 | 510.5106
0.68379 | 1023476.7561 | 57.4881 | 138.7885 | 510.335
0.67858 | 1005013.2622 | 74.0021 | 178.6569 | 489.3368
0.69843 | 1065731.6374 | 84.4655 | 203.9177 | 478.21
0.68637 | 1037959.7668 | 89.6974 | 216.5486 | 496.4444
0.67754 | 995616.8167 | 75.0738 | 181.2442 | 495.9903
0.6839 | 1043814.9202 | 87.5602 | 211.389 | 470.3498
0.70016 | 1059558.6955 | 62.4351 | 150.7316 | 480.1204
0.72147 | 1122046.5496 | 72.2154 | 174.3434 | 516.2159
0.68008 | 1021660.2999 | 53.0751 | 128.1347 | 499.1901
0.70441 | 1077231.9461 | 90.7074 | 218.987 | 463.0027
0.68651 | 1032740.6532 | 83.2716 | 201.0353 | 499.6128
0.70499 | 1079553.5272 | 58.8149 | 141.9916 | 508.095
0.70657 | 1064337.1746 | 55.5858 | 134.1959 | 485.0496
0.67991 | 1017114.0369 | 69.4636 | 167.7001 | 484.9155
0.68301 | 1018494.3408 | 67.395 | 162.7059 | 471.5609
0.67823 | 1021191.3253 | 81.7881 | 197.454 | 500.8642
0.66348 | 982694.8328 | 63.1116 | 152.3649 | 492.8143
0.68824 | 1023493.0489 | 79.2107 | 191.2315 | 491.4861
0.69524 | 1040436.0417 | 73.0553 | 176.371 | 475.3843
0.69573 | 1074508.8152 | 71.7146 | 173.1344 | 479.0257
0.71092 | 1070591.3835 | 64.168 | 154.9154 | 468.6511
0.68034 | 1010600.4454 | 93.6206 | 226.0202 | 457.3592
0.68234 | 1010905.8615 | 74.0265 | 178.7158 | 481.3335
0.69235 | 1058433.4979 | 72.8311 | 175.8299 | 488.7313
0.70895 | 1087178.158 | 71.2071 | 171.9092 | 482.3741
0.6877 | 1028377.413 | 82.7771 | 199.8415 | 515.4948
0.68633 | 1031939.4429 | 80.0867 | 193.3463 | 524.5982
0.67723 | 1026413.8265 | 80.7283 | 194.8954 | 463.3039
0.65214 | 951899.4287 | 72.9331 | 176.0761 | 472.2007
0.69005 | 1045328.3922 | 83.8543 | 202.4421 | 465.0874
0.69748 | 1044423.1203 | 67.6766 | 163.3858 | 485.785
0.68737 | 1044725.9929 | 77.6505 | 187.4649 | 501.6762
0.71336 | 1090231.4455 | 72.8093 | 175.7771 | 499.6946
0.69687 | 1061859.9718 | 76.1366 | 183.8101 | 470.9806
0.6595 | 958503.9459 | 91.6008 | 221.1439 | 474.3673
0.68846 | 1037159.6054 | 89.3568 | 215.7264 | 502.579
0.6941 | 1050801.9877 | 73.9046 | 178.4215 | 481.3557
0.68855 | 1063836.9434 | 77.2 | 186.3772 | 509.1636
0.68061 | 1013298.4633 | 82.412 | 198.9602 | 485.7922
0.67271 | 1000362.9498 | 89.616 | 216.3521 | 498.9483
0.68945 | 1037843.5208 | 78.5606 | 189.6621 | 498.4962
0.69669 | 1045988.3895 | 59.8416 | 144.4704 | 476.4803
0.71245 | 1101566.327 | 71.795 | 173.3284 | 480.3595
0.68997 | 1049324.5481 | 84.1849 | 203.2404 | 453.4542
0.70543 | 1069060.4154 | 79.0901 | 190.9404 | 469.7832
0.6798 | 1012794.1959 | 77.7519 | 187.7096 | 501.9286
0.6778 | 1010901.2811 | 76.1966 | 183.9548 | 497.4098
0.66198 | 978481.4814 | 70.4186 | 170.0056 | 480.5575
0.68183 | 1028104.7071 | 71.8359 | 173.4271 | 477.6304
0.69822 | 1045709.9793 | 79.774 | 192.5915 | 505.1807
0.67934 | 1032675.5473 | 68.7024 | 165.8624 | 482.4427
0.72112 | 1124925.3707 | 79.6983 | 192.4087 | 462.7776
0.71159 | 1095391.4231 | 80.3361 | 193.9486 | 487.3048
0.68655 | 1039453.3484 | 82.2414 | 198.5483 | 490.1792
0.68299 | 1036482.7889 | 70.4667 | 170.1216 | 473.1247
0.67418 | 999831.1216 | 77.3967 | 186.8521 | 478.4334
0.67549 | 1001970.539 | 61.1984 | 147.7459 | 502.4978
0.69997 | 1054321.7883 | 64.1816 | 154.9481 | 495.0872
0.67248 | 994696.8657 | 70.4674 | 170.1233 | 491.5238
0.72282 | 1130096.3235 | 62.3167 | 150.4457 | 501.4013
0.69668 | 1054856.51 | 86.8628 | 209.7055 | 504.6481
0.69397 | 1047687.9098 | 83.5689 | 201.7532 | 474.6542
0.64828 | 953011.8242 | 58.5939 | 141.4583 | 499.5408
0.67938 | 1016755.1297 | 81.2943 | 196.2617 | 483.1357
0.67349 | 1002177.0127 | 72.6417 | 175.3727 | 488.5157
0.67691 | 1013558.4941 | 76.4982 | 184.6829 | 492.1232
0.68533 | 1038237.3088 | 83.8591 | 202.4538 | 483.2263
0.69547 | 1055287.3889 | 75.8169 | 183.0382 | 477.9702
0.70604 | 1072898.5536 | 64.7511 | 156.323 | 459.4699
0.72465 | 1133957.6299 | 50.0078 | 120.7296 | 482.8404
0.67661 | 1017750.0278 | 73.4303 | 177.2763 | 459.6441
0.70406 | 1080789.0213 | 88.0657 | 212.6093 | 486.6702
0.68769 | 1042236.0333 | 75.7267 | 182.8205 | 487.3349
0.69986 | 1062499.2951 | 70.4944 | 170.1885 | 512.6086
0.71387 | 1094207.0449 | 58.8717 | 142.1287 | 506.2791
0.68921 | 1040666.4109 | 83.5817 | 201.7841 | 528.8951
0.71519 | 1099642.3097 | 71.3954 | 172.3637 | 484.3241
0.68359 | 1038309.5397 | 77.6004 | 187.3438 | 515.3547
0.70025 | 1071604.5392 | 78.1824 | 188.7491 | 498.892
0.6989 | 1059009.5748 | 68.4827 | 165.3318 | 519.7046
0.7051 | 1068757.3738 | 60.4152 | 145.8552 | 496.4985
0.69345 | 1059791.2082 | 72.8086 | 175.7755 | 507.4742
0.70297 | 1064707.8714 | 70.3285 | 169.7881 | 471.2301
0.68682 | 1033053.4815 | 76.8445 | 185.519 | 469.0229
0.68999 | 1043370.1915 | 80.5447 | 194.4521 | 486.1395
0.70657 | 1073894.8265 | 66.7559 | 161.1629 | 504.5057
0.67651 | 1006031.9827 | 72.596 | 175.2622 | 472.1578
0.67808 | 1022730.895 | 57.9299 | 139.8552 | 485.4567
0.68182 | 1007773.9232 | 64.9198 | 156.7302 | 490.9309
0.68846 | 1047079.2936 | 79.3737 | 191.6251 | 489.7245
0.68197 | 1022662.1841 | 79.8739 | 192.8326 | 486.6955
0.71482 | 1097647.8821 | 80.1382 | 193.4707 | 495.7032
0.69278 | 1034292.8252 | 63.6156 | 153.5817 | 492.9117
0.67907 | 1012640.1791 | 72.7967 | 175.7468 | 487.4818
0.71961 | 1110410.3058 | 73.5124 | 177.4747 | 488.0611
0.69412 | 1050063.6596 | 92.5799 | 223.5076 | 479.178
0.66862 | 993235.2272 | 75.8927 | 183.2212 | 536.1486
0.71339 | 1085255.0315 | 85.7545 | 207.0297 | 473.0454
0.69313 | 1050032.8343 | 64.5516 | 155.8412 | 506.0532
0.67129 | 1010452.8695 | 71.711 | 173.1257 | 519.3955
0.66647 | 978682.5331 | 76.6368 | 185.0177 | 450.0348
0.67564 | 1009193.4258 | 69.3575 | 167.4439 | 469.594
0.71058 | 1074228.3986 | 75.5367 | 182.3618 | 494.6482
0.69193 | 1031520.0192 | 72.7763 | 175.6976 | 522.4197
0.7037 | 1068759.2217 | 92.9461 | 224.3917 | 476.5091
0.68313 | 1035947.8966 | 66.6064 | 160.802 | 491.6371
0.69108 | 1038515.0317 | 84.7105 | 204.5092 | 492.2013
0.71791 | 1116529.4254 | 74.0927 | 178.8757 | 488.7372
0.69089 | 1037935.2705 | 71.5714 | 172.7886 | 476.7518
0.67692 | 997411.7286 | 71.9304 | 173.6554 | 474.8083
0.70008 | 1063933.2389 | 68.6855 | 165.8215 | 496.5981
0.69608 | 1061617.7156 | 88.9906 | 214.8422 | 473.9553
0.68038 | 1026864.3092 | 83.5436 | 201.6922 | 456.3213
0.68309 | 1022615.4969 | 87.3523 | 210.8872 | 492.9312
0.69322 | 1053462.025 | 95.7772 | 231.2267 | 505.8999
0.69228 | 1047132.0304 | 77.2774 | 186.5641 | 487.2083
0.68228 | 1022997.8908 | 71.4316 | 172.4511 | 495.9837
0.68766 | 1027599.2746 | 81.1727 | 195.9683 | 503.7075
0.67417 | 991475.7608 | 54.269 | 131.017 | 473.0849
0.68499 | 1023859.9046 | 79.6152 | 192.2081 | 506.0935
0.67237 | 1005316.2874 | 72.9905 | 176.2146 | 511.4862
0.68543 | 1044369.5449 | 77.8434 | 187.9307 | 497.9681
0.68989 | 1049170.238 | 70.616 | 170.482 | 511.0672
0.69881 | 1067110.0689 | 70.9064 | 171.1833 | 484.13
0.70284 | 1073523.5277 | 58.6651 | 141.6302 | 489.7556
0.688 | 1044374.9735 | 72.3175 | 174.5899 | 513.6983
0.6905 | 1072099.6984 | 73.9689 | 178.5768 | 465.9101
0.70865 | 1090449.4882 | 67.5857 | 163.1664 | 479.1798
0.68424 | 1031668.8024 | 68.1441 | 164.5143 | 495.1056
0.6963 | 1059147.1968 | 76.0042 | 183.4904 | 501.0476
0.69238 | 1065140.3463 | 71.4328 | 172.454 | 504.1691
0.69909 | 1062405.994 | 80.9457 | 195.4202 | 496.1545
0.70135 | 1061718.7444 | 64.5046 | 155.7278 | 500.7356
0.70766 | 1092319.8025 | 77.1261 | 186.1988 | 518.4558
0.70148 | 1062065.5528 | 64.5141 | 155.7508 | 477.2577
0.70246 | 1075382.4092 | 74.7332 | 180.4219 | 471.2224
0.68937 | 1039640.1585 | 73.6214 | 177.7379 | 499.4536
0.69922 | 1068450.6589 | 78.0671 | 188.4707 | 480.4534
0.69886 | 1063947.797 | 72.5282 | 175.0987 | 476.1701
0.69646 | 1040921.2185 | 72.9613 | 176.1442 | 514.0613
0.66044 | 973085.5134 | 64.7146 | 156.2348 | 494.7079
0.68629 | 1034901.4395 | 71.8404 | 173.4381 | 475.4537
0.65862 | 973653.77 | 68.6452 | 165.7241 | 482.1062
0.68687 | 1026419.5367 | 76.3455 | 184.3143 | 499.2736
0.70261 | 1055105.714 | 71.5417 | 172.7169 | 497.0877
0.69304 | 1055004.1935 | 79.4453 | 191.7978 | 493.4049
0.67925 | 1019934.0219 | 71.5353 | 172.7016 | 484.6332
0.69542 | 1059948.1929 | 68.8861 | 166.3057 | 506.1856
0.6845 | 1006391.3489 | 66.4494 | 160.4231 | 465.6284
0.67973 | 1034993.0434 | 60.69 | 146.5186 | 490.2749
0.70098 | 1044114.7167 | 72.7139 | 175.547 | 507.265
0.70638 | 1064418.3983 | 80.2624 | 193.7707 | 484.6956
0.70215 | 1059585.0809 | 78.7554 | 190.1324 | 485.5906
0.67473 | 1010190.9786 | 71.3268 | 172.1981 | 507.1532
0.68157 | 1031491.2632 | 66.4722 | 160.4781 | 484.7276
0.69748 | 1060492.5145 | 82.5911 | 199.3926 | 502.9833
0.6582 | 967705.1753 | 58.601 | 141.4752 | 516.277
0.68378 | 1018225.0172 | 70.9419 | 171.2689 | 487.1859
0.69843 | 1052422.1081 | 81.851 | 197.6059 | 521.3012
0.6709 | 979410.2949 | 69.5482 | 167.9042 | 486.1754
0.6755 | 1010270.8573 | 63.7281 | 153.8533 | 509.3774
0.70921 | 1101042.6069 | 70.1659 | 169.3954 | 462.04
0.6971 | 1046034.8613 | 74.9768 | 181.0099 | 506.8293
0.68697 | 1025590.4635 | 63.6144 | 153.5787 | 478.7549
0.69404 | 1046230.4198 | 77.0147 | 185.9299 | 503.195
0.70554 | 1068436.2877 | 72.1452 | 174.1739 | 512.0564
0.67781 | 1013628.3749 | 69.4399 | 167.6427 | 510.9868
0.68713 | 1030332.5662 | 67.7694 | 163.6098 | 502.3647
0.69109 | 1036669.0119 | 79.8992 | 192.8937 | 476.5752
0.6955 | 1071499.5917 | 73.3963 | 177.1943 | 468.26
0.67905 | 1008452.9249 | 77.5048 | 187.1131 | 516.3271
0.7159 | 1089471.755 | 72.5008 | 175.0324 | 472.0021
0.69836 | 1063037.8269 | 62.802 | 151.6173 | 485.7802
0.68768 | 1036574.4927 | 71.8111 | 173.3674 | 473.3075
0.68799 | 1026092.4259 | 72.8077 | 175.7734 | 480.4996
0.66954 | 988417.0935 | 70.33 | 169.7917 | 496.7762
0.68747 | 1028413.6851 | 79.6807 | 192.3661 | 469.0509
0.67965 | 1034736.7658 | 84.2754 | 203.4587 | 473.431
0.7069 | 1072594.9978 | 80.8194 | 195.1153 | 486.5242
0.70288 | 1070150.0214 | 76.0747 | 183.6607 | 504.5521
0.70002 | 1056041.4063 | 73.6449 | 177.7945 | 521.961
0.68607 | 1046059.6098 | 77.8963 | 188.0583 | 485.5813
0.69211 | 1045339.582 | 68.2616 | 164.798 | 505.6547
0.68317 | 1030033.2213 | 80.3084 | 193.8816 | 493.0192
0.67017 | 993962.5034 | 80.5637 | 194.498 | 499.9813
0.69257 | 1037902.5474 | 67.7313 | 163.5179 | 495.7384
0.68182 | 1012666.5166 | 82.2145 | 198.4835 | 473.1852
0.70153 | 1071924.5175 | 82.9991 | 200.3776 | 506.177
0.66658 | 989529.946 | 70.7224 | 170.739 | 477.5129
0.69303 | 1039145.927 | 83.0001 | 200.38 | 450.8871
0.66896 | 988280.3309 | 62.9499 | 151.9744 | 530.6632
0.69225 | 1051962.9647 | 58.3664 | 140.909 | 465.7578
0.69817 | 1054181.4702 | 70.1123 | 169.2661 | 512.3776
0.70086 | 1061861.055 | 75.7027 | 182.7625 | 484.6335
0.70924 | 1077260.4632 | 70.5289 | 170.2719 | 466.207
0.67447 | 1008522.3646 | 84.2905 | 203.4953 | 490.6305
0.69881 | 1071848.7074 | 69.3221 | 167.3585 | 492.2763
0.68634 | 1028749.959 | 67.7431 | 163.5464 | 490.6029
0.7061 | 1085022.291 | 76.7799 | 185.3632 | 497.3804
0.66414 | 977107.0527 | 49.5602 | 119.6489 | 504.0594
0.67652 | 1005450.5492 | 77.563 | 187.2536 | 456.9433
0.6774 | 1008203.1912 | 69.1097 | 166.8456 | 498.6272
0.7263 | 1125993.459 | 91.9553 | 221.9997 | 494.7043
0.68488 | 1038462.2299 | 77.4652 | 187.0174 | 500.4591
0.69181 | 1045792.0972 | 60.2527 | 145.463 | 498.0418
0.69358 | 1065143.8918 | 63.3593 | 152.9629 | 490.316
0.67094 | 987493.8535 | 60.7227 | 146.5976 | 509.0926
0.7165 | 1089930.7613 | 87.2684 | 210.6846 | 491.9409
0.68099 | 1019770.1089 | 76.5049 | 184.6992 | 502.1574
0.69998 | 1068757.6851 | 72.7397 | 175.6092 | 497.3331
0.70487 | 1086391.3858 | 67.6959 | 163.4324 | 494.7802
0.69817 | 1051008.1123 | 81.7717 | 197.4144 | 502.922
0.69603 | 1052217.3183 | 66.3991 | 160.3017 | 498.7615
0.68143 | 1030573.9746 | 77.2662 | 186.5371 | 470.1912
0.67392 | 1002463.9446 | 77.6308 | 187.4172 | 456.5734
0.69397 | 1045445.3025 | 68.086 | 164.3742 | 475.3396
0.71494 | 1104013.3282 | 75.3833 | 181.9914 | 493.2231
0.69827 | 1073732.909 | 67.5392 | 163.054 | 505.2301
0.69046 | 1049294.6441 | 70.5667 | 170.363 | 479.331
0.69843 | 1055604.224 | 88.1887 | 212.9064 | 485.3147
0.68058 | 1020078.2118 | 76.5745 | 184.8671 | 482.2446
0.68552 | 1028882.99 | 81.5272 | 196.8241 | 497.9932
0.67703 | 1014676.3026 | 73.1632 | 176.6316 | 528.5856
0.72159 | 1118643.923 | 70.0647 | 169.1511 | 478.8825
0.67852 | 1002921.8492 | 75.515 | 182.3094 | 476.0308
0.7142 | 1100886.8584 | 75.2484 | 181.6658 | 498.6733
0.68379 | 1021119.5141 | 65.5899 | 158.3481 | 509.6786
0.6669 | 1000569.0265 | 71.4761 | 172.5586 | 470.3597
0.71829 | 1108869.6736 | 77.1364 | 186.2237 | 473.5157
0.72043 | 1125128.2453 | 81.5051 | 196.7707 | 500.0962
0.70066 | 1063152.7752 | 68.5778 | 165.5613 | 467.3359
0.67754 | 997679.661 | 79.9856 | 193.1024 | 491.3951
0.71879 | 1090991.5772 | 70.4787 | 170.1506 | 460.0414
0.66939 | 999351.8246 | 67.4733 | 162.895 | 489.0603
0.69747 | 1047390.5031 | 65.9578 | 159.2363 | 479.3083
0.67064 | 1018397.2391 | 80.8846 | 195.2727 | 508.7129
0.69644 | 1045527.5274 | 69.7089 | 168.2921 | 487.927
0.67572 | 995021.0227 | 71.885 | 173.5457 | 490.9314
0.70238 | 1046459.5849 | 65.4385 | 157.9825 | 480.2835
0.69761 | 1047713.3459 | 76.7291 | 185.2405 | 474.318
0.67506 | 994828.2229 | 76.9097 | 185.6763 | 489.8886
0.67923 | 1013142.5555 | 69.2616 | 167.2122 | 487.1767
0.68654 | 1037279.4227 | 86.6772 | 209.2574 | 514.174
0.69231 | 1040552.1626 | 57.7278 | 139.3673 | 476.3549
0.69973 | 1070465.9664 | 65.1864 | 157.374 | 520.1017
0.71195 | 1113605.9668 | 78.1548 | 188.6824 | 487.9005
0.68885 | 1034874.632 | 76.7526 | 185.2971 | 471.6276
0.6877 | 1029564.7767 | 67.663 | 163.353 | 481.2503
0.69543 | 1043684.5905 | 73.2639 | 176.8747 | 491.8567
0.68474 | 1027093.4249 | 69.6497 | 168.1491 | 505.8036
0.683 | 1037552.4826 | 74.1585 | 179.0346 | 482.3273
0.70144 | 1068515.1322 | 77.8254 | 187.8873 | 497.6549
0.71161 | 1101225.6137 | 90.2965 | 217.9951 | 500.2222
0.68994 | 1040987.5079 | 71.8804 | 173.5347 | 495.2172
0.69358 | 1044372.65 | 76.9294 | 185.724 | 470.2729
0.65565 | 956657.0631 | 72.906 | 176.0108 | 476.6186
0.6888 | 1027140.3057 | 69.2429 | 167.1671 | 502.1741
0.67783 | 1009515.3671 | 67.311 | 162.5032 | 476.2126
0.68104 | 1027115.8648 | 90.8863 | 219.4189 | 503.7332
0.6931 | 1042244.8303 | 66.0472 | 159.452 | 500.957
0.72418 | 1123979.0308 | 82.6709 | 199.5853 | 522.3562
0.7026 | 1061509.0536 | 64.3772 | 155.4203 | 451.9943
0.66707 | 993100.2436 | 79.2972 | 191.4404 | 490.0386
0.71534 | 1106779.3116 | 67.9567 | 164.0621 | 493.797
0.70021 | 1070281.296 | 89.1879 | 215.3186 | 524.9289
0.70254 | 1067279.4167 | 77.6131 | 187.3747 | 513.189
0.67418 | 1009199.0068 | 61.7597 | 149.101 | 464.4973
0.66957 | 988312.0714 | 75.4099 | 182.0555 | 499.1659
0.70773 | 1094994.2123 | 77.3078 | 186.6374 | 493.2413
0.67519 | 1009947.5567 | 69.4561 | 167.6818 | 506.3663
0.70327 | 1063889.3648 | 68.371 | 165.0621 | 499.2412
0.68476 | 1019533.9627 | 77.7026 | 187.5908 | 472.1106
0.68591 | 1058666.7062 | 80.55 | 194.465 | 503.7691
0.65852 | 970330.4378 | 68.4213 | 165.1837 | 499.0833
0.70111 | 1061883.1987 | 82.2708 | 198.6192 | 499.6407
0.6757 | 1019314.7984 | 68.1793 | 164.5995 | 471.087
0.69822 | 1057204.2795 | 60.3506 | 145.6993 | 484.1384
0.69242 | 1033746.9827 | 74.597 | 180.0931 | 524.7355
0.71238 | 1085829.9928 | 62.8856 | 151.8192 | 470.2688
0.68141 | 1022080.9561 | 67.7914 | 163.6629 | 481.8792
0.67882 | 997055.4893 | 71.7029 | 173.1061 | 495.111
0.70995 | 1096002.5434 | 69.7721 | 168.4447 | 505.3314
0.70432 | 1069198.9131 | 66.6763 | 160.9708 | 531.9556
0.67726 | 1004134.9796 | 75.7468 | 182.8691 | 491.6885
0.69657 | 1055049.1079 | 73.9404 | 178.5079 | 517.2286
0.69047 | 1048736.4111 | 82.5045 | 199.1836 | 476.6275
0.7074 | 1083040.3105 | 76.4994 | 184.6859 | 491.5318
0.66788 | 1005683.9003 | 73.5078 | 177.4636 | 498.2373
0.708 | 1089778.1486 | 69.5661 | 167.9474 | 482.0076
0.69268 | 1046751.0222 | 80.491 | 194.3225 | 522.303
0.69871 | 1070304.6317 | 67.2807 | 162.4301 | 486.4576
0.67304 | 1022265.6373 | 68.6467 | 165.7278 | 489.0254
0.7135 | 1082951.4643 | 82.4467 | 199.0439 | 471.274
0.66177 | 970746.6835 | 76.5544 | 184.8186 | 490.5365
0.69562 | 1056923.6378 | 75.9182 | 183.2827 | 506.665
0.70842 | 1091299.378 | 59.1664 | 142.8402 | 510.6212
0.68786 | 1045888.3617 | 64.7892 | 156.4149 | 518.3237
0.67529 | 1013103.9317 | 73.2378 | 176.8118 | 483.7236
0.68262 | 1037742.8038 | 79.3386 | 191.5404 | 513.5089
0.68322 | 1019033.5954 | 71.0565 | 171.5456 | 492.7618
0.70313 | 1067211.6845 | 56.0088 | 135.2173 | 487.4671
0.68527 | 1016613.9412 | 65.1595 | 157.3089 | 507.5789
0.68991 | 1032811.3516 | 74.8001 | 180.5835 | 487.2187
0.71062 | 1082216.0201 | 77.805 | 187.8378 | 488.1731
0.68794 | 1043112.22 | 58.2378 | 140.5985 | 498.7301
0.70518 | 1058857.6213 | 57.6718 | 139.232 | 505.7572
0.64413 | 938185.7395 | 65.828 | 158.9229 | 479.3671
0.67238 | 1012244.6784 | 62.1112 | 149.9498 | 478.0811
0.68883 | 1036762.0466 | 73.4746 | 177.3835 | 487.7526
0.6674 | 997248.6577 | 63.6702 | 153.7136 | 477.8415
0.67877 | 1021232.318 | 76.866 | 185.5709 | 489.0034
0.70801 | 1083904.6403 | 79.8962 | 192.8865 | 488.4719
0.66418 | 984331.5092 | 86.4757 | 208.7708 | 471.723
0.71201 | 1091342.0996 | 87.9326 | 212.2881 | 497.6215
0.66732 | 988892.161 | 66.9749 | 161.6916 | 477.2843
0.68662 | 1034523.5894 | 57.8695 | 139.7092 | 487.4028
0.69239 | 1041921.3572 | 69.8752 | 168.6938 | 492.0016
0.69861 | 1072557.7642 | 78.4358 | 189.3607 | 449.9056
0.68447 | 1033841.0038 | 60.2032 | 145.3435 | 481.8288
0.67826 | 1024906.7065 | 70.1662 | 169.3961 | 500.6917
0.68114 | 1026613.1082 | 68.1186 | 164.4529 | 511.419
0.69007 | 1032099.0283 | 65.6922 | 158.595 | 477.6836
0.69125 | 1049769.6769 | 77.5876 | 187.3131 | 466.709
0.67634 | 1020507.5205 | 77.0866 | 186.1036 | 492.6425
0.69556 | 1057132.3171 | 74.1938 | 179.1197 | 463.7575
0.687 | 1049022.0562 | 84.5242 | 204.0594 | 491.6201
0.68341 | 1028238.6566 | 62.0747 | 149.8617 | 495.339
0.69088 | 1041006.1779 | 64.2052 | 155.005 | 553.4741
0.71393 | 1088904.6822 | 70.8225 | 170.9807 | 488.8143
0.6871 | 1034229.9912 | 76.9001 | 185.6533 | 496.9641
0.68455 | 1036283.8137 | 75.6397 | 182.6104 | 513.814
0.68942 | 1036201.0849 | 62.0095 | 149.7043 | 483.8805
0.692 | 1048329.47 | 86.9094 | 209.8178 | 483.1781
0.69207 | 1056460.4528 | 70.8824 | 171.1254 | 461.9606
0.6892 | 1041367.1354 | 70.1613 | 169.3844 | 501.2996
0.69822 | 1072087.0757 | 76.7212 | 185.2214 | 480.7216
0.67366 | 987883.6301 | 84.1279 | 203.1027 | 510.2089
0.68447 | 1026128.5454 | 76.5911 | 184.9074 | 510.988
0.68453 | 1031581.9342 | 86.4539 | 208.7183 | 470.8411
0.71946 | 1123626.3319 | 68.5962 | 165.6058 | 497.4544
0.67103 | 1005139.4558 | 74.8444 | 180.6903 | 487.5183
0.70316 | 1073830.1157 | 81.3436 | 196.3808 | 467.0076
0.70276 | 1054404.2343 | 73.2438 | 176.8262 | 482.1
0.70155 | 1059455.4789 | 70.1229 | 169.2918 | 461.5677
0.68965 | 1024593.5799 | 88.1218 | 212.7449 | 491.4145
0.70505 | 1073807.5923 | 64.8966 | 156.6743 | 493.2258
0.67145 | 1020647.1028 | 72.417 | 174.8302 | 494.3608
0.68841 | 1027183.0155 | 81.0513 | 195.6752 | 497.0268
0.69743 | 1066080.132 | 84.6143 | 204.277 | 505.3686
0.70565 | 1091431.9759 | 91.7412 | 221.4827 | 493.5036
0.68629 | 1030493.9478 | 83.1178 | 200.6642 | 510.9762
0.71102 | 1103465.6629 | 61.8577 | 149.3377 | 460.3543
0.70003 | 1067724.1608 | 85.9796 | 207.5731 | 479.8653
0.69641 | 1041798.3513 | 71.2679 | 172.056 | 527.2146
0.69639 | 1045332.7789 | 76.6886 | 185.1426 | 501.4707
0.69769 | 1064522.5555 | 72.6897 | 175.4884 | 488.5544
0.68005 | 1015997.0446 | 79.5772 | 192.1163 | 519.7924
0.70405 | 1070025.4176 | 75.9389 | 183.3328 | 479.3721
0.70519 | 1058301.9142 | 81.519 | 196.8044 | 513.278
0.6934 | 1033548.305 | 65.4678 | 158.0532 | 486.6077
0.69278 | 1033969.0688 | 65.5534 | 158.2598 | 503.631
0.70488 | 1085176.783 | 73.3432 | 177.0662 | 514.5158
0.6947 | 1034687.9824 | 50.1716 | 121.1249 | 516.3865
0.69685 | 1060746.7915 | 77.768 | 187.7486 | 492.2926
0.69943 | 1061334.0227 | 68.5814 | 165.5702 | 485.2926
0.68622 | 1031939.8689 | 73.4115 | 177.2311 | 507.052
0.70368 | 1075691.9873 | 60.9297 | 147.0974 | 500.4624
0.68893 | 1054875.3961 | 67.2529 | 162.3628 | 490.1644
0.69081 | 1044289.1368 | 75.5335 | 182.3541 | 489.3866
0.68703 | 1030468.6394 | 69.6465 | 168.1415 | 478.4777
0.69857 | 1067026.8001 | 67.0886 | 161.9662 | 475.204
0.67748 | 1029320.6676 | 68.6672 | 165.7772 | 518.5039
0.67736 | 1016565.4009 | 65.9172 | 159.1381 | 505.8681
0.6817 | 1013806.2644 | 74.3081 | 179.3957 | 490.2889
0.71044 | 1082625.9999 | 66.9324 | 161.5892 | 460.8214
0.67628 | 1020080.6956 | 70.0395 | 169.0904 | 471.5009
0.68003 | 1014550.5667 | 76.2212 | 184.0143 | 499.3908
0.66792 | 982225.6165 | 71.7891 | 173.3143 | 506.5593
0.70525 | 1081957.2778 | 76.656 | 185.064 | 510.3097
0.67424 | 994045.7653 | 66.6324 | 160.8648 | 472.6164
0.71386 | 1101821.7968 | 87.0057 | 210.0505 | 478.6107
0.68729 | 1006359.7798 | 69.7579 | 168.4104 | 501.5154
0.68 | 1003330.0669 | 66.6859 | 160.9939 | 481.0898
0.68523 | 1010328.8812 | 81.6165 | 197.0397 | 461.6918
0.70887 | 1094890.0013 | 78.9101 | 190.5059 | 478.5556
0.673 | 998630.7095 | 74.1428 | 178.9965 | 501.7243
0.65982 | 957412.8711 | 80.3204 | 193.9107 | 509.2357
0.70523 | 1076397.4542 | 60.8362 | 146.8716 | 483.8163
0.69385 | 1053373.0086 | 80.8642 | 195.2235 | 485.3567
0.69859 | 1068131.6707 | 78.115 | 188.5864 | 524.6013
0.67774 | 1025526.4434 | 70.2626 | 169.629 | 480.6606
0.69568 | 1058505.4951 | 68.7672 | 166.0187 | 512.4291
0.68136 | 1019331.2061 | 61.7224 | 149.011 | 499.0971
0.69854 | 1042484.4823 | 84.1257 | 203.0974 | 483.8122
0.68692 | 1038930.9011 | 64.4464 | 155.5875 | 440.7928
0.69275 | 1040362.388 | 67.8738 | 163.8618 | 468.9786
0.71229 | 1101994.0689 | 79.0409 | 190.8216 | 494.0362
0.6854 | 1042714.0706 | 77.7234 | 187.6408 | 499.4509
0.71165 | 1091041.1981 | 65.3268 | 157.7128 | 481.367
0.7055 | 1090335.2583 | 82.3189 | 198.7354 | 491.1262
0.66742 | 1002578.4593 | 73.332 | 177.0392 | 465.7262
0.68031 | 1024223.909 | 81.7791 | 197.4322 | 476.1047
0.68209 | 1026624.4428 | 62.097 | 149.9154 | 509.6865
0.68733 | 1040190.0161 | 58.1262 | 140.329 | 520.496
0.65991 | 968608.3594 | 73.3403 | 177.0592 | 468.9509
0.68886 | 1039995.8609 | 64.331 | 155.3087 | 463.3869
0.69114 | 1046191.5658 | 73.8957 | 178.4 | 486.7843
0.6891 | 1049688.5857 | 56.2112 | 135.7059 | 461.8604
0.6694 | 1009775.4071 | 78.7254 | 190.06 | 477.1013
0.68927 | 1041891.0858 | 83.6739 | 202.0066 | 492.6528
0.68575 | 1032230.4412 | 81.7654 | 197.3991 | 501.4768
0.67861 | 997562.7085 | 84.3325 | 203.5967 | 485.1002
0.67782 | 1029352.5023 | 84.0628 | 202.9456 | 500.829
0.66757 | 996763.5857 | 71.133 | 171.7304 | 480.928
0.67973 | 1009809.8617 | 70.7754 | 170.8668 | 502.5438
0.70241 | 1064963.9434 | 83.9258 | 202.6147 | 485.8124
0.70312 | 1083442.1412 | 86.8616 | 209.7023 | 472.9528
0.66836 | 991258.9202 | 86.5398 | 208.9256 | 507.2979
0.67779 | 1013279.4862 | 78.0207 | 188.3586 | 464.0133
0.68774 | 1015346.6741 | 64.8445 | 156.5484 | 487.2058
0.68655 | 1028308.2875 | 85.926 | 207.4438 | 500.6034
0.68967 | 1042313.2084 | 86.7019 | 209.3168 | 478.7072
0.70288 | 1062870.1434 | 77.2555 | 186.5112 | 491.7706
0.70438 | 1062985.4066 | 81.9745 | 197.9038 | 476.5041
0.709 | 1078853.0125 | 76.0314 | 183.556 | 456.4877
0.70856 | 1093185.9579 | 88.5097 | 213.6813 | 481.3622
0.70778 | 1067440.3164 | 62.2224 | 150.2182 | 483.3077
0.69853 | 1051596.1382 | 81.7755 | 197.4236 | 506.9893
0.6693 | 990836.9001 | 59.0063 | 142.4537 | 468.3179
0.70895 | 1094595.0117 | 71.6924 | 173.0808 | 491.8439
0.68923 | 1036756.9576 | 64.096 | 154.7414 | 473.3464
0.67075 | 984320.4363 | 71.6387 | 172.951 | 487.425
0.67656 | 1003731.2807 | 67.051 | 161.8755 | 519.2323
0.69211 | 1039065.2076 | 86.1277 | 207.9306 | 507.9633
0.70064 | 1063478.652 | 75.8425 | 183.1 | 497.4924
0.7158 | 1109946.4362 | 85.8345 | 207.2228 | 476.8137
0.69234 | 1032527.5679 | 67.4024 | 162.7237 | 467.2353
0.67657 | 1021130.933 | 57.9686 | 139.9486 | 505.8555
0.70132 | 1061801.3317 | 79.0129 | 190.7541 | 489.7669
0.67326 | 1003086.1773 | 76.1174 | 183.7636 | 491.4753
0.68936 | 1047183.4666 | 74.1566 | 179.0298 | 507.2502
0.68196 | 1034395.3152 | 80.7922 | 195.0495 | 493.4216
0.68878 | 1042004.8087 | 82.1757 | 198.3898 | 453.9803
0.68132 | 996625.4425 | 73.0897 | 176.454 | 510.2543
0.67547 | 1000157.5432 | 86.3973 | 208.5816 | 482.5932
0.67819 | 1027160.9637 | 79.0419 | 190.8241 | 516.3558
0.69532 | 1049917.9693 | 77.144 | 186.2422 | 506.8495
0.6946 | 1038550.4131 | 69.4275 | 167.6127 | 501.0148
0.65982 | 966066.3144 | 46.7899 | 112.9608 | 500.6806
0.66804 | 983518.7765 | 66.5387 | 160.6385 | 504.6404
0.68663 | 1041546.627 | 79.1643 | 191.1194 | 467.9053
0.70561 | 1079347.3615 | 78.1498 | 188.6702 | 478.2071
0.67778 | 1008557.9394 | 69.0953 | 166.8109 | 478.2796
0.69338 | 1060260.8121 | 76.4047 | 184.4573 | 485.5985
0.66621 | 996244.2957 | 65.6804 | 158.5666 | 518.7274
0.70089 | 1062709.6015 | 68.5271 | 165.439 | 490.0111
0.69433 | 1058822.4979 | 87.4841 | 211.2052 | 505.3289
0.70995 | 1080039.3784 | 72.4752 | 174.9705 | 486.0905
0.70247 | 1063182.4459 | 68.4031 | 165.1397 | 498.7938
0.67526 | 1011867.0489 | 73.1523 | 176.6052 | 500.7021
0.69706 | 1056819.686 | 72.2288 | 174.3758 | 520.7264
0.7275 | 1131826.5591 | 81.1933 | 196.018 | 490.9976
0.67399 | 996457.9164 | 84.8166 | 204.7654 | 505.6707
0.70124 | 1082907.4148 | 77.1613 | 186.2838 | 508.2884
0.70107 | 1084158.3853 | 77.2694 | 186.5448 | 484.5772
0.65085 | 943790.6762 | 67.3334 | 162.5571 | 507.737
0.67717 | 993645.7849 | 73.3182 | 177.0059 | 484.5847
0.71641 | 1099875.2278 | 71.2887 | 172.1061 | 482.0142
0.67835 | 1011955.9404 | 44.9468 | 108.5113 | 470.962
0.696 | 1043571.3707 | 81.4465 | 196.6293 | 471.2724
0.69334 | 1043152.0567 | 71.677 | 173.0436 | 495.2731
0.70564 | 1070470.1956 | 74.0572 | 178.79 | 500.0202
0.68908 | 1031816.9519 | 61.7776 | 149.1443 | 461.8658
0.69481 | 1036063.9543 | 79.4249 | 191.7488 | 471.092
0.68209 | 1023478.9487 | 65.9056 | 159.1102 | 474.2504
0.69253 | 1048872.4515 | 72.1541 | 174.1955 | 494.6632
0.68082 | 1004587.1446 | 80.152 | 193.5041 | 498.955
0.69248 | 1034690.1244 | 71.1838 | 171.853 | 492.6322
0.68925 | 1036613.8067 | 84.7801 | 204.6774 | 502.6458
0.68772 | 1043367.6248 | 66.6361 | 160.8737 | 485.829
0.67542 | 1014780.1223 | 76.7592 | 185.3131 | 471.5543
0.69332 | 1056108.3346 | 90.351 | 218.1267 | 518.3248
0.68374 | 1021826.9895 | 72.9562 | 176.1319 | 483.682
0.68848 | 1035903.6938 | 76.8702 | 185.5812 | 509.3911
0.66833 | 994625.8049 | 70.9 | 171.1677 | 506.2525
0.6956 | 1051105.4168 | 78.8608 | 190.3868 | 474.225
0.67282 | 1010670.1223 | 83.1499 | 200.7417 | 488.4533
0.68352 | 1031010.6074 | 77.0623 | 186.0449 | 475.0743
0.67226 | 995297.4085 | 83.4582 | 201.4858 | 508.6124
0.68722 | 1038651.6655 | 75.518 | 182.3166 | 508.2032
0.69098 | 1059847.6833 | 76.6403 | 185.026 | 499.7426
0.69572 | 1055927.6216 | 70.4756 | 170.1432 | 457.1864
0.66853 | 970963.5042 | 74.2014 | 179.1381 | 468.8585
0.68858 | 1031134.6103 | 81.3424 | 196.378 | 521.2946
0.69313 | 1037259.6182 | 76.4543 | 184.5769 | 488.0817
0.69488 | 1053452.4906 | 68.5178 | 165.4167 | 461.3873
0.68393 | 1034291.3617 | 73.8251 | 178.2296 | 501.2432
0.72335 | 1124550.7035 | 81.1341 | 195.875 | 503.962
0.69237 | 1046296.7286 | 85.1861 | 205.6573 | 491.5215
0.69162 | 1037093.5804 | 63.8946 | 154.2552 | 463.3827
0.71107 | 1096471.4425 | 65.7056 | 158.6274 | 488.1264
0.70262 | 1076766.3033 | 73.009 | 176.2593 | 498.5612
0.70099 | 1075465.3731 | 65.3739 | 157.8265 | 510.3012
0.69046 | 1042276.5463 | 89.5365 | 216.1602 | 476.2396
0.69434 | 1047942.4799 | 71.902 | 173.5869 | 478.2987
0.68969 | 1051551.5812 | 74.4306 | 179.6914 | 472.0414
0.71339 | 1093687.5179 | 69.2885 | 167.2773 | 505.3523
0.6729 | 1006027.5615 | 71.226 | 171.9547 | 473.9762
0.67139 | 994227.1913 | 87.8349 | 212.0521 | 487.8923
0.7116 | 1084547.5942 | 74.4017 | 179.6217 | 485.8106
0.69179 | 1052793.3508 | 53.767 | 129.805 | 498.536
0.69425 | 1048428.6768 | 68.2528 | 164.7768 | 489.3797
0.68415 | 1026605.7404 | 89.0048 | 214.8765 | 484.428
0.69877 | 1053276.1128 | 90.8374 | 219.3008 | 525.6003
0.67377 | 1010309.9126 | 65.8607 | 159.0017 | 479.1968
0.70583 | 1076941.0655 | 76.2371 | 184.0527 | 497.289
0.68365 | 1014427.2348 | 61.2196 | 147.7972 | 468.7733
0.68557 | 1026812.2098 | 74.8366 | 180.6715 | 504.2359
0.70186 | 1088509.9153 | 66.084 | 159.5408 | 488.1781
0.69305 | 1053697.3004 | 80.9821 | 195.508 | 502.0439
0.70115 | 1084419.0524 | 87.1847 | 210.4825 | 487.3549
0.66713 | 970805.4521 | 63.9148 | 154.3039 | 527.4407
0.68832 | 1049510.1956 | 72.6393 | 175.3668 | 500.7106
0.69516 | 1057909.4704 | 62.786 | 151.5789 | 483.9296
0.71591 | 1098692.8372 | 71.2615 | 172.0405 | 518.3536
0.68714 | 1028228.8781 | 59.6432 | 143.9914 | 500.128
0.69909 | 1049123.4729 | 67.8246 | 163.743 | 465.4848
0.68439 | 1014366.2365 | 89.3198 | 215.6371 | 496.4259
0.67042 | 986150.6537 | 58.6297 | 141.5445 | 523.3172
0.67672 | 1014432.9085 | 81.4279 | 196.5844 | 477.8224
0.70529 | 1073859.345 | 62.8173 | 151.6544 | 537.3531
0.69727 | 1055452.5343 | 73.8945 | 178.3971 | 491.0382
0.70658 | 1088159.1657 | 73.3686 | 177.1275 | 500.2302
0.69976 | 1064412.4306 | 85.7461 | 207.0094 | 475.6772
0.67426 | 1007119.6368 | 65.6047 | 158.3837 | 481.5469
0.69553 | 1065007.5836 | 67.2542 | 162.366 | 482.6688
0.68247 | 1015548.8554 | 77.6592 | 187.4858 | 498.5648
0.69414 | 1061067.8545 | 79.7761 | 192.5965 | 524.0868
0.68311 | 1009721.3468 | 75.516 | 182.3118 | 490.0653
0.70195 | 1077948.8428 | 58.9816 | 142.3942 | 492.9873
0.70404 | 1084375.6798 | 69.6112 | 168.0562 | 476.2009
0.66222 | 975652.3794 | 80.029 | 193.2072 | 493.5424
0.70226 | 1073594.1567 | 76.3874 | 184.4154 | 491.7037
0.67517 | 1004999.0979 | 74.375 | 179.5572 | 493.4552
0.68524 | 1021539.7126 | 67.3761 | 162.6604 | 490.9989
0.69736 | 1068327.5594 | 69.2382 | 167.1558 | 483.2545
0.70716 | 1058821.7992 | 85.8974 | 207.3747 | 503.3392
0.69315 | 1051913.0709 | 79.1999 | 191.2056 | 519.1302
0.67062 | 1008000.7088 | 64.0476 | 154.6247 | 479.6438
0.68672 | 1039045.7548 | 85.2915 | 205.9119 | 466.1591
0.68707 | 1037017.8538 | 75.4153 | 182.0687 | 474.6321
0.69127 | 1047880.2729 | 70.3881 | 169.9318 | 511.6499
0.69441 | 1071364.7921 | 73.5024 | 177.4504 | 485.7261
0.72229 | 1129025.8516 | 81.5766 | 196.9433 | 477.9064
0.68426 | 1030146.8723 | 69.9059 | 168.7677 | 478.3414
0.7113 | 1078111.6668 | 75.7332 | 182.8361 | 517.6466
0.67763 | 1022702.6213 | 77.1056 | 186.1493 | 486.8432
0.68443 | 1034691.9363 | 66.3992 | 160.3019 | 491.3466
0.69524 | 1057342.3611 | 87.2482 | 210.6358 | 483.0256
0.67369 | 997555.1163 | 68.7562 | 165.9921 | 478.4601
0.67787 | 995671.6066 | 69.9757 | 168.9363 | 489.8611
0.69681 | 1063421.083 | 77.9685 | 188.2325 | 482.193
0.66363 | 975877.1321 | 73.4756 | 177.3857 | 498.7319
0.69898 | 1082109.2337 | 73.7539 | 178.0576 | 484.9692
0.68964 | 1046455.3318 | 77.5222 | 187.1552 | 468.9931
0.69403 | 1052464.4068 | 83.9426 | 202.6553 | 525.052
0.69127 | 1066059.6047 | 82.1435 | 198.3119 | 521.5508
0.66848 | 1003812.0581 | 79.9835 | 193.0973 | 494.1816
0.70607 | 1063525.2767 | 68.0767 | 164.3517 | 502.1529
0.68679 | 1057122.4896 | 73.3296 | 177.0332 | 527.4364
0.69129 | 1054364.6512 | 74.4562 | 179.7532 | 464.9074
0.70659 | 1074630.1451 | 60.6193 | 146.348 | 471.4984
0.68496 | 1048526.889 | 79.8889 | 192.8688 | 470.3161
0.6802 | 997515.6799 | 68.0958 | 164.3979 | 519.4646
0.68368 | 1014408.0003 | 67.0588 | 161.8942 | 511.5274
0.69788 | 1065783.138 | 69.0212 | 166.6318 | 472.5663
0.66682 | 995820.0147 | 67.4945 | 162.9461 | 496.7835
0.68417 | 1029665.91 | 69.8829 | 168.7124 | 488.4117
0.67551 | 990455.0048 | 73.2624 | 176.871 | 481.0962
0.66668 | 976561.8387 | 73.1913 | 176.6994 | 474.9328
0.6875 | 1030001.3651 | 67.7112 | 163.4693 | 478.1147
0.70476 | 1082795.8426 | 67.482 | 162.9159 | 483.6601
0.68674 | 1031822.3468 | 80.0294 | 193.208 | 470.3497
0.6876 | 1049599.1053 | 62.9776 | 152.0414 | 464.8736
0.65644 | 969340.1306 | 72.0871 | 174.0337 | 480.1391
0.69046 | 1026549.0039 | 72.5795 | 175.2223 | 492.3369
0.67919 | 1020294.0279 | 75.0609 | 181.213 | 504.3773
0.68221 | 1010864.9258 | 82.0215 | 198.0175 | 485.7754
0.67368 | 1014882.6859 | 75.2724 | 181.7238 | 515.668
0.70165 | 1077088.1754 | 89.5706 | 216.2427 | 500.8579
0.70148 | 1068851.5195 | 82.768 | 199.8195 | 512.7415
0.70413 | 1073136.9559 | 66.94 | 161.6075 | 479.1987
0.69339 | 1053462.4234 | 68.6653 | 165.7727 | 494.0756
0.67408 | 1012651.8839 | 69.7568 | 168.4077 | 485.5304
0.69789 | 1064330.7025 | 72.6551 | 175.4049 | 504.4825
0.70557 | 1079178.858 | 74.6662 | 180.2602 | 477.5841
0.69174 | 1036494.1608 | 67.7206 | 163.4921 | 476.9507
0.68241 | 1020563.5704 | 65.6865 | 158.5812 | 490.8045
0.68418 | 1038050.8071 | 70.4509 | 170.0835 | 471.9571
0.69335 | 1057601.089 | 77.9296 | 188.1388 | 483.6021
0.66613 | 994411.4434 | 76.2669 | 184.1247 | 516.0145
0.69512 | 1044445.2256 | 89.1217 | 215.1589 | 504.7321
0.70426 | 1061339.1218 | 66.1916 | 159.8007 | 464.7678
0.67816 | 1018467.4067 | 71.5467 | 172.7289 | 498.4134
0.69197 | 1032747.8357 | 87.0855 | 210.2429 | 488.6287
0.69098 | 1056754.305 | 78.1338 | 188.6318 | 477.6611
0.67912 | 1021173.4295 | 75.1331 | 181.3873 | 497.2902
0.64419 | 931545.6786 | 70.593 | 170.4266 | 486.091
0.7238 | 1115654.7181 | 63.5935 | 153.5283 | 531.4291
0.68544 | 1029742.8691 | 83.1599 | 200.7659 | 461.7994
0.69259 | 1045970.7596 | 75.4609 | 182.1786 | 483.2938
0.69973 | 1069119.4022 | 76.3439 | 184.3105 | 490.1623
0.66584 | 996890.7049 | 73.0892 | 176.4529 | 498.3356
0.66894 | 980094.2503 | 69.7999 | 168.512 | 490.6182
0.69091 | 1040238.1803 | 68.0469 | 164.2797 | 479.1456
0.68176 | 1019618.4111 | 78.0517 | 188.4336 | 512.0097
0.68882 | 1051993.9629 | 74.5557 | 179.9934 | 488.5636
0.68803 | 1037210.3116 | 57.901 | 139.7855 | 482.5853
0.68212 | 1027672.7695 | 71.7747 | 173.2794 | 486.8346
0.67601 | 1020613.0651 | 85.1064 | 205.465 | 488.9507
0.70701 | 1087160.8625 | 71.4017 | 172.3789 | 496.895
0.69202 | 1043151.0743 | 71.8131 | 173.3723 | 512.6613
0.67567 | 1017377.1459 | 75.2355 | 181.6345 | 469.3808
0.69329 | 1065633.3694 | 81.8712 | 197.6547 | 455.1719
0.70833 | 1093224.1514 | 78.8499 | 190.3606 | 532.534
0.70413 | 1070282.861 | 82.5659 | 199.3317 | 501.6653
0.69002 | 1037218.5796 | 64.0505 | 154.6315 | 492.2328
0.72911 | 1147206.4884 | 88.6956 | 214.1302 | 477.5082
0.67938 | 1013005.0667 | 82.5373 | 199.2627 | 499.6694
0.69415 | 1047792.8281 | 73.6875 | 177.8974 | 523.4341
0.71936 | 1097076.9211 | 76.2017 | 183.9671 | 489.8126
0.67568 | 1015233.3191 | 65.9943 | 159.3244 | 482.4955
0.68041 | 1020605.1548 | 71.4603 | 172.5203 | 474.8185
0.67693 | 1015081.1692 | 65.8631 | 159.0076 | 494.0483
0.70017 | 1068749.293 | 56.3269 | 135.9852 | 511.9945
0.69469 | 1050014.6776 | 89.6524 | 216.4399 | 494.3539
0.6697 | 1005060.1543 | 67.8276 | 163.7503 | 489.444
0.72406 | 1125046.2444 | 77.6375 | 187.4336 | 486.9039
0.69078 | 1054135.8179 | 68.9683 | 166.5042 | 501.4048
0.71899 | 1084370.3695 | 70.0315 | 169.0709 | 465.215
0.67572 | 1014390.0137 | 71.5176 | 172.6588 | 498.8562
0.71548 | 1104322.6361 | 71.6696 | 173.0257 | 494.8551
0.67164 | 985545.2524 | 73.3194 | 177.0087 | 485.907
0.68692 | 1038091.6815 | 86.6974 | 209.306 | 503.8469
0.71161 | 1089993.3108 | 72.9728 | 176.1719 | 529.1744
0.71425 | 1100759.7487 | 89.9248 | 217.0976 | 492.8105
0.70029 | 1068920.5148 | 77.9824 | 188.2662 | 501.1187
0.6771 | 1021823.0581 | 75.4046 | 182.0427 | 496.0298
0.65737 | 962338.9422 | 66.1628 | 159.731 | 469.308
0.69524 | 1050502.6898 | 60.2255 | 145.3972 | 492.7424
0.70208 | 1075822.2844 | 83.2366 | 200.9508 | 510.0197
0.69523 | 1059132.5242 | 76.1202 | 183.7704 | 474.8908
0.70424 | 1072268.9612 | 72.5376 | 175.1213 | 478.6035
0.68408 | 1025000.6702 | 79.8099 | 192.678 | 501.0839
0.71714 | 1109693.5184 | 71.8515 | 173.4649 | 478.5038
0.69768 | 1071859.5568 | 81.8318 | 197.5594 | 480.909
0.68262 | 1001293.5852 | 68.1116 | 164.436 | 476.0952
0.67757 | 1013629.1968 | 84.3294 | 203.5892 | 482.2178
0.70826 | 1095924.5866 | 69.9336 | 168.8346 | 466.2759
0.70457 | 1089137.4007 | 65.8682 | 159.0198 | 514.9837
0.70316 | 1059244.2537 | 76.9617 | 185.802 | 476.8582
0.66492 | 965057.5791 | 61.8448 | 149.3065 | 476.7272
0.70284 | 1067323.0537 | 87.1195 | 210.325 | 498.9318
0.69417 | 1058204.9288 | 77.882 | 188.0238 | 499.2884
0.68613 | 1040862.5104 | 68.5618 | 165.5228 | 479.9399
0.68924 | 1023201.9211 | 65.2038 | 157.4158 | 484.6606
0.68146 | 1021854.8663 | 84.5282 | 204.069 | 475.938
0.69266 | 1067822.3722 | 77.0288 | 185.9641 | 513.6444
0.6787 | 1021615.3012 | 83.323 | 201.1596 | 495.2208
0.67739 | 1005684.5134 | 61.2521 | 147.8757 | 485.2457
0.68008 | 1019193.4743 | 84.479 | 203.9503 | 474.4596
0.70104 | 1056839.3036 | 73.5932 | 177.6698 | 501.2757
0.68776 | 1040978.2397 | 77.3272 | 186.6844 | 515.2766
0.69003 | 1039299.4179 | 69.9375 | 168.8442 | 491.8017
0.69017 | 1038644.8867 | 59.9832 | 144.8124 | 500.8049
0.66771 | 988166.6159 | 84.9021 | 204.9717 | 479.4347
0.70171 | 1052822.1099 | 93.2003 | 225.0055 | 522.1794
0.69413 | 1045776.6444 | 69.9576 | 168.8926 | 500.5351
0.6957 | 1064488.3563 | 87.7789 | 211.9169 | 495.9987
0.69336 | 1033281.8491 | 65.4551 | 158.0227 | 483.6187
0.68555 | 1014966.4609 | 63.0855 | 152.3018 | 481.1334
0.68144 | 1015173.7215 | 61.4627 | 148.384 | 470.0416
0.6886 | 1037231.9786 | 71.7604 | 173.2448 | 496.3825
0.69123 | 1046063.1403 | 60.319 | 145.6229 | 525.3281
0.68339 | 1042228.9232 | 72.6561 | 175.4073 | 508.7824
0.68594 | 1033150.3823 | 64.1796 | 154.9433 | 469.3716
0.69865 | 1066447.7127 | 79.6629 | 192.3233 | 479.4108
0.70091 | 1079802.6611 | 82.0583 | 198.1063 | 503.6309
0.71324 | 1091326.7321 | 78.1722 | 188.7244 | 479.2658
0.71167 | 1100049.5745 | 67.892 | 163.9057 | 480.3908
0.69376 | 1057040.7575 | 79.5168 | 191.9706 | 477.0613
0.70199 | 1076429.8621 | 69.3027 | 167.3116 | 473.5141
0.68111 | 1020770.1365 | 78.2143 | 188.8259 | 498.5632
0.68893 | 1039600.8164 | 75.3994 | 182.0302 | 492.6329
0.70566 | 1069867.5251 | 70.3998 | 169.96 | 507.4312
0.68321 | 1020554.5322 | 61.6284 | 148.784 | 504.1679
0.67915 | 1028485.6127 | 80.6629 | 194.7375 | 501.1929
0.69786 | 1054899.107 | 85.2108 | 205.717 | 475.5776
0.69418 | 1051839.5288 | 77.4119 | 186.8888 | 467.494
0.6813 | 1018994.0598 | 87.2712 | 210.6913 | 492.8772
0.66018 | 971092.4167 | 82.9178 | 200.1812 | 512.683
0.67903 | 1013276.799 | 65.5426 | 158.2339 | 481.7871
0.67632 | 1005020.8914 | 77.5526 | 187.2285 | 486.3803
0.69624 | 1065475.1132 | 71.0598 | 171.5536 | 486.9748
0.68986 | 1042963.5704 | 73.8045 | 178.1797 | 528.909
0.67578 | 1006881.6485 | 63.4873 | 153.2719 | 475.1267
0.69334 | 1044197.117 | 87.8765 | 212.1526 | 494.8503
0.67601 | 1003002.3827 | 77.432 | 186.9375 | 498.7308
0.68966 | 1042852.8255 | 68.3419 | 164.9918 | 475.9702
0.68299 | 1033044.2867 | 70.2335 | 169.5586 | 467.5194
0.68687 | 1046295.7324 | 72.5702 | 175.1999 | 496.2664
0.69219 | 1050023.5659 | 79.1681 | 191.1287 | 499.3207
0.68645 | 1052157.3552 | 86.4018 | 208.5924 | 466.3235
0.6954 | 1052001.6039 | 70.2461 | 169.589 | 483.301
0.68889 | 1035648.2639 | 73.3897 | 177.1783 | 501.0306
0.68195 | 1032487.5925 | 75.6498 | 182.6348 | 472.0248
0.68718 | 1035657.3663 | 71.8001 | 173.3408 | 524.8921
0.69438 | 1047084.4895 | 73.6138 | 177.7194 | 475.4395
0.68838 | 1036978.8671 | 62.1504 | 150.0444 | 508.3804
0.67618 | 997736.9524 | 64.809 | 156.4628 | 492.1527
0.67012 | 990310.2307 | 70.4033 | 169.9686 | 496.3763
0.69066 | 1032453.3096 | 82.7489 | 199.7735 | 506.8024
0.68461 | 1018458.5342 | 69.0778 | 166.7687 | 472.9322
0.68378 | 1028042.7058 | 67.9498 | 164.0454 | 487.5305
0.70016 | 1046483.4626 | 78.0023 | 188.3142 | 489.6446
0.66305 | 964925.7726 | 83.4873 | 201.5562 | 450.032
0.68798 | 1028992.804 | 71.2543 | 172.0231 | 527.1561
0.6763 | 1003371.1198 | 52.4162 | 126.544 | 495.6773
0.69864 | 1068531.6107 | 78.1386 | 188.6433 | 501.0359
0.70893 | 1084701.2679 | 72.8924 | 175.9778 | 513.2362
0.67602 | 1020215.479 | 77.1017 | 186.1399 | 479.0323
0.70507 | 1072947.0364 | 68.7767 | 166.0416 | 485.0631
0.69043 | 1023070.12 | 47.3539 | 114.3224 | 479.9765
0.67377 | 991471.7996 | 68.329 | 164.9608 | 498.5359
0.69515 | 1056422.4193 | 77.6412 | 187.4425 | 485.2505
0.69173 | 1036093.882 | 71.3571 | 172.2712 | 484.8662
0.67312 | 975644.3803 | 63.6619 | 153.6934 | 508.1022
0.70743 | 1078654.8037 | 72.6405 | 175.3697 | 536.7244
0.69966 | 1077305.8789 | 80.357 | 193.999 | 458.2013
0.7088 | 1104201.4655 | 66.3804 | 160.2566 | 489.5876
0.69749 | 1046616.1683 | 76.662 | 185.0785 | 481.9282
0.72035 | 1115907.1194 | 79.143 | 191.0681 | 477.8768
0.69895 | 1065800.1912 | 70.8057 | 170.9401 | 494.8735
0.69337 | 1058689.3478 | 89.0925 | 215.0882 | 501.3681
0.69293 | 1044244.7367 | 83.4192 | 201.3917 | 509.6855
0.67566 | 1007393.7391 | 69.2326 | 167.1422 | 495.9632
0.67585 | 1016432.7962 | 57.7636 | 139.4537 | 466.6665
0.69375 | 1050759.9205 | 70.3292 | 169.7897 | 474.3117
0.68485 | 1005177.028 | 64.4661 | 155.6349 | 472.0373
0.6918 | 1043323.6787 | 73.6226 | 177.7408 | 453.7401
0.68677 | 1025314.0675 | 84.1581 | 203.1757 | 459.4189
0.67664 | 1000893.4005 | 80.3504 | 193.9831 | 495.2227
0.69807 | 1054799.0267 | 70.1631 | 169.3886 | 537.1238
0.67851 | 993097.0659 | 69.7293 | 168.3414 | 472.3787
0.66126 | 983194.1694 | 57.1072 | 137.8689 | 472.1714
0.69478 | 1062951.5299 | 74.6866 | 180.3094 | 498.35
0.68332 | 1039988.6188 | 60.6624 | 146.4519 | 492.8482
0.67059 | 1013019.8792 | 85.1557 | 205.5841 | 497.5395
0.71136 | 1096961.0797 | 83.7642 | 202.2247 | 502.8734
0.71576 | 1101037.1127 | 67.5642 | 163.1144 | 483.9838
0.65777 | 965632.3047 | 75.0007 | 181.0677 | 498.3047
0.67472 | 1007715.9839 | 74.6263 | 180.1638 | 467.691
0.68493 | 1023380.6792 | 79.5843 | 192.1334 | 486.5506
0.70128 | 1063991.4931 | 74.0663 | 178.8119 | 500.4752
0.7292 | 1140154.2567 | 68.675 | 165.7961 | 478.0892
0.68813 | 1030410.7029 | 74.8194 | 180.63 | 475.4502
0.6901 | 1046257.2561 | 62.5501 | 151.0092 | 499.6125
0.6808 | 1006874.7067 | 73.8976 | 178.4046 | 500.9009
0.6988 | 1070484.1519 | 67.2842 | 162.4384 | 499.9277
0.67785 | 1023288.0105 | 77.1246 | 186.1952 | 524.0183
0.66522 | 1001632.7196 | 80.2012 | 193.6227 | 459.2742
0.68941 | 1054691.9521 | 63.1174 | 152.3789 | 487.1158
0.69591 | 1079328.8796 | 73.8885 | 178.3825 | 504.1961
0.68192 | 1025903.9492 | 78.514 | 189.5495 | 505.7719
0.67979 | 1002934.504 | 64.8007 | 156.4428 | 480.2693
0.70459 | 1078927.1637 | 76.5123 | 184.7171 | 483.2316
0.67916 | 1033560.3932 | 77.7488 | 187.7023 | 472.7647
0.69874 | 1078133.3366 | 71.646 | 172.9687 | 518.6627
0.69425 | 1048947.3473 | 80.4776 | 194.29 | 492.2456
0.6929 | 1045298.3037 | 73.6452 | 177.7954 | 465.2213
0.68408 | 1022458.5584 | 74.6818 | 180.2979 | 482.8193
0.7105 | 1090153.4942 | 65.3789 | 157.8387 | 510.6837
0.6873 | 1020966.8388 | 54.8157 | 132.3368 | 482.3322
0.6784 | 1023124.5882 | 60.1116 | 145.1223 | 501.4121
0.66846 | 990638.735 | 82.0337 | 198.0468 | 487.3713
0.71354 | 1107179.7316 | 85.8924 | 207.3625 | 474.6755
0.67829 | 1005322.0322 | 78.9464 | 190.5934 | 481.0088
0.68941 | 1044112.5515 | 66.5205 | 160.5946 | 487.1105
0.66592 | 979029.33 | 71.7707 | 173.2698 | 487.659
0.67654 | 1010393.5354 | 79.5496 | 192.0496 | 476.2476
0.68326 | 1036747.828 | 77.5206 | 187.1512 | 497.3486
0.7008 | 1070434.3871 | 65.0852 | 157.1295 | 470.4468
0.6936 | 1035554.3615 | 82.6451 | 199.5229 | 502.552
0.67048 | 1000115.6935 | 63.5597 | 153.4467 | 500.5291
0.69001 | 1038482.1816 | 66.1042 | 159.5897 | 467.2442
0.68776 | 1042126.2826 | 83.8757 | 202.4938 | 490.6358
0.69088 | 1036787.2205 | 73.259 | 176.8628 | 500.2256
0.69704 | 1058549.2974 | 70.5165 | 170.2419 | 521.8785
0.70194 | 1074353.2682 | 74.7167 | 180.3821 | 509.9262
0.70364 | 1065957.3081 | 89.8567 | 216.9332 | 462.1927
0.64148 | 938849.4173 | 62.3636 | 150.559 | 509.2234
0.68536 | 1037465.9857 | 78.9999 | 190.7226 | 490.1265
0.67584 | 1005032.097 | 69.0517 | 166.7056 | 501.8649
0.68654 | 1029824.4859 | 63.9168 | 154.3088 | 514.3918
0.66897 | 1000417.2483 | 71.2449 | 172.0004 | 498.0262
0.66755 | 987737.7301 | 72.2325 | 174.3846 | 510.1853
0.66033 | 967004.3029 | 64.321 | 155.2846 | 474.3236
0.66299 | 970024.1701 | 69.7037 | 168.2795 | 460.3248
0.7023 | 1092136.8037 | 77.7022 | 187.5896 | 498.774
0.67915 | 1018774.9911 | 74.9532 | 180.953 | 473.7859
0.69719 | 1050844.5508 | 78.6612 | 189.905 | 485.0461
0.69991 | 1067276.299 | 81.2219 | 196.0871 | 499.8305
0.69733 | 1040645.4401 | 71.0959 | 171.6407 | 489.5481
0.66763 | 981718.9484 | 70.0537 | 169.1246 | 496.9983
0.67459 | 996171.0366 | 72.0882 | 174.0364 | 501.044
0.68487 | 1037393.7574 | 75.6849 | 182.7195 | 481.9408
0.67893 | 1010882.6453 | 61.4296 | 148.3042 | 506.5696
0.6808 | 1024357.691 | 69.8671 | 168.6741 | 471.9748
0.69495 | 1042673.9131 | 65.8984 | 159.0929 | 482.7076
0.6996 | 1061428.6435 | 72.2798 | 174.4988 | 523.6678
0.68493 | 1043478.9323 | 88.8938 | 214.6087 | 513.9581
0.67944 | 1007047.4701 | 74.059 | 178.7944 | 493.0366
0.66478 | 977832.2622 | 70.0754 | 169.177 | 514.6536
0.68302 | 1014739.2989 | 74.233 | 179.2144 | 494.8154
0.69316 | 1058961.4663 | 78.0506 | 188.4309 | 500.5281
0.70197 | 1071684.2013 | 72.7316 | 175.5895 | 455.4674
0.68605 | 1021424.9303 | 70.9301 | 171.2405 | 489.4441
0.71239 | 1106873.8971 | 79.3794 | 191.6389 | 496.8599
0.64967 | 945884.1728 | 58.031 | 140.0992 | 503.398
0.68317 | 1015611.092 | 72.5729 | 175.2065 | 493.9464
0.67055 | 989249.2167 | 72.0196 | 173.8707 | 478.3441
0.67202 | 1001890.8298 | 77.8058 | 187.8397 | 486.0488
0.67934 | 1019943.2762 | 89.053 | 214.993 | 500.7963
0.68116 | 1021581.0275 | 79.8842 | 192.8576 | 483.2316
0.67834 | 1005178.7008 | 72.3387 | 174.641 | 482.4896
0.68897 | 1052495.6195 | 67.0973 | 161.9871 | 471.4816
0.68825 | 1033474.7028 | 75.0296 | 181.1375 | 486.2596
0.67315 | 1006808.6111 | 79.6667 | 192.3325 | 503.747
0.70605 | 1088760.2142 | 86.9266 | 209.8593 | 496.8417
0.68584 | 1037343.2914 | 80.5633 | 194.4971 | 516.0952
0.6883 | 1041139.379 | 66.4379 | 160.3954 | 507.3606
0.71841 | 1103985.4266 | 78.5094 | 189.5384 | 519.9847
0.67228 | 1021288.5477 | 82.5189 | 199.2182 | 506.4719
0.68657 | 1025819.6651 | 76.3343 | 184.2873 | 528.7011
0.69036 | 1033342.5444 | 76.4715 | 184.6185 | 481.5887
0.67238 | 997627.3824 | 72.4319 | 174.866 | 517.0309
0.67561 | 1003750.6377 | 74.2344 | 179.2177 | 448.9717
0.67844 | 997164.7017 | 66.428 | 160.3714 | 501.4617
0.68789 | 1040849.0092 | 62.4263 | 150.7103 | 512.3767
0.69249 | 1050686.7868 | 80.8161 | 195.1074 | 474.5285
0.68206 | 1031058.8871 | 85.214 | 205.7248 | 474.9952
0.68003 | 1019791.9495 | 81.4591 | 196.6598 | 504.838
0.68131 | 1006672.9422 | 68.6094 | 165.6377 | 497.9173
0.70429 | 1063380.502 | 62.2928 | 150.3882 | 475.1211
0.68098 | 1022348.3188 | 67.8517 | 163.8086 | 485.3388
0.66971 | 993830.6871 | 80.6093 | 194.6081 | 483.8639
0.69825 | 1063017.7901 | 77.7791 | 187.7753 | 511.4121
0.71764 | 1106953.7872 | 90.0042 | 217.2893 | 478.9583
0.72236 | 1126683.1495 | 72.4249 | 174.8491 | 484.5006
0.69403 | 1049809.3452 | 73.8422 | 178.2709 | 501.6789
0.7048 | 1071684.9418 | 80.9923 | 195.5326 | 455.3093
0.68412 | 1012812.7832 | 66.9102 | 161.5356 | 514.2244
0.68497 | 1020759.7856 | 86.9535 | 209.9243 | 477.5609
0.71008 | 1104148.0528 | 77.6954 | 187.5732 | 454.0701
0.6729 | 993538.6741 | 52.165 | 125.9375 | 502.4268
0.69817 | 1059077.0253 | 58.2625 | 140.6581 | 479.1784
0.70296 | 1067405.6111 | 82.4811 | 199.127 | 510.3641
0.67775 | 1029918.0806 | 77.636 | 187.4299 | 489.4778
0.69184 | 1055455.5518 | 79.6179 | 192.2147 | 472.2988
0.6895 | 1023858.7254 | 79.0938 | 190.9492 | 516.1816
0.69051 | 1040804.3906 | 71.8088 | 173.3618 | 487.9943
0.66551 | 993987.8666 | 79.5552 | 192.0633 | 510.3864
0.68113 | 1030288.8032 | 85.2668 | 205.8522 | 506.4685
0.68865 | 1048332.8132 | 66.8623 | 161.4199 | 481.4832
0.68528 | 1039616.0184 | 79.1731 | 191.1407 | 469.6222
0.67124 | 1000956.27 | 69.2432 | 167.1679 | 493.1148
0.70138 | 1065506.3732 | 70.5408 | 170.3005 | 497.1199
0.67686 | 1016562.8417 | 71.551 | 172.7394 | 497.5858
0.68841 | 1034013.8989 | 81.5899 | 196.9755 | 474.082
0.68786 | 1032952.5071 | 55.4147 | 133.7828 | 468.7603
0.68807 | 1054972.455 | 73.0447 | 176.3456 | 490.6346
0.69498 | 1058415.6 | 77.8058 | 187.8399 | 466.5012
0.69105 | 1031892.4184 | 61.6064 | 148.7311 | 489.3067
0.70782 | 1085184.1503 | 81.662 | 197.1495 | 508.4394
0.71225 | 1115931.7098 | 88.3199 | 213.223 | 475.8473
0.69607 | 1051918.8414 | 62.3493 | 150.5245 | 498.8473
0.67326 | 998064.3032 | 63.1676 | 152.5001 | 468.9512
0.6823 | 1016375.5256 | 86.9195 | 209.8423 | 494.1582
0.6864 | 1041208.1135 | 66.6255 | 160.8481 | 507.2311
0.70039 | 1072309.0469 | 63.1661 | 152.4964 | 483.6071
0.72062 | 1131908.9742 | 90.1109 | 217.5469 | 495.3081
0.6978 | 1059759.8621 | 84.0674 | 202.9567 | 505.6411
0.69386 | 1041276.7656 | 76.67 | 185.0977 | 477.5507
0.7031 | 1079159.3402 | 74.9275 | 180.891 | 438.8357
0.68431 | 1018426.8266 | 83.8937 | 202.5372 | 472.1573
0.68068 | 1022968.1997 | 76.7798 | 185.3628 | 501.1841
0.69871 | 1050318.3198 | 76.7985 | 185.408 | 496.1122
0.68925 | 1042576.9157 | 82.1657 | 198.3656 | 525.0391
0.69223 | 1040882.4561 | 76.8209 | 185.4621 | 469.0582
0.67975 | 1027297.6011 | 73.777 | 178.1135 | 506.1927
0.70144 | 1066855.5278 | 72.4591 | 174.9318 | 481.1386
0.6517 | 949326.4216 | 71.1114 | 171.678 | 508.9945
0.70904 | 1084342.7389 | 74.6563 | 180.2362 | 499.8913
0.69932 | 1068147.7655 | 70.9242 | 171.2262 | 500.686
0.68724 | 1028017.7638 | 71.2643 | 172.0472 | 501.9762
0.66407 | 977573.9419 | 78.4109 | 189.3006 | 482.9732
0.6996 | 1053246.3533 | 71.7224 | 173.1531 | 503.3252
0.70592 | 1068847.832 | 58.1543 | 140.397 | 502.2653
0.69491 | 1044754.9067 | 67.949 | 164.0434 | 481.8355
0.6805 | 1011812.2241 | 67.2001 | 162.2355 | 521.4113
0.67676 | 1012563.9008 | 79.6948 | 192.4002 | 524.237
0.67831 | 1020227.0737 | 71.6548 | 172.99 | 487.9118
0.66475 | 970857.6376 | 66.0077 | 159.3567 | 478.7495
0.67693 | 1012135.2027 | 79.6402 | 192.2683 | 505.56
0.67475 | 1004059.8503 | 54.911 | 132.5669 | 508.5782
0.67706 | 996833.0936 | 64.6664 | 156.1186 | 486.323
0.70555 | 1074571.2826 | 81.1237 | 195.85 | 491.4196
0.69066 | 1043068.2955 | 73.7261 | 177.9906 | 495.9819
0.70299 | 1063964.2624 | 78.9867 | 190.6908 | 467.038
0.70383 | 1059072.6651 | 81.7361 | 197.3283 | 482.3272
0.69395 | 1050311.2434 | 85.3187 | 205.9776 | 518.4628
0.68315 | 1025433.4162 | 59.8287 | 144.4393 | 529.812
0.66295 | 981389.7117 | 78.8494 | 190.3594 | 471.6759
0.67852 | 1015758.4312 | 76.4822 | 184.6443 | 505.3826
0.68292 | 1021203.9251 | 65.6792 | 158.5637 | 503.5553
0.683 | 1041739.5324 | 85.9787 | 207.5709 | 468.3364
0.69447 | 1039391.6447 | 71.3018 | 172.1377 | 492.4708
0.67767 | 1018414.7339 | 83.1884 | 200.8347 | 476.7036
0.69286 | 1030192.1847 | 74.2932 | 179.3597 | 509.7916
0.67138 | 996837.5545 | 72.482 | 174.987 | 490.1586
0.69075 | 1054304.4612 | 65.0372 | 157.0138 | 506.8522
0.71829 | 1113939.7808 | 85.8455 | 207.2494 | 476.6325
0.69312 | 1063301.3666 | 74.8891 | 180.7982 | 514.681
0.68878 | 1029686.8937 | 50.3587 | 121.5766 | 476.6236
0.6808 | 995252.5671 | 53.9456 | 130.2362 | 503.1398
0.6862 | 1019555.092 | 77.5411 | 187.2009 | 469.5746
0.69692 | 1056014.9673 | 64.5663 | 155.8768 | 491.2709
0.67111 | 1014427.4951 | 75.527 | 182.3383 | 504.5466
0.71376 | 1105700.6767 | 79.1436 | 191.0695 | 491.1554
0.6612 | 964704.8858 | 65.9417 | 159.1974 | 475.7045
0.698 | 1063400.9805 | 76.1231 | 183.7773 | 507.0484
0.67645 | 1013362.9742 | 67.7052 | 163.4548 | 517.2971
0.69704 | 1050891.1524 | 86.4911 | 208.8079 | 493.3803
0.67372 | 993403.0637 | 80.0944 | 193.3649 | 491.6152
0.72129 | 1105640.4794 | 69.2791 | 167.2546 | 511.2334
0.69362 | 1064567.0296 | 74.9247 | 180.8841 | 470.1131
0.70219 | 1086791.3235 | 75.5554 | 182.4068 | 472.3359
0.67112 | 996148.7745 | 85.153 | 205.5776 | 498.6906
0.72979 | 1130171.2332 | 73.7774 | 178.1144 | 487.4522
0.69375 | 1067245.9304 | 76.8661 | 185.5713 | 502.1537
0.69994 | 1058432.5492 | 85.2426 | 205.7938 | 495.7023
0.69107 | 1030203.9617 | 63.5212 | 153.3538 | 484.139
0.69233 | 1039160.2126 | 73.2724 | 176.8953 | 514.5908
0.65565 | 948287.6112 | 59.672 | 144.061 | 479.7591
0.66639 | 972236.4344 | 72.8988 | 175.9933 | 491.7427
0.68469 | 1032441.6215 | 70.0178 | 169.0379 | 477.0689
0.68754 | 1035587.0381 | 87.8867 | 212.1773 | 499.2046
0.67667 | 1011656.8523 | 71.2254 | 171.9532 | 468.5015
0.69407 | 1044672.541 | 64.1352 | 154.8361 | 488.3233
0.69257 | 1045690.5579 | 80.4232 | 194.1588 | 487.2222
0.6923 | 1047469.0029 | 80.591 | 194.5639 | 510.7594
0.71269 | 1088086.6787 | 78.0003 | 188.3093 | 495.5911
0.68446 | 1006076.724 | 76.7112 | 185.1971 | 511.4434
0.67476 | 997916.5262 | 66.8409 | 161.3682 | 492.0283
0.6635 | 986487.0909 | 75.9592 | 183.3817 | 514.9044
0.69497 | 1051273.6813 | 58.1001 | 140.2659 | 502.0053
0.6644 | 989132.5471 | 65.3895 | 157.8643 | 493.746
0.67315 | 1009672.2502 | 84.551 | 204.1242 | 484.0699
0.70452 | 1079627.2996 | 79.9753 | 193.0776 | 457.4699
0.70377 | 1064731.8767 | 68.3764 | 165.0753 | 530.7945
0.68862 | 1033565.8197 | 63.9553 | 154.4018 | 464.2255
0.69348 | 1044884.7356 | 86.4388 | 208.6816 | 493.003
0.70263 | 1089676.7725 | 66.0827 | 159.5376 | 469.0261
0.69161 | 1031247.2796 | 67.3898 | 162.6934 | 502.0444
0.68297 | 1037841.5013 | 66.7437 | 161.1336 | 498.9448
0.67791 | 1006045.2827 | 52.4313 | 126.5804 | 482.5052
0.69941 | 1055309.6659 | 70.0629 | 169.1469 | 497.5881
0.70208 | 1064903.6424 | 88.5041 | 213.6679 | 488.5471
0.71185 | 1091804.3296 | 71.2191 | 171.9381 | 494.8443
0.68611 | 1027324.4769 | 73.0948 | 176.4664 | 460.9798
0.6822 | 1036862.3871 | 79.3196 | 191.4943 | 512.2956
0.66476 | 977886.5377 | 71.8121 | 173.3697 | 499.794
0.67802 | 1009575.8734 | 77.7088 | 187.6058 | 480.6557
0.69976 | 1054654.7316 | 76.9998 | 185.8939 | 505.6302
0.68551 | 1040062.163 | 75.8093 | 183.0199 | 476.7157
0.70861 | 1085126.1862 | 78.236 | 188.8783 | 515.5984
0.69375 | 1052766.3453 | 80.8293 | 195.1392 | 466.866
0.68459 | 1032535.2597 | 77.0559 | 186.0293 | 492.4112
0.6661 | 973957.0907 | 77.882 | 188.0239 | 483.7105
0.70209 | 1073869.352 | 65.1743 | 157.3447 | 488.1765
0.67374 | 987186.8726 | 74.8576 | 180.7223 | 461.9771
0.68569 | 1028295.1742 | 78.2639 | 188.9457 | 481.323
0.67619 | 1012423.0272 | 67.6692 | 163.3679 | 492.0309
0.6877 | 1039875.362 | 82.1366 | 198.2954 | 512.1823
0.69673 | 1065704.6171 | 71.0246 | 171.4685 | 482.1916
0.67038 | 993178.0853 | 66.9584 | 161.6519 | 488.5071
0.69588 | 1054328.2022 | 82.3 | 198.6899 | 480.8303
0.67892 | 1012961.7523 | 67.4328 | 162.7972 | 501.0853
0.70398 | 1085688.5684 | 73.4723 | 177.3777 | 481.3876
0.7024 | 1076410.2852 | 71.1593 | 171.7938 | 503.1687
0.716 | 1108016.0253 | 85.6618 | 206.806 | 500.9707
0.69624 | 1037715.6426 | 71.0058 | 171.4231 | 480.4529
0.68819 | 1025982.7922 | 76.1666 | 183.8824 | 493.6219
0.68972 | 1036325.5549 | 80.0323 | 193.215 | 496.0371
0.71377 | 1090648.046 | 76.3492 | 184.3234 | 522.9339
0.70056 | 1077259.907 | 88.2158 | 212.9718 | 493.7105
0.67215 | 1016943.1917 | 80.5882 | 194.5572 | 492.4224
0.69848 | 1056987.0751 | 72.0934 | 174.0489 | 507.566
0.68886 | 1034723.1844 | 86.1724 | 208.0386 | 487.2697
0.69778 | 1061055.8694 | 74.2327 | 179.2135 | 513.7135
0.68585 | 1010764.573 | 54.3829 | 131.292 | 492.8854
0.70177 | 1056909.9569 | 68.0669 | 164.3281 | 529.8895
0.69068 | 1046835.3002 | 75.7315 | 182.832 | 487.625
0.67551 | 1000914.6225 | 63.1368 | 152.4257 | 514.0109
0.69298 | 1032360.9134 | 74.4362 | 179.705 | 503.2954
0.70128 | 1071419.4425 | 70.9424 | 171.2702 | 500.5002
0.70544 | 1092284.6053 | 75.5877 | 182.4849 | 509.7763
0.68499 | 1019914.8569 | 47.5829 | 114.8752 | 491.7416
0.71284 | 1091578.7513 | 83.1799 | 200.814 | 505.6768
0.70591 | 1069087.5112 | 70.3241 | 169.7774 | 505.927
0.71034 | 1090119.603 | 65.3917 | 157.8694 | 486.2455
0.68156 | 1013259.6578 | 70.8262 | 170.9896 | 477.3734
0.68447 | 1024060.3186 | 73.0372 | 176.3274 | 498.2282
0.68343 | 1017167.4885 | 74.3664 | 179.5364 | 493.0085
0.6855 | 1030849.6814 | 73.4113 | 177.2306 | 503.9402
0.66945 | 995103.1109 | 59.2468 | 143.0344 | 518.4821
0.69029 | 1027709.0434 | 80.8715 | 195.2411 | 496.0285
0.69939 | 1056799.3092 | 71.9861 | 173.7897 | 490.2418
0.70128 | 1069539.9414 | 62.9224 | 151.908 | 462.9863
0.68561 | 1032647.0787 | 78.6047 | 189.7686 | 513.4876
0.67103 | 998301.0057 | 76.0065 | 183.4959 | 488.1608
0.70653 | 1084727.1735 | 78.4369 | 189.3635 | 483.495
0.67814 | 1023241.448 | 71.8843 | 173.544 | 497.2838
0.69231 | 1045370.4321 | 71.9362 | 173.6693 | 486.9508
0.67435 | 1003699.1809 | 72.6594 | 175.4154 | 526.0272
0.71771 | 1092925.3674 | 60.6962 | 146.5336 | 490.5554
0.66915 | 973788.0558 | 65.7712 | 158.7857 | 478.7872
0.68961 | 1048461.8404 | 91.6233 | 221.1983 | 523.4364
0.71799 | 1101020.166 | 78.3444 | 189.1401 | 505.9577
0.70222 | 1089671.5058 | 73.9262 | 178.4736 | 500.9462
0.6641 | 970319.8947 | 64.1744 | 154.9308 | 497.5886
0.68085 | 1021022.2827 | 78.2637 | 188.9452 | 482.5821
0.69828 | 1046028.9944 | 67.7032 | 163.45 | 474.7443
0.66313 | 958383.8017 | 61.5822 | 148.6727 | 459.435
0.6886 | 1031300.9789 | 63.8169 | 154.0675 | 510.4917
0.67349 | 1007737.7347 | 78.0892 | 188.5241 | 474.1308
0.70582 | 1073069.965 | 77.6723 | 187.5175 | 493.4771
0.71118 | 1096527.9177 | 61.4552 | 148.3661 | 484.604
0.68111 | 1014657.4375 | 63.4775 | 153.2482 | 514.1911
0.68782 | 1053841.0845 | 76.5556 | 184.8217 | 463.4734
0.67309 | 1011416.2956 | 75.0811 | 181.2618 | 508.2867
0.72018 | 1123310.431 | 79.6872 | 192.3819 | 481.7894
0.68974 | 1028417.5653 | 65.4883 | 158.1028 | 503.5769
0.70511 | 1076927.0952 | 79.7084 | 192.4332 | 486.8327
0.71105 | 1086234.0881 | 86.1141 | 207.8979 | 482.2884
0.6946 | 1058091.363 | 76.1138 | 183.7549 | 481.266
0.70096 | 1064191.2101 | 66.6818 | 160.9841 | 506.9885
0.68045 | 1000772.5199 | 75.1401 | 181.4041 | 485.4717
0.68549 | 1033394.0434 | 70.0005 | 168.9961 | 490.4325
0.68534 | 1024591.5245 | 59.5882 | 143.8588 | 495.6728
0.70044 | 1082280.4819 | 71.7473 | 173.2132 | 548.2418
0.70501 | 1067810.1261 | 83.8406 | 202.4092 | 507.382
0.69154 | 1037745.888 | 71.4911 | 172.5949 | 494.9018
0.67595 | 1003557.3542 | 80.6463 | 194.6973 | 495.4077
0.67223 | 1000937.979 | 72.6885 | 175.4856 | 485.3126
0.69205 | 1056880.4902 | 74.5525 | 179.9857 | 475.0364
0.69981 | 1062520.451 | 51.1455 | 123.4761 | 513.1973
0.68177 | 1022548.4449 | 74.942 | 180.926 | 448.2986
0.7036 | 1060176.7856 | 62.166 | 150.0821 | 514.7881
0.69205 | 1039132.9888 | 73.3007 | 176.9635 | 467.132
0.69051 | 1038224.437 | 67.9541 | 164.0557 | 493.5829
0.70061 | 1065043.1444 | 72.4221 | 174.8424 | 472.1444
0.68882 | 1044678.9888 | 74.0369 | 178.7409 | 506.0117
0.67806 | 1030901.06 | 56.1033 | 135.4454 | 493.3245
0.68966 | 1033690.8428 | 67.7413 | 163.5419 | 458.5065
0.69048 | 1047622.0631 | 91.4346 | 220.7426 | 516.8708
0.6722 | 1009117.9537 | 59.8362 | 144.4573 | 514.5559
0.70044 | 1079908.4781 | 71.5899 | 172.8333 | 497.0471
0.69299 | 1039581.6035 | 69.6572 | 168.1675 | 473.2875
0.6892 | 1050594.0916 | 70.7829 | 170.8851 | 490.7172
0.68676 | 1021420.5649 | 75.9851 | 183.4442 | 516.8808
0.71033 | 1080077.6361 | 68.7215 | 165.9084 | 450.7218
0.69982 | 1066445.3386 | 72.5989 | 175.2691 | 493.4877
0.70689 | 1079356.732 | 78.0092 | 188.3308 | 478.7159
0.66604 | 996354.0911 | 69.4623 | 167.6967 | 486.6741
0.67332 | 983146.4714 | 83.7666 | 202.2306 | 500.7109
0.66651 | 983611.0902 | 87.5337 | 211.325 | 493.8455
0.68578 | 1040223.2555 | 64.6333 | 156.0385 | 480.2752
0.68517 | 1011372.1219 | 68.5406 | 165.4717 | 510.8751
0.69939 | 1061030.9507 | 66.6329 | 160.8661 | 496.8885
0.70571 | 1090588.6713 | 67.8758 | 163.8667 | 487.1582
0.66997 | 991160.5655 | 72.2212 | 174.3575 | 494.0299
0.68562 | 1029908.3502 | 75.0572 | 181.204 | 487.3727
0.68524 | 1039708.3292 | 72.2844 | 174.5101 | 485.5351
0.71344 | 1105285.8385 | 70.0184 | 169.0393 | 465.1572
0.69158 | 1038102.4284 | 67.8409 | 163.7825 | 521.9121
0.69164 | 1062847.9999 | 94.0432 | 227.0403 | 490.466
0.70903 | 1087380.4459 | 86.4514 | 208.7123 | 511.5409
0.67081 | 981797.6314 | 60.3842 | 145.7803 | 485.9358
0.69603 | 1056523.7314 | 77.5503 | 187.223 | 488.8946
0.68423 | 1025791.6955 | 71.4225 | 172.4292 | 466.4334
0.70109 | 1073308.5895 | 78.2134 | 188.8238 | 464.4058
0.68939 | 1045462.1908 | 75.7738 | 182.9341 | 518.0104
0.68107 | 1004849.0212 | 81.6707 | 197.1705 | 498.5188
0.67616 | 1004179.9564 | 59.1939 | 142.9068 | 497.1968
0.67695 | 1013080.4572 | 69.49 | 167.7636 | 477.651
0.66431 | 977458.4956 | 88.6709 | 214.0704 | 448.7758
0.69685 | 1053215.2418 | 66.3832 | 160.2632 | 500.3818
0.70208 | 1079389.1903 | 65.5105 | 158.1564 | 468.8447
0.70302 | 1081703.5415 | 80.3002 | 193.8619 | 481.7869
0.69786 | 1043666.742 | 78.0492 | 188.4274 | 486.7221
0.67851 | 981400.6074 | 73.8955 | 178.3995 | 511.9815
0.69023 | 1043793.3484 | 78.9211 | 190.5323 | 485.959
0.72143 | 1125439.0909 | 79.7005 | 192.4139 | 503.0174
0.69987 | 1063384.2389 | 64.7423 | 156.3019 | 529.6129
0.67085 | 994736.1458 | 74.0899 | 178.8688 | 479.1588
0.69514 | 1052230.0858 | 78.9971 | 190.7159 | 506.3237
0.66428 | 980981.1929 | 80.2166 | 193.6599 | 508.317
0.7036 | 1070435.6551 | 68.7595 | 166.0001 | 499.8782
0.70787 | 1081616.5088 | 65.1084 | 157.1856 | 489.7431
0.6681 | 985168.7268 | 76.9561 | 185.7884 | 498.0538
0.70744 | 1089325.4673 | 69.5931 | 168.0126 | 513.7365
0.69367 | 1055607.0713 | 58.4488 | 141.1079 | 517.6505
0.68908 | 1052563.5047 | 81.3613 | 196.4236 | 508.4374
0.70592 | 1069630.4958 | 73.9694 | 178.578 | 519.3591
0.68231 | 1003108.3234 | 68.2662 | 164.8093 | 503.5827
0.6899 | 1048416.5969 | 53.6898 | 129.6187 | 481.6901
0.7209 | 1120188.0714 | 59.6816 | 144.0842 | 499.8796
0.71631 | 1089080.6276 | 80.9836 | 195.5117 | 496.0287
0.70532 | 1077818.1883 | 74.5987 | 180.0972 | 498.8308
0.69654 | 1068213.8128 | 65.3479 | 157.7639 | 494.3761
0.68997 | 1036212.1431 | 69.7363 | 168.3583 | 482.075
0.70764 | 1087782.0315 | 67.5166 | 162.9995 | 453.2606
0.66347 | 969065.3249 | 62.2369 | 150.2531 | 500.2777
0.70099 | 1058270.2634 | 68.2309 | 164.724 | 504.2577
0.67087 | 1001985.3387 | 72.9116 | 176.0242 | 489.4328
0.6977 | 1054515.6576 | 66.3841 | 160.2654 | 479.216
0.70388 | 1073123.6157 | 70.7924 | 170.908 | 481.9374
0.67553 | 1005356.0011 | 72.6519 | 175.3973 | 507.8539
0.67396 | 994487.515 | 70.2977 | 169.7137 | 500.5181
0.7152 | 1131669.6332 | 67.4227 | 162.7728 | 446.6552
0.68459 | 1034820.6845 | 58.058 | 140.1643 | 496.3387
0.70473 | 1064695.8823 | 58.0269 | 140.0893 | 475.8897
0.69318 | 1029898.6476 | 68.9904 | 166.5575 | 459.7398
0.69061 | 1052346.9447 | 79.5242 | 191.9884 | 481.0613
0.6828 | 1025397.9707 | 63.9903 | 154.4863 | 472.829
0.6648 | 981979.1476 | 78.6776 | 189.9446 | 508.2647
0.7173 | 1108737.1133 | 78.9692 | 190.6484 | 512.0395
0.66421 | 974366.4633 | 49.4161 | 119.301 | 487.6313
0.66144 | 979565.4968 | 56.8814 | 137.3238 | 473.2622
0.69854 | 1051935.8628 | 66.1429 | 159.683 | 478.4667
0.68149 | 1031710.3669 | 60.3419 | 145.6781 | 474.3358
0.7118 | 1097379.0591 | 75.2609 | 181.6959 | 488.1974
0.71985 | 1113336.6166 | 59.403 | 143.4115 | 479.7176
0.71162 | 1101381.3705 | 82.438 | 199.0228 | 467.4069
0.69615 | 1067740.8513 | 74.8746 | 180.7633 | 489.927
0.69913 | 1057778.1931 | 78.9606 | 190.6279 | 516.0417
0.71155 | 1099456.0243 | 80.6394 | 194.6808 | 477.6995
0.6803 | 1025612.1193 | 73.596 | 177.6765 | 500.86
0.67955 | 1018658.4736 | 84.5012 | 204.004 | 479.1438
0.67955 | 1021432.3507 | 90.5827 | 218.686 | 521.9933
0.6814 | 1005289.4254 | 68.5056 | 165.3872 | 481.6249
0.70083 | 1052426.21 | 63.9871 | 154.4786 | 495.5357
0.6806 | 1032887.9557 | 70.4693 | 170.128 | 497.3312
0.67901 | 1007940.344 | 65.1989 | 157.404 | 499.5701
0.68175 | 1023671.7638 | 73.3004 | 176.9627 | 511.0069
0.69014 | 1028919.3637 | 69.2277 | 167.1304 | 503.555
0.67484 | 992455.6451 | 71.7931 | 173.3238 | 484.702
0.69218 | 1055717.502 | 79.5976 | 192.1655 | 497.372
0.69476 | 1053683.0031 | 68.5445 | 165.481 | 524.2179
0.68789 | 1036322.3049 | 73.544 | 177.551 | 488.9108
0.68754 | 1036856.4456 | 70.0256 | 169.0567 | 485.9273
0.69121 | 1035449.1817 | 79.7913 | 192.6332 | 497.4095
0.70886 | 1098681.9182 | 77.4442 | 186.9668 | 485.9928
0.67439 | 1005276.1424 | 52.6645 | 127.1434 | 479.4702
0.68772 | 1026577.8355 | 76.2627 | 184.1146 | 469.0177
0.7254 | 1128559.415 | 68.3332 | 164.971 | 515.1581
0.71623 | 1109378.4398 | 80.8356 | 195.1544 | 497.5917
0.67345 | 994185.8301 | 73.7629 | 178.0794 | 501.8881
0.70045 | 1057456.9099 | 71.922 | 173.6351 | 495.9108
0.68922 | 1038699.8187 | 76.9949 | 185.8822 | 498.3666
0.70152 | 1079013.1645 | 83.5493 | 201.7059 | 479.5893
0.64986 | 952722.9836 | 76.5101 | 184.7117 | 512.9389
0.69164 | 1043895.1883 | 73.6804 | 177.8803 | 481.446
0.69847 | 1057072.4927 | 78.9823 | 190.6801 | 489.0362
0.67374 | 1009815.3586 | 71.4045 | 172.3856 | 475.7275
0.70523 | 1083035.7196 | 81.8001 | 197.4829 | 487.6941
0.67536 | 1001315.1988 | 78.9724 | 190.6562 | 502.4476
0.69075 | 1049545.0093 | 64.7436 | 156.3049 | 512.0581
0.67852 | 1016359.1187 | 67.699 | 163.4399 | 504.3825
0.68512 | 1032456.0572 | 68.1083 | 164.428 | 486.8315
0.69679 | 1047224.7306 | 66.0112 | 159.3652 | 502.2414
0.71011 | 1097842.5762 | 73.6528 | 177.8136 | 492.2224
0.68696 | 1037709.0492 | 66.3475 | 160.1771 | 478.0019
0.70113 | 1070008.0531 | 76.2545 | 184.0947 | 487.4641
0.68508 | 1040836.7322 | 62.6821 | 151.3281 | 491.8683
0.67999 | 1009124.5456 | 67.0713 | 161.9243 | 496.8865
0.68481 | 1020959.4761 | 76.1674 | 183.8843 | 506.8358
0.72214 | 1140697.4389 | 75.2531 | 181.6771 | 488.6234
0.69207 | 1056215.8054 | 72.7073 | 175.5311 | 495.8861
0.70134 | 1061629.2636 | 80.6272 | 194.6513 | 478.3944
0.69671 | 1060287.3582 | 90.2676 | 217.9253 | 473.2566
0.6859 | 1035447.2776 | 75.4298 | 182.1037 | 504.8871
0.6966 | 1057948.3234 | 62.6588 | 151.2718 | 494.7961
0.69745 | 1074109.364 | 63.5387 | 153.396 | 506.2395
0.71072 | 1085150.0883 | 89.0903 | 215.083 | 481.5979
0.68896 | 1061439.7331 | 78.71 | 190.0227 | 499.9314
0.68148 | 1011729.3833 | 77.4794 | 187.0519 | 485.2866
0.67694 | 1013237.796 | 76.4706 | 184.6163 | 500.5933
0.684 | 1032554.768 | 66.3044 | 160.073 | 491.7908
0.68846 | 1038735.7539 | 71.8796 | 173.5327 | 486.1468
0.70762 | 1099985.8246 | 82.7993 | 199.8953 | 483.2285
0.69953 | 1052913.5537 | 81.7601 | 197.3863 | 498.9402
0.69814 | 1056925.5852 | 68.4221 | 165.1856 | 445.3625
0.69882 | 1070781.4368 | 76.5622 | 184.8374 | 468.0676
0.68168 | 1001983.3893 | 62.6322 | 151.2075 | 507.7676
0.69823 | 1057351.3916 | 86.875 | 209.7347 | 492.3623
0.68737 | 1050750.6963 | 69.0883 | 166.7939 | 493.0447
0.69677 | 1073285.4933 | 80.3901 | 194.0788 | 486.9976
0.68739 | 1049488.8285 | 69.0038 | 166.5899 | 500.5124
0.69531 | 1050167.0955 | 75.4667 | 182.1927 | 490.6257
0.68645 | 1044421.7474 | 77.7319 | 187.6614 | 499.4882
0.68695 | 1035337.1961 | 69.9829 | 168.9538 | 517.1768
0.68335 | 1016101.8859 | 79.2017 | 191.2098 | 498.2637
0.68681 | 1027037.9478 | 82.7759 | 199.8387 | 477.4036
0.66836 | 991123.937 | 80.5518 | 194.4692 | 497.6782
0.6858 | 1011314.7655 | 68.4743 | 165.3116 | 473.9624
0.68044 | 1015455.1172 | 77.8839 | 188.0283 | 508.3113
0.70984 | 1092102.1252 | 76.8739 | 185.5899 | 510.792
0.70952 | 1096469.9278 | 77.2164 | 186.4169 | 500.2554
0.6749 | 987016.5171 | 71.3937 | 172.3596 | 462.3446
0.69063 | 1040960.7157 | 75.3125 | 181.8204 | 508.7151
0.71543 | 1114820.1442 | 66.3976 | 160.2979 | 473.7064
0.67548 | 1023732.2825 | 72.2028 | 174.313 | 478.9773
0.68529 | 1020917.2894 | 65.4329 | 157.969 | 492.8791
0.67904 | 1025711.3582 | 68.269 | 164.8159 | 472.417
0.69386 | 1058930.5404 | 65.0296 | 156.9954 | 471.4495
0.71911 | 1130145.8947 | 69.0272 | 166.6465 | 506.6959
0.68571 | 1006414.9077 | 76.5966 | 184.9206 | 501.5509
0.68997 | 1022047.9412 | 72.3127 | 174.5783 | 455.3604
0.69641 | 1063698.5584 | 56.3631 | 136.0726 | 482.4169
0.68812 | 1048320.6469 | 81.1458 | 195.9034 | 474.9179
0.68896 | 1038008.7531 | 83.3989 | 201.3428 | 515.5234
0.6755 | 1001771.4812 | 80.5568 | 194.4814 | 510.7348
0.6898 | 1041484.1304 | 80.4211 | 194.1537 | 455.46
0.68744 | 1045660.8104 | 93.088 | 224.7343 | 503.9023
0.70149 | 1061764.5149 | 75.6285 | 182.5834 | 513.9289
0.7109 | 1072510.6148 | 86.5503 | 208.9508 | 480.9971
0.67178 | 987513.9901 | 84.3311 | 203.5932 | 520.98
0.68871 | 1052719.2482 | 70.9762 | 171.3518 | 493.134
0.69029 | 1046624.3363 | 82.3606 | 198.836 | 475.7021
0.70585 | 1094981.2819 | 82.6288 | 199.4835 | 547.1409
0.69431 | 1041994.3131 | 78.3219 | 189.0857 | 461.5599
0.69931 | 1081064.2545 | 85.5731 | 206.5918 | 508.4306
0.70082 | 1058458.6113 | 70.1307 | 169.3104 | 486.7187
0.69569 | 1055711.3313 | 72.2944 | 174.5342 | 482.8556
0.68639 | 1036159.4874 | 75.4282 | 182.0997 | 477.755
0.68467 | 1033786.2122 | 78.7993 | 190.2383 | 478.5165
0.68775 | 1055256.2052 | 83.3461 | 201.2153 | 491.2148
0.69278 | 1061067.896 | 83.1961 | 200.8532 | 464.5824
0.70223 | 1057798.7341 | 78.6138 | 189.7906 | 479.6228
0.71255 | 1101142.8238 | 88.655 | 214.0321 | 465.9872
0.67686 | 997647.4996 | 64.7378 | 156.291 | 495.4563
0.70101 | 1072502.9266 | 87.7703 | 211.8962 | 481.6405
0.69143 | 1038930.1099 | 59.8907 | 144.5888 | 472.1442
0.70696 | 1086887.6511 | 75.0522 | 181.1921 | 476.0964
0.67933 | 1018961.8599 | 61.8843 | 149.4019 | 478.4272
0.69234 | 1036722.0157 | 74.8203 | 180.6322 | 508.7683
0.6847 | 1029988.4608 | 71.6656 | 173.016 | 489.5314
0.70921 | 1086502.3381 | 63.2061 | 152.5931 | 490.7232
0.68486 | 1027392.9906 | 76.6413 | 185.0284 | 493.8257
0.71955 | 1109496.3623 | 80.4706 | 194.2732 | 497.6683
0.6981 | 1066818.573 | 67.1204 | 162.0429 | 503.2699
0.65403 | 953593.0799 | 74.4892 | 179.8329 | 508.0608
0.70714 | 1082675.1294 | 67.2469 | 162.3484 | 509.0354
0.68748 | 1041886.9195 | 64.5769 | 155.9025 | 480.5625
0.66882 | 1001171.0678 | 63.7846 | 153.9896 | 496.4708
0.68228 | 1021702.1229 | 74.823 | 180.6387 | 482.8174
0.67028 | 988863.4499 | 76.9935 | 185.8787 | 499.8408
0.68967 | 1037808.5444 | 75.189 | 181.5222 | 447.3635
0.71645 | 1121120.6065 | 83.3495 | 201.2235 | 516.1207
0.68696 | 1011081.5361 | 73.5413 | 177.5444 | 485.309
0.67673 | 1005120.0377 | 79.1121 | 190.9935 | 474.6192
0.67587 | 1006500.3581 | 88.3307 | 213.2491 | 503.5937
0.70889 | 1088991.5354 | 73.1849 | 176.684 | 497.9294
0.70591 | 1069040.9136 | 64.53 | 155.7892 | 484.23
0.67666 | 1011403.3572 | 63.6896 | 153.7603 | 510.4189
0.67022 | 988279.0234 | 73.6787 | 177.8761 | 478.6975
0.7015 | 1076253.3418 | 74.8959 | 180.8146 | 510.123
0.70237 | 1069100.2632 | 66.8547 | 161.4016 | 516.6266
0.6828 | 1021806.0941 | 81.2892 | 196.2494 | 523.3198
0.70059 | 1053947.5517 | 86.3037 | 208.3557 | 503.4665
0.6778 | 1007154.7648 | 67.1572 | 162.1318 | 487.5382
0.67982 | 1020142.0451 | 65.9095 | 159.1196 | 478.9039
0.67532 | 1002264.5876 | 65.1865 | 157.3742 | 501.1461
0.69776 | 1049717.2813 | 71.6714 | 173.0302 | 502.0972
0.6964 | 1060376.4923 | 82.71 | 199.6797 | 484.8754
0.70035 | 1070140.879 | 76.3484 | 184.3213 | 476.7307
0.67076 | 989162.7712 | 58.7781 | 141.9029 | 522.2751
0.70954 | 1086249.2188 | 74.735 | 180.4264 | 481.615
0.70607 | 1094885.5507 | 63.977 | 154.454 | 507.1327
0.6906 | 1035519.0817 | 71.4343 | 172.4577 | 515.4347
0.70097 | 1057042.027 | 71.8406 | 173.4386 | 478.9057
0.67074 | 1010080.9447 | 76.9367 | 185.7416 | 488.8697
0.70793 | 1096737.7021 | 75.629 | 182.5845 | 483.2143
0.68963 | 1052227.7687 | 82.8807 | 200.0916 | 481.8086
0.71519 | 1107122.018 | 57.1239 | 137.9092 | 486.4414
0.69361 | 1056403.6562 | 79.486 | 191.8961 | 507.2926
0.68883 | 1038250.9738 | 55.6201 | 134.2789 | 504.2069
0.68685 | 1048045.9287 | 70.8524 | 171.0528 | 485.6227
0.68977 | 1048687.2214 | 82.585 | 199.3778 | 509.3571
0.67857 | 1020790.4906 | 76.8555 | 185.5455 | 475.1733
0.70205 | 1078312.01 | 67.4624 | 162.8686 | 485.1262
0.68838 | 1025621.7266 | 72.7335 | 175.5943 | 493.1516
0.68629 | 1053689.2012 | 80.127 | 193.4437 | 494.2997
0.69628 | 1053182.171 | 85.2653 | 205.8486 | 520.7877
0.68498 | 1023325.7393 | 74.6467 | 180.2131 | 468.8853
0.68904 | 1034291.944 | 57.6529 | 139.1864 | 494.9771
0.68432 | 1026192.4715 | 73.2054 | 176.7334 | 494.4155
0.66289 | 974340.0077 | 62.6307 | 151.2038 | 522.0975
0.68997 | 1050793.2944 | 60.7023 | 146.5484 | 499.1889
0.70459 | 1094680.8975 | 77.6295 | 187.4142 | 499.5014
0.70029 | 1078640.116 | 80.3845 | 194.0654 | 491.2493
0.67113 | 989323.0288 | 56.6182 | 136.6884 | 451.011
0.69608 | 1055192.9544 | 80.744 | 194.9332 | 477.4199
0.68389 | 1016325.3784 | 75.924 | 183.2967 | 488.9942
0.69215 | 1043359.2593 | 78.6692 | 189.9242 | 497.4982
0.70902 | 1078640.1423 | 68.7564 | 165.9926 | 477.3468
0.68765 | 1036134.3218 | 71.784 | 173.3018 | 505.6818
0.70027 | 1067614.0934 | 64.9625 | 156.8333 | 525.202
0.69514 | 1066621.1638 | 69.3481 | 167.421 | 486.7087
0.68811 | 1046748.6753 | 70.6355 | 170.5291 | 498.7356
0.71286 | 1102142.044 | 71.7908 | 173.3184 | 498.586
0.68185 | 1019339.9464 | 65.6616 | 158.5212 | 482.9356
0.69775 | 1058957.2519 | 66.6594 | 160.9299 | 497.2308
0.6828 | 1011674.5631 | 82.0161 | 198.0043 | 485.3289
0.67915 | 1016901.7971 | 79.7539 | 192.543 | 483.6667
0.69401 | 1048935.9201 | 70.1809 | 169.4317 | 476.7891
0.68425 | 1042392.2908 | 87.3953 | 210.9909 | 473.7364
0.68753 | 1025857.1749 | 67.4358 | 162.8043 | 513.8291
0.6991 | 1052720.0732 | 71.7385 | 173.1921 | 486.7534
0.72005 | 1105447.1262 | 81.7742 | 197.4204 | 511.3013
0.68452 | 1039374.4562 | 72.3378 | 174.6388 | 508.279
0.67916 | 1003228.4892 | 66.7613 | 161.1761 | 490.0318
0.70366 | 1066164.2837 | 70.3559 | 169.8542 | 492.4273
0.72726 | 1134225.8263 | 76.1347 | 183.8055 | 493.3364
0.67931 | 1011978.7515 | 74.4351 | 179.7023 | 488.9058
0.71012 | 1083404.4786 | 72.6836 | 175.4738 | 488.066
0.69218 | 1053885.4437 | 72.5352 | 175.1156 | 495.0926
0.69052 | 1048375.1989 | 68.6177 | 165.6578 | 506.528
0.69022 | 1043785.9249 | 73.8241 | 178.2272 | 488.0667
0.70497 | 1071949.8297 | 66.7865 | 161.2369 | 502.8615
0.70578 | 1086808.6968 | 83.6448 | 201.9363 | 473.3172
0.70405 | 1041359.367 | 77.1923 | 186.3586 | 489.7132
0.68036 | 1026076.1345 | 71.8487 | 173.4582 | 469.6778
0.69981 | 1070379.425 | 79.8236 | 192.7111 | 497.3353
0.69867 | 1042278.6016 | 52.7057 | 127.2429 | 494.1308
0.66355 | 973712.7502 | 68.1391 | 164.5022 | 522.4064
0.68868 | 1031729.8783 | 71.3712 | 172.3054 | 465.5212
0.69343 | 1051138.501 | 72.2162 | 174.3452 | 484.6996
0.69963 | 1081930.1126 | 80.7058 | 194.8411 | 472.6016
0.67671 | 1012696.7137 | 87.3195 | 210.8079 | 509.9817
0.69387 | 1043358.1319 | 69.8219 | 168.5649 | 503.5752
0.67383 | 1006411.8159 | 69.3525 | 167.4317 | 471.7473
0.69842 | 1051605.2645 | 75.2338 | 181.6304 | 490.3413
0.71203 | 1097031.5491 | 69.6332 | 168.1094 | 448.0998
0.6899 | 1052125.7031 | 64.9829 | 156.8825 | 497.6463
0.67592 | 1006285.655 | 85.4085 | 206.1944 | 487.4399
0.69412 | 1040236.1091 | 79.8243 | 192.7129 | 506.2752
0.67524 | 1003725.9858 | 85.2769 | 205.8765 | 518.1016
0.67809 | 1022308.6992 | 66.218 | 159.8645 | 486.8381
0.70004 | 1058163.4184 | 55.7788 | 134.662 | 469.7566
0.65453 | 958841.001 | 69.7905 | 168.4892 | 458.6847
0.69016 | 1053316.1151 | 86.0321 | 207.6999 | 476.6687
0.69315 | 1052010.2178 | 64.6712 | 156.1301 | 479.6686
0.69778 | 1060359.3463 | 64.752 | 156.3252 | 506.3177
0.69571 | 1048034.2522 | 60.0436 | 144.9581 | 479.2575
0.68815 | 1030257.4596 | 78.1338 | 188.6317 | 486.8007
0.68259 | 1019612.9613 | 79.5635 | 192.0833 | 499.4755
0.7124 | 1103963.5852 | 79.1706 | 191.1347 | 467.4083
0.67469 | 1023892.9621 | 60.488 | 146.0309 | 524.5211
0.68898 | 1040178.1456 | 78.7901 | 190.2162 | 476.9037
0.68027 | 1026279.0639 | 71.8725 | 173.5156 | 527.6043
0.67563 | 1008161.383 | 63.5827 | 153.5022 | 505.9536
0.68785 | 1043635.9209 | 68.573 | 165.5499 | 489.7581
0.69442 | 1039659.159 | 73.0876 | 176.449 | 515.8716
0.71189 | 1093418.2345 | 71.3205 | 172.1828 | 514.5513
0.68119 | 1042430.4038 | 89.2922 | 215.5706 | 496.935
0.69129 | 1044658.3657 | 73.7709 | 178.0988 | 505.5547
0.69367 | 1056172.8147 | 70.7001 | 170.6851 | 498.4641
0.68762 | 1028960.2916 | 76.4502 | 184.5672 | 478.0283
0.70517 | 1075268.7665 | 77.9027 | 188.0737 | 489.0264
0.70229 | 1069655.5755 | 48.6741 | 117.5096 | 490.3918
0.68917 | 1034123.8949 | 74.9755 | 181.007 | 501.3518
0.66658 | 993482.438 | 67.3606 | 162.6228 | 475.123
0.69568 | 1045922.9686 | 69.8116 | 168.5401 | 506.3111
0.69054 | 1041033.5582 | 77.0208 | 185.9447 | 508.5754
0.69174 | 1041515.1969 | 73.5013 | 177.4479 | 481.0673
0.70446 | 1077794.2847 | 71.0423 | 171.5112 | 498.715
0.6954 | 1052005.1982 | 66.4451 | 160.4128 | 489.2005
0.68863 | 1036873.134 | 73.5045 | 177.4556 | 478.6758
0.68985 | 1036799.338 | 69.7875 | 168.4819 | 492.7282
0.68673 | 1026002.9241 | 82.7016 | 199.6594 | 478.9118
0.70767 | 1076752.2512 | 74.2781 | 179.3232 | 482.5493
0.70293 | 1060929.045 | 81.9842 | 197.9273 | 484.1402
0.71234 | 1102630.2057 | 73.7428 | 178.0309 | 458.9089
0.67065 | 998052.5022 | 79.2675 | 191.3686 | 515.4971
0.7014 | 1052938.95 | 59.8347 | 144.4538 | 484.029
0.69157 | 1029579.2827 | 67.2849 | 162.4401 | 479.4744
0.69321 | 1041327.9492 | 61.9763 | 149.6241 | 520.2694
0.71801 | 1127339.9899 | 60.304 | 145.5866 | 467.0857
0.70854 | 1084264.7159 | 86.2633 | 208.2581 | 515.2663
0.69287 | 1043651.1297 | 78.6095 | 189.7802 | 490.8937
0.68535 | 1022347.7023 | 74.5357 | 179.945 | 510.3339
0.68355 | 1027855.9287 | 73.407 | 177.2202 | 490.744
0.70858 | 1090298.2375 | 76.8833 | 185.6126 | 485.5491
0.70797 | 1074797.4775 | 75.8555 | 183.1313 | 471.0253
0.68684 | 1029185.8531 | 79.8139 | 192.6879 | 509.122
0.68349 | 1028325.1993 | 75.1747 | 181.4878 | 467.6651
0.70785 | 1093492.5729 | 80.6211 | 194.6365 | 502.2098
0.68741 | 1028928.7697 | 67.6493 | 163.3198 | 481.3656
0.68949 | 1047065.3126 | 90.0754 | 217.4612 | 517.6469
0.68828 | 1057032.1361 | 58.9856 | 142.4039 | 494.6515
0.66892 | 989745.0473 | 70.4901 | 170.1782 | 510.0602
0.69346 | 1047864.9429 | 82.081 | 198.161 | 535.3175
0.69783 | 1061885.5837 | 66.8359 | 161.3561 | 520.51
0.68395 | 1013200.1895 | 63.123 | 152.3924 | 491.6548
0.67825 | 1020412.6009 | 69.8595 | 168.6558 | 458.3061
0.68216 | 1020179.4806 | 71.2472 | 172.0059 | 494.5708
0.67878 | 1025185.7403 | 85.1386 | 205.5427 | 451.2567
0.69049 | 1042429.6708 | 80.4131 | 194.1345 | 490.5184
0.65977 | 969061.0434 | 69.9741 | 168.9324 | 468.5957
0.70147 | 1078479.1843 | 89.7102 | 216.5795 | 506.4068
0.68598 | 1024718.1612 | 76.3175 | 184.2467 | 497.6519
0.70466 | 1080312.0699 | 83.2779 | 201.0506 | 470.9115
0.69409 | 1039266.6031 | 67.672 | 163.3746 | 522.7679
0.69141 | 1043044.7759 | 73.6448 | 177.7942 | 511.6294
0.68924 | 1043772.4562 | 58.1647 | 140.4221 | 521.7802
0.71804 | 1100837.3307 | 54.1653 | 130.7666 | 481.1579
0.69284 | 1060946.4833 | 78.2336 | 188.8726 | 509.3885
0.65744 | 960713.5654 | 62.2665 | 150.3247 | 482.3558
0.69599 | 1032545.9185 | 72.4238 | 174.8465 | 468.4414
0.67599 | 1010274.538 | 76.3535 | 184.3338 | 516.644
0.68928 | 1027334.0838 | 57.2449 | 138.2015 | 503.1325
0.65856 | 976945.717 | 70.6734 | 170.6206 | 502.4693
0.68289 | 1032587.5992 | 77.2517 | 186.5022 | 478.573
0.68674 | 1040633.8723 | 83.8807 | 202.5058 | 496.4373
0.71483 | 1081055.8896 | 68.3509 | 165.0136 | 483.4339
0.71003 | 1083568.0458 | 89.0462 | 214.9766 | 492.4329
0.69752 | 1064259.7785 | 82.8795 | 200.0887 | 477.1348
0.69343 | 1049300.3659 | 77.3085 | 186.6393 | 485.4903
0.71015 | 1079251.1351 | 72.4016 | 174.793 | 483.5339
0.69069 | 1037516.8231 | 76.7434 | 185.275 | 486.7718
0.68816 | 1041558.5358 | 80.6186 | 194.6305 | 509.3435
0.66728 | 984473.9964 | 68.5072 | 165.3909 | 468.2832
0.66378 | 979757.3442 | 90.3301 | 218.0761 | 487.3707
0.67728 | 1019108.52 | 83.293 | 201.0871 | 519.7875
0.68198 | 1013461.3771 | 72.5448 | 175.1387 | 458.7359
0.67581 | 1002202.286 | 77.3936 | 186.8446 | 492.5212
0.68517 | 1030164.4651 | 73.3971 | 177.1962 | 516.7282
0.67287 | 1002026.6361 | 77.7177 | 187.6271 | 539.8291
0.70338 | 1095571.4919 | 82.9644 | 200.2939 | 463.0645
0.6797 | 1010163.8968 | 54.8423 | 132.4011 | 475.915
0.7031 | 1083407.0915 | 63.5736 | 153.4802 | 510.3453
0.70184 | 1078715.4736 | 71.3685 | 172.2987 | 488.9799
0.67676 | 1001997.0165 | 83.4614 | 201.4936 | 503.291
0.70149 | 1065283.6518 | 72.3476 | 174.6626 | 506.6307
0.70304 | 1076115.6105 | 84.729 | 204.5539 | 485.1718
0.68254 | 1029218.5512 | 71.838 | 173.4323 | 512.896
0.69909 | 1072101.0946 | 69.31 | 167.3292 | 501.9532
0.70128 | 1077679.3056 | 85.0648 | 205.3645 | 489.9496
0.70631 | 1097871.9414 | 77.6661 | 187.5025 | 492.3082
0.69537 | 1060292.1997 | 70.8019 | 170.931 | 471.6692
0.68385 | 1013708.1317 | 59.0837 | 142.6407 | 476.0411
0.6797 | 1010156.4944 | 61.3159 | 148.0297 | 512.3154
0.69593 | 1049372.5626 | 64.2514 | 155.1165 | 515.8701
0.67075 | 982556.7568 | 68.0196 | 164.2139 | 523.1765
0.6855 | 1023821.3415 | 74.6268 | 180.1651 | 477.2543
0.66834 | 1008563.7318 | 70.0722 | 169.1692 | 492.7238
0.69821 | 1062985.0678 | 60.6819 | 146.4991 | 498.414
0.69222 | 1056577.1613 | 68.9847 | 166.5437 | 489.3662
0.69068 | 1051396.8088 | 72.5214 | 175.0821 | 485.6815
0.68901 | 1060146.1563 | 94.1623 | 227.3278 | 472.8413
0.71994 | 1120936.1213 | 74.1754 | 179.0753 | 514.0685
0.69002 | 1040581.2429 | 62.2175 | 150.2064 | 481.6794
0.68561 | 1030541.4303 | 75.1045 | 181.3183 | 484.6757
0.68555 | 1033728.0322 | 84.9996 | 205.2071 | 471.2975
0.67327 | 974287.5166 | 77.0793 | 186.0859 | 475.13
0.65755 | 983818.1596 | 79.6945 | 192.3997 | 490.5127
0.7145 | 1103565.666 | 87.0654 | 210.1945 | 508.8302
0.66836 | 995679.2484 | 74.437 | 179.7067 | 496.6555
0.66572 | 984203.6445 | 91.8367 | 221.7133 | 468.9659
0.71493 | 1117502.8679 | 60.0355 | 144.9386 | 491.1934
0.67688 | 1021995.0748 | 80.4223 | 194.1566 | 481.4279
0.71503 | 1092154.4972 | 82.9688 | 200.3043 | 482.1227
0.68579 | 1043153.7735 | 84.1998 | 203.2764 | 477.6503
0.66957 | 983698.5385 | 71.7198 | 173.1469 | 491.362
0.68738 | 1062088.4383 | 82.3012 | 198.6926 | 479.8704
0.6956 | 1061605.6895 | 67.0328 | 161.8314 | 536.7443
0.69474 | 1057488.849 | 76.0549 | 183.6128 | 487.3894
0.70435 | 1068893.6167 | 70.4669 | 170.122 | 477.2464
0.69649 | 1051980.516 | 75.4695 | 182.1996 | 505.5904
0.69253 | 1052379.6983 | 77.6368 | 187.4317 | 482.4865
0.66544 | 974841.2224 | 68.0246 | 164.2259 | 492.2951
0.70366 | 1075521.1963 | 67.133 | 162.0733 | 496.4317
0.70642 | 1066272.2943 | 70.0879 | 169.2071 | 516.604
0.67482 | 1003952.1383 | 76.0075 | 183.4984 | 504.8494
0.69266 | 1029239.0685 | 75.9147 | 183.2742 | 494.5351
0.70199 | 1057322.6835 | 78.6049 | 189.7691 | 514.2297
0.68966 | 1041390.1343 | 64.3425 | 155.3366 | 501.3606
0.68975 | 1033978.4568 | 79.8472 | 192.7681 | 473.9721
0.70232 | 1070100.3344 | 77.8436 | 187.9311 | 483.5698
0.68639 | 1026459.1113 | 57.4373 | 138.666 | 516.3641
0.70553 | 1071302.449 | 70.5078 | 170.2208 | 517.7991
0.70029 | 1062447.168 | 72.2779 | 174.4944 | 495.8543
0.70654 | 1078295.4771 | 70.8666 | 171.0871 | 534.2547
0.6735 | 997250.2543 | 63.0934 | 152.3208 | 476.7613
0.71379 | 1078612.0438 | 66.3707 | 160.2331 | 528.9246
0.6615 | 980756.3896 | 56.9453 | 137.4782 | 503.7575
0.68255 | 1033357.8569 | 62.4623 | 150.7974 | 493.8715
0.68263 | 1021845.9063 | 67.6548 | 163.3332 | 512.5608
0.70182 | 1089718.1106 | 82.7039 | 199.665 | 481.4853
0.69696 | 1054747.5138 | 73.8778 | 178.3569 | 473.687
0.68907 | 1039160.6303 | 71.845 | 173.4493 | 502.8417
0.68049 | 1021565.5983 | 76.6961 | 185.1608 | 484.4017
0.69778 | 1060778.4172 | 74.2132 | 179.1665 | 491.7302
0.68404 | 1019849.4327 | 63.5289 | 153.3724 | 490.5802
0.66989 | 990879.5901 | 64.9684 | 156.8475 | 471.3947
0.67156 | 982528.669 | 64.2599 | 155.137 | 485.4639
0.68313 | 1015462.8614 | 74.2741 | 179.3135 | 513.249
0.6961 | 1065124.4402 | 72.8963 | 175.9873 | 479.4085
0.68567 | 1055041.6816 | 73.305 | 176.9738 | 485.0562
0.68448 | 1020624.5881 | 72.7911 | 175.7332 | 455.7547
0.68318 | 1021709.8176 | 87.2627 | 210.6707 | 507.8763
0.67904 | 1010890.8248 | 69.127 | 166.8874 | 509.4572
0.68354 | 1026698.3415 | 73.3254 | 177.0231 | 491.2093
0.68574 | 1040240.0233 | 74.0537 | 178.7814 | 459.4215
0.71512 | 1105638.0608 | 80.6882 | 194.7986 | 498.4954
0.70346 | 1061644.8918 | 78.7619 | 190.1481 | 475.4151
0.70332 | 1068438.8073 | 79.2815 | 191.4024 | 470.7785
0.69209 | 1049463.9157 | 66.6019 | 160.7912 | 475.5473
0.68335 | 1020507.8855 | 71.1462 | 171.7622 | 516.4241
0.66897 | 985445.4701 | 78.4629 | 189.4261 | 503.9745
0.71848 | 1109774.2483 | 71.7226 | 173.1537 | 488.3556
0.69832 | 1030395.5525 | 87.5207 | 211.2937 | 498.9639
0.67336 | 998694.1376 | 68.8691 | 166.2648 | 478.9456
0.70879 | 1078903.5407 | 74.3082 | 179.3959 | 493.1475
0.67675 | 1016215.4204 | 73.3043 | 176.9723 | 461.1755
0.70273 | 1059758.1411 | 71.845 | 173.4493 | 461.7034
0.68704 | 1036559.6725 | 81.9537 | 197.8538 | 456.7605
0.68346 | 1019335.2231 | 63.551 | 153.4256 | 497.5911
0.6914 | 1044663.5606 | 75.4963 | 182.2642 | 497.9712
0.67273 | 1010331.2103 | 75.0414 | 181.1659 | 501.5486
0.66422 | 988956.2156 | 80.8765 | 195.2532 | 483.0823
0.66214 | 981209.8301 | 74.0435 | 178.7568 | 484.5669
0.69369 | 1065147.0768 | 88.0883 | 212.6639 | 492.903
0.70246 | 1095944.1891 | 84.7471 | 204.5977 | 472.3302
0.70097 | 1077355.7435 | 71.367 | 172.2952 | 500.7093
0.66862 | 992123.3478 | 80.6022 | 194.5909 | 512.236
0.72176 | 1112015.3281 | 88.3873 | 213.3857 | 483.0295
0.68301 | 1029522.493 | 68.1244 | 164.4669 | 491.2605
0.69204 | 1046796.0661 | 75.5802 | 182.4668 | 498.1705
0.70547 | 1097771.4731 | 70.3286 | 169.7882 | 491.9404
0.69282 | 1055674.5593 | 77.8466 | 187.9384 | 485.6729
0.67357 | 989564.4379 | 60.6124 | 146.3312 | 493.4589
0.68718 | 1036567.1627 | 78.7224 | 190.0528 | 495.4778
0.69075 | 1051783.6296 | 72.5408 | 175.1289 | 503.49
0.69897 | 1064083.3848 | 76.0662 | 183.6401 | 489.5916
0.68841 | 1059401.8798 | 59.107 | 142.6969 | 483.3254
0.69695 | 1065722.4918 | 78.1216 | 188.6023 | 477.4403
0.66759 | 977134.8318 | 71.4528 | 172.5022 | 464.5396
0.69118 | 1034500.7833 | 74.7852 | 180.5474 | 472.8744
0.72074 | 1111059.3118 | 71.5808 | 172.8114 | 487.2928
0.69921 | 1059154.0679 | 64.5459 | 155.8276 | 484.1904
0.69393 | 1050883.7928 | 72.147 | 174.1783 | 474.7093
0.70566 | 1083990.5258 | 88.1263 | 212.7556 | 513.5886
0.66937 | 1001992.8264 | 67.6537 | 163.3304 | 506.4076
0.69087 | 1051772.5957 | 77.7773 | 187.771 | 478.4753
0.7003 | 1063538.3817 | 60.5322 | 146.1377 | 463.4515
0.68513 | 1025640.6588 | 78.1257 | 188.6121 | 491.8215
0.70524 | 1087721.055 | 69.7074 | 168.2887 | 514.1224
0.7029 | 1068369.0191 | 72.4109 | 174.8155 | 526.0581
0.67365 | 1005279.0418 | 69.2932 | 167.2885 | 479.7601
0.69141 | 1030475.6325 | 63.5069 | 153.3191 | 495.3157
0.70505 | 1082891.0367 | 71.3187 | 172.1787 | 526.8159
0.6803 | 1019151.2485 | 69.5625 | 167.9387 | 477.1389
0.69783 | 1055729.0874 | 73.8543 | 178.3002 | 508.3394
0.69024 | 1029648.0858 | 87.0413 | 210.1363 | 495.154
0.69236 | 1033609.2071 | 71.1213 | 171.702 | 498.9429
0.7041 | 1078289.6321 | 84.7729 | 204.66 | 476.9993
0.6807 | 1020978.3826 | 71.4558 | 172.5096 | 492.4935
0.71088 | 1075130.9881 | 62.3977 | 150.6413 | 473.2151
0.69521 | 1035796.0595 | 67.3372 | 162.5664 | 460.9844
0.67966 | 1005144.2608 | 70.4554 | 170.0943 | 488.888
0.66182 | 962758.706 | 67.5488 | 163.0772 | 514.4674
0.68536 | 1022414.5462 | 81.0339 | 195.6332 | 501.136
0.69158 | 1046703.1724 | 77.3233 | 186.6751 | 473.6773
0.67503 | 1013188.5384 | 74.9969 | 181.0586 | 487.8375
0.6768 | 1017774.7726 | 73.4115 | 177.2311 | 488.9048
0.69657 | 1063800.1754 | 72.757 | 175.651 | 479.4104
0.69175 | 1054870.0563 | 76.1955 | 183.9521 | 485.4727
0.6864 | 1037870.0184 | 64.4291 | 155.5456 | 480.1826
0.68357 | 1032102.3091 | 64.3363 | 155.3216 | 475.6904
0.70561 | 1064871.3434 | 70.0758 | 169.1779 | 500.3646
0.67238 | 1009617.4422 | 65.9512 | 159.2202 | 490.7296
0.68909 | 1046469.5799 | 77.5855 | 187.3079 | 474.3852
0.71469 | 1097256.4737 | 58.968 | 142.3614 | 503.7217
0.69348 | 1033071.1776 | 63.4769 | 153.2468 | 501.1361
0.71099 | 1095667.3648 | 85.3043 | 205.9427 | 476.621
0.69924 | 1060839.2311 | 95.6179 | 230.842 | 499.0411
0.69908 | 1051336.2889 | 73.4153 | 177.2403 | 490.2453
0.69283 | 1047934.64 | 73.8382 | 178.2612 | 500.9791
0.69936 | 1059618.4763 | 79.0722 | 190.8972 | 452.1884
0.6903 | 1049596.1016 | 58.3997 | 140.9893 | 473.8904
0.6965 | 1038732.781 | 71.0495 | 171.5287 | 489.4139
0.70395 | 1085423.4 | 63.6145 | 153.579 | 461.0146
0.69964 | 1064366.8501 | 72.4759 | 174.9723 | 484.5978
0.66167 | 976627.7172 | 87.5474 | 211.3581 | 524.821
0.7162 | 1122238.5107 | 81.776 | 197.4247 | 454.0774
0.68663 | 1036044.8623 | 90.9309 | 219.5267 | 510.4416
0.68297 | 1032048.1088 | 78.8627 | 190.3915 | 520.3884
0.69503 | 1039349.4255 | 74.7192 | 180.3882 | 464.3459
0.6891 | 1025599.1151 | 77.6727 | 187.5185 | 494.3269
0.66581 | 977809.7738 | 75.0297 | 181.1378 | 520.0691
0.69237 | 1040420.3603 | 71.5894 | 172.8321 | 478.7626
0.68376 | 1029676.7636 | 68.0601 | 164.3116 | 484.8566
0.698 | 1065531.1651 | 74.5214 | 179.9106 | 490.6824
0.68907 | 1037721.0358 | 68.0634 | 164.3196 | 508.0997
0.69377 | 1073021.4554 | 75.5037 | 182.2821 | 516.3975
0.67143 | 1001545.0516 | 63.5604 | 153.4483 | 486.815
0.67337 | 1002391.7559 | 74.3809 | 179.5715 | 496.7928
0.70215 | 1045838.2836 | 67.3183 | 162.5208 | 483.7592
0.70079 | 1089624.9299 | 70.2724 | 169.6525 | 495.6203
0.67318 | 1001016.6365 | 72.5916 | 175.2517 | 520.9497
0.7103 | 1098720.761 | 65.9115 | 159.1244 | 511.1017
0.69168 | 1051900.3358 | 87.2461 | 210.6308 | 474.9231
0.70686 | 1078676.3978 | 72.7002 | 175.5137 | 474.0386
0.6722 | 1014596.1362 | 70.0857 | 169.2018 | 497.3668
0.70185 | 1061026.2272 | 58.0638 | 140.1785 | 470.9854
0.67907 | 1002057.4374 | 80.1555 | 193.5124 | 492.104
0.70578 | 1083443.4343 | 76.7034 | 185.1783 | 519.5368
0.6635 | 967847.1403 | 69.4929 | 167.7706 | 465.8582
0.69462 | 1053635.8266 | 68.5554 | 165.5074 | 477.8201
0.7014 | 1052441.6256 | 78.308 | 189.0523 | 492.1734
0.67812 | 1013110.1608 | 85.7085 | 206.9185 | 509.519
0.68503 | 1027218.3912 | 69.6286 | 168.0984 | 491.3027
0.6703 | 987494.1917 | 62.774 | 151.5499 | 499.3314
0.7002 | 1083106.3348 | 63.8186 | 154.0717 | 513.3884
0.7216 | 1119748.3693 | 54.8238 | 132.3564 | 507.4573
0.69112 | 1042684.171 | 78.7356 | 190.0845 | 492.4729
0.71519 | 1092870.0034 | 90.6366 | 218.816 | 479.4944
0.6727 | 1010769.6066 | 91.9542 | 221.9971 | 503.0136
0.69824 | 1052634.3879 | 80.2999 | 193.861 | 478.7544
0.7124 | 1107234.0915 | 66.2263 | 159.8844 | 482.1579
0.68369 | 999041.604 | 86.6729 | 209.2469 | 540.0569
0.72405 | 1129067.422 | 72.4089 | 174.8105 | 462.94
0.68737 | 1029856.2314 | 69.5703 | 167.9575 | 503.0816
0.6909 | 1040854.8023 | 76.9671 | 185.8151 | 502.8706
0.69191 | 1037117.8422 | 59.7447 | 144.2364 | 481.353
0.69181 | 1039043.7753 | 79.9993 | 193.1353 | 474.0356
0.70575 | 1062317.1078 | 94.8687 | 229.0334 | 484.6803
0.68995 | 1048549.877 | 79.3135 | 191.4797 | 505.3955
0.69531 | 1050875.6997 | 78.8766 | 190.4251 | 509.6621
0.66963 | 999817.2515 | 71.1315 | 171.7267 | 507.9287
0.65986 | 986815.9489 | 59.608 | 143.9065 | 468.719
0.69216 | 1035315.8765 | 79.6716 | 192.3444 | 483.1757
0.69544 | 1060211.03 | 76.1742 | 183.9008 | 504.4762
0.67278 | 988605.9186 | 81.3397 | 196.3713 | 485.4926
0.6884 | 1029438.6126 | 63.7197 | 153.833 | 495.9206
0.69225 | 1045061.6231 | 70.0996 | 169.2354 | 478.6383
0.67881 | 1032201.4326 | 70.6079 | 170.4625 | 500.74
0.68434 | 1022765.0145 | 76.3549 | 184.337 | 486.6843
0.69865 | 1060858.1457 | 80.8796 | 195.2606 | 517.3743
0.70163 | 1071646.8992 | 87.6227 | 211.54 | 492.2243
0.72488 | 1127625.3197 | 70.306 | 169.7336 | 487.1988
0.69486 | 1041996.0857 | 73.205 | 176.7325 | 475.7838
0.69195 | 1056680.0985 | 58.6536 | 141.6023 | 500.56
0.67476 | 998146.744 | 66.8566 | 161.406 | 502.0684
0.68664 | 1018434.4414 | 68.8717 | 166.2709 | 500.3075
0.69324 | 1051338.6179 | 75.7762 | 182.9399 | 480.2202
0.66551 | 991146.4658 | 70.9514 | 171.2918 | 464.9866
0.68938 | 1030745.0991 | 69.427 | 167.6116 | 490.0453
0.67196 | 993042.9559 | 73.8419 | 178.2701 | 475.2022
0.68385 | 1024450.5589 | 79.7939 | 192.6394 | 511.1163
0.66412 | 971460.9211 | 66.1115 | 159.6072 | 474.86
0.6773 | 1008549.9272 | 69.1436 | 166.9274 | 506.5769
0.67737 | 1024664.13 | 77.2837 | 186.5793 | 499.125
0.71851 | 1112705.5064 | 80.8201 | 195.1169 | 507.0307
0.70108 | 1073257.3261 | 81.5413 | 196.8582 | 482.749
0.68146 | 1015253.0608 | 65.4361 | 157.9768 | 509.235
0.6965 | 1059285.9925 | 74.2448 | 179.2428 | 526.0059
0.7068 | 1086730.4553 | 85.1094 | 205.4722 | 489.4506
0.69917 | 1071259.4957 | 72.9939 | 176.2228 | 499.7328
0.71004 | 1085968.4622 | 72.8641 | 175.9095 | 485.0234
0.70117 | 1066167.9271 | 90.1909 | 217.7401 | 493.0789
0.66848 | 995333.4565 | 63.4446 | 153.1687 | 504.6354
0.69784 | 1056007.4957 | 77.311 | 186.6453 | 506.2472
0.6807 | 1006940.8262 | 80.3534 | 193.9902 | 475.7369
0.69113 | 1038008.5293 | 63.9965 | 154.5012 | 503.2525
0.68988 | 1048429.7854 | 77.6424 | 187.4454 | 525.7365
0.69639 | 1066551.8884 | 72.9632 | 176.1487 | 468.9207
0.65808 | 967581.1178 | 79.867 | 192.8159 | 471.4436
0.68768 | 1046236.3028 | 84.2166 | 203.3169 | 489.045
0.70106 | 1083886.6667 | 83.9818 | 202.7499 | 501.0946
0.6942 | 1072692.491 | 61.5897 | 148.6907 | 500.3408
0.69796 | 1067453.6856 | 79.8168 | 192.6949 | 506.4858
0.7125 | 1095780.5661 | 86.6951 | 209.3004 | 504.8105
0.70914 | 1075001.5095 | 75.7454 | 182.8655 | 468.5422
0.67648 | 996000.4746 | 73.5669 | 177.6063 | 484.4664
0.70252 | 1064149.6377 | 60.3759 | 145.7602 | 508.8254
0.67123 | 1003708.825 | 81.2261 | 196.0971 | 475.5229
0.67438 | 994906.5749 | 58.8024 | 141.9616 | 517.3624
0.67332 | 992235.8974 | 68.4274 | 165.1983 | 499.2069
0.70236 | 1065250.3786 | 85.4957 | 206.405 | 477.3509
0.69932 | 1063527.366 | 87.2466 | 210.6319 | 493.7319
0.7009 | 1070698.8171 | 93.4863 | 225.696 | 483.0789
0.7265 | 1136386.5878 | 76.3596 | 184.3484 | 488.4171
0.70706 | 1077396.6681 | 66.9274 | 161.577 | 503.0649
0.69614 | 1065610.9306 | 74.7262 | 180.405 | 503.1328
0.68319 | 1022411.5182 | 61.7424 | 149.0593 | 516.9773
0.69759 | 1067702.5804 | 83.14 | 200.7178 | 489.258
0.6652 | 979330.0725 | 73.5616 | 177.5935 | 486.5418
0.71447 | 1095895.1903 | 73.0734 | 176.4149 | 493.4147
0.68321 | 1019843.2684 | 61.586 | 148.6818 | 512.1249
0.70442 | 1068450.7827 | 72.26 | 174.4511 | 471.5939
0.69277 | 1033242.6808 | 81.0158 | 195.5894 | 482.5084
0.68998 | 1041770.1829 | 77.1373 | 186.226 | 483.1621
0.69901 | 1053485.9919 | 62.1794 | 150.1143 | 470.0446
0.67934 | 1002029.802 | 62.851 | 151.7358 | 465.3754
0.68531 | 1036149.4686 | 72.0302 | 173.8962 | 493.974
0.68409 | 1037889.9698 | 62.5949 | 151.1174 | 495.1633
0.70806 | 1078733.5535 | 85.8808 | 207.3345 | 464.6991
0.68697 | 1065138.5029 | 79.2136 | 191.2385 | 482.072
0.68076 | 1024456.8422 | 62.0867 | 149.8905 | 470.0932
0.68046 | 1017918.2634 | 79.6799 | 192.3642 | 523.9709
0.68039 | 999668.5206 | 68.2233 | 164.7057 | 475.4964
0.6847 | 1031766.6354 | 68.1052 | 164.4204 | 489.7843
0.67884 | 1002455.9596 | 73.7856 | 178.1342 | 506.2234
0.70795 | 1076200.7362 | 56.0958 | 135.4273 | 513.4
0.681 | 1023559.8472 | 71.6049 | 172.8694 | 501.2497
0.69769 | 1079142.802 | 84.0181 | 202.8377 | 480.4219
0.68541 | 1030673.4622 | 84.7662 | 204.6436 | 492.1999
0.69004 | 1050009.6104 | 70.6575 | 170.5823 | 457.1031
0.67909 | 1031553.3019 | 77.6059 | 187.3572 | 486.229
0.71137 | 1079589.4805 | 66.6573 | 160.925 | 480.7951
0.69859 | 1067994.1934 | 76.3879 | 184.4166 | 471.2406
0.69217 | 1058705.6949 | 69.129 | 166.8923 | 483.9464
0.69221 | 1056987.5121 | 76.2967 | 184.1965 | 487.2536
0.69808 | 1072909.2591 | 67.5054 | 162.9724 | 481.1039
0.67243 | 992142.4763 | 83.173 | 200.7973 | 475.1764
0.6822 | 1024014.0444 | 83.0901 | 200.5972 | 510.1158
0.67901 | 1025628.6416 | 70.1375 | 169.327 | 500.3335
0.67487 | 1008100.4498 | 77.957 | 188.2049 | 505.1348
0.68407 | 1014234.2537 | 61.1704 | 147.6784 | 500.2429
0.69342 | 1038959.3321 | 78.195 | 188.7794 | 475.725
0.68687 | 1043816.1748 | 68.6819 | 165.8128 | 497.6938
0.69691 | 1034216.5267 | 63.2345 | 152.6616 | 506.3875
0.70286 | 1079141.8251 | 78.4403 | 189.3716 | 493.024
0.6951 | 1061399.7152 | 72.746 | 175.6245 | 476.5551
0.70572 | 1084449.9045 | 75.0154 | 181.1032 | 501.3136
0.68155 | 1027312.355 | 73.5834 | 177.6461 | 501.8346
0.69675 | 1065541.3728 | 60.9555 | 147.1597 | 513.3853
0.67523 | 983448.8207 | 58.0522 | 140.1505 | 500.5041
0.6941 | 1033942.8535 | 74.4304 | 179.6908 | 482.8965
0.70409 | 1068828.0349 | 75.85 | 183.1181 | 498.6579
0.68962 | 1043622.817 | 60.319 | 145.623 | 496.7427
0.67649 | 1009797.974 | 69.312 | 167.334 | 491.6524
0.68339 | 1014347.095 | 74.1643 | 179.0484 | 486.0139
0.68425 | 995944.9287 | 74.3622 | 179.5262 | 495.7065
0.6651 | 991528.9308 | 79.4074 | 191.7064 | 515.5807
0.67322 | 1019966.5423 | 72.1781 | 174.2533 | 492.6755
0.71314 | 1107505.9655 | 75.8127 | 183.028 | 522.399
0.68718 | 1036234.0091 | 75.3467 | 181.9031 | 492.3227
0.70506 | 1082894.1856 | 71.3178 | 172.1765 | 499.6846
0.68683 | 1026499.0197 | 59.5702 | 143.8152 | 515.1916
0.69441 | 1039388.2304 | 50.4211 | 121.7274 | 456.3636
0.69258 | 1056739.7188 | 81.2712 | 196.2061 | 511.3175
0.67407 | 1008471.1628 | 65.7113 | 158.6411 | 474.2907
0.67379 | 1009669.8127 | 74.8741 | 180.7621 | 481.1554
0.69783 | 1063976.4641 | 76.1831 | 183.9224 | 496.7252
0.67853 | 999613.7839 | 70.1849 | 169.4413 | 507.3221
0.71354 | 1086110.1083 | 81.0769 | 195.7371 | 508.7299
0.7182 | 1101878.4841 | 78.382 | 189.231 | 493.9078
0.69285 | 1062805.9419 | 56.605 | 136.6565 | 484.9477
0.70019 | 1051839.5567 | 78.3989 | 189.2716 | 485.3657
0.68796 | 1034054.8265 | 73.3763 | 177.1461 | 505.6218
0.70244 | 1066521.5393 | 70.5019 | 170.2066 | 468.1291
0.67888 | 1000201.7318 | 76.8897 | 185.6282 | 478.5408
0.71206 | 1101533.0202 | 79.0338 | 190.8044 | 515.7709
0.69398 | 1066078.0773 | 71.3309 | 172.2081 | 493.9531
0.69219 | 1045487.1798 | 89.5889 | 216.2867 | 465.9888
0.68348 | 1009035.0243 | 75.4247 | 182.0914 | 511.8616
0.68607 | 1027667.0362 | 74.8451 | 180.6921 | 498.9319
0.67733 | 1003429.991 | 59.048 | 142.5544 | 476.5508
0.68828 | 1033255.0959 | 67.8478 | 163.7991 | 511.1699
0.6847 | 1025354.6396 | 71.3429 | 172.2369 | 509.2253
0.68122 | 1035628.6672 | 75.9615 | 183.3872 | 463.2346
0.69262 | 1051421.4034 | 74.1066 | 178.9091 | 496.7455
0.65749 | 964252.5644 | 76.5372 | 184.7772 | 493.6046
0.72656 | 1127239.5046 | 70.1194 | 169.2833 | 502.4127
0.70412 | 1085758.3849 | 78.7799 | 190.1914 | 521.7898
0.68808 | 1046339.0602 | 80.9981 | 195.5468 | 506.8128
0.67958 | 1020669.1532 | 78.3827 | 189.2325 | 458.8217
0.7056 | 1085364.311 | 80.2772 | 193.8064 | 490.9205
0.67488 | 990087.6643 | 81.2088 | 196.0555 | 504.0586
0.69319 | 1052399.1884 | 62.6367 | 151.2183 | 491.5022
0.68437 | 1040559.9567 | 77.6795 | 187.5349 | 495.1321
0.67852 | 1016162.2262 | 87.3846 | 210.9651 | 498.1764
0.67632 | 1014010.6334 | 74.9152 | 180.8612 | 476.0315
0.69017 | 1037993.7009 | 66.0576 | 159.4773 | 492.576
0.6781 | 1004425.6093 | 94.681 | 228.5802 | 508.8481
0.69401 | 1039915.0339 | 65.8133 | 158.8874 | 474.7456
0.67621 | 1018111.1843 | 69.9125 | 168.7836 | 477.6833
0.70338 | 1073179.4676 | 67.0141 | 161.7863 | 482.7336
0.69345 | 1060050.0211 | 72.8267 | 175.8192 | 491.7417
0.67154 | 990756.0024 | 70.2864 | 169.6864 | 509.7526
0.73087 | 1157565.539 | 77.3176 | 186.6611 | 490.8058
0.68898 | 1038879.3183 | 84.9986 | 205.2047 | 476.2497
0.67742 | 1006257.5237 | 72.5148 | 175.0662 | 489.6105
0.69962 | 1061693.5622 | 74.0819 | 178.8496 | 494.308
0.68051 | 1025563.5377 | 83.422 | 201.3985 | 484.8717
0.66621 | 985705.6681 | 78.8085 | 190.2606 | 494.8019
0.68928 | 1047673.6264 | 68.695 | 165.8444 | 472.69
0.6994 | 1078614.4722 | 56.9092 | 137.3909 | 460.4831
0.68744 | 1032384.1607 | 75.0385 | 181.159 | 500.5355
0.70161 | 1076407.4817 | 83.337 | 201.1934 | 472.8323
0.69394 | 1049953.5104 | 73.8618 | 178.3182 | 480.3411
0.71475 | 1099305.7591 | 67.5533 | 163.0881 | 488.2965
0.69655 | 1061499.6718 | 79.5323 | 192.0079 | 522.788
0.701 | 1042954.3253 | 76.0111 | 183.507 | 492.9511
0.69969 | 1051204.4882 | 75.0684 | 181.2312 | 502.7688
0.7016 | 1080115.4103 | 83.6255 | 201.8899 | 481.1513
0.71289 | 1101929.3836 | 82.3323 | 198.7677 | 482.0224
0.69885 | 1056418.2321 | 77.229 | 186.4474 | 518.8762
0.70265 | 1076006.6893 | 78.2361 | 188.8787 | 476.3971
0.68855 | 1021012.8707 | 75.7574 | 182.8946 | 471.4062
0.67712 | 997945.1486 | 71.9475 | 173.6967 | 488.9693
0.68763 | 1023859.9279 | 70.9352 | 171.2528 | 497.179
0.69982 | 1053215.6954 | 76.889 | 185.6264 | 513.5863
0.69774 | 1058475.161 | 75.7989 | 182.9948 | 491.5852
0.70509 | 1082162.423 | 74.9245 | 180.8837 | 509.6477
0.67607 | 1031673.3308 | 65.0226 | 156.9783 | 490.946
0.67367 | 1011248.6924 | 76.6909 | 185.1482 | 500.9194
0.72062 | 1098822.5764 | 77.9029 | 188.0742 | 493.4879
0.68889 | 1055379.5959 | 71.1376 | 171.7414 | 476.62
0.65787 | 955655.3448 | 67.4525 | 162.8448 | 505.0189
0.69644 | 1062458.499 | 68.9482 | 166.4557 | 480.2028
0.69299 | 1046811.733 | 86.6595 | 209.2145 | 506.0463
0.67346 | 1025144.3096 | 82.6858 | 199.6211 | 501.0793
0.66668 | 1000240.2879 | 74.9661 | 180.9841 | 504.9967
0.71287 | 1109810.5804 | 74.1678 | 179.0569 | 503.2101
0.69837 | 1078517.1898 | 80.5966 | 194.5775 | 474.9236
0.6694 | 995380.0137 | 72.59 | 175.2477 | 467.0481
0.70158 | 1065388.1519 | 88.607 | 213.9163 | 457.1657
0.69325 | 1056493.2335 | 60.7432 | 146.6471 | 488.8208
0.68019 | 1029869.185 | 58.1535 | 140.395 | 481.0841
0.66991 | 983436.0477 | 73.3514 | 177.0859 | 482.46
0.68186 | 1024998.7553 | 66.0258 | 159.4003 | 500.6849
0.68843 | 1046069.675 | 75.9232 | 183.2949 | 518.0859
0.68452 | 1021856.7553 | 77.9073 | 188.085 | 470.5031
0.69272 | 1044422.4994 | 64.2447 | 155.1005 | 491.7325
0.70023 | 1065876.4261 | 72.518 | 175.074 | 495.3397
0.70674 | 1087972.2039 | 78.6483 | 189.8738 | 492.8318
0.67027 | 992469.8639 | 78.8689 | 190.4064 | 486.3283
0.67622 | 1008123.2749 | 72.7773 | 175.6999 | 486.4394
0.69845 | 1082015.665 | 68.0426 | 164.2693 | 520.3795
0.7244 | 1128661.6084 | 77.851 | 187.949 | 487.9071
0.69608 | 1071712.0777 | 73.3498 | 177.0821 | 466.8395
0.67964 | 1039282.557 | 79.8047 | 192.6656 | 471.316
0.69138 | 1037139.2883 | 73.231 | 176.7953 | 490.7047
0.71085 | 1104911.6032 | 70.2499 | 169.5984 | 506.2986
0.7027 | 1068851.1697 | 79.3815 | 191.6439 | 471.357
0.66993 | 1000815.6482 | 72.9288 | 176.0656 | 466.7897
0.6996 | 1060701.5947 | 66.5922 | 160.7679 | 491.5889
0.72071 | 1098083.3887 | 76.129 | 183.7916 | 483.6305
0.67675 | 1016123.4214 | 86.1372 | 207.9537 | 488.3381
0.68425 | 1028379.5695 | 76.7833 | 185.3714 | 482.1063
0.69552 | 1069168.6772 | 69.4758 | 167.7295 | 501.5349
0.71648 | 1103393.633 | 78.6782 | 189.9461 | 506.0727
0.68751 | 1034723.0493 | 68.0211 | 164.2174 | 462.0989
0.69123 | 1039081.7255 | 73.3837 | 177.1639 | 507.3591
0.68561 | 1018543.407 | 72.523 | 175.086 | 508.4474
0.69055 | 1046487.5693 | 68.4915 | 165.3531 | 490.5234
0.69439 | 1052043.2714 | 70.3503 | 169.8407 | 537.4925
0.69167 | 1028377.0202 | 83.8114 | 202.3386 | 499.8551
0.71304 | 1100692.5382 | 69.7662 | 168.4306 | 485.8296
0.69985 | 1060412.8906 | 80.7063 | 194.8422 | 484.5929
0.68459 | 1018145.712 | 75.9819 | 183.4366 | 470.6729
0.71943 | 1132869.5114 | 76.8714 | 185.584 | 474.2738
0.66978 | 1001783.0769 | 63.7316 | 153.8618 | 470.407
0.7058 | 1080086.7357 | 67.2138 | 162.2685 | 525.8252
0.70787 | 1085510.4729 | 67.3544 | 162.6079 | 500.6328
0.69767 | 1049805.2203 | 75.1852 | 181.5132 | 465.5846
0.68347 | 1022223.5288 | 87.2697 | 210.6877 | 463.8324
0.66587 | 981978.0312 | 80.1074 | 193.3964 | 516.6542
0.73164 | 1128389.8037 | 59.2362 | 143.0088 | 506.12
0.66205 | 968683.6551 | 66.1281 | 159.6473 | 489.847
0.68637 | 1031992.5891 | 64.0667 | 154.6707 | 511.2487
0.67546 | 987321.4432 | 85.2898 | 205.9077 | 498.1663
0.668 | 997922.5423 | 67.5179 | 163.0027 | 505.1697
0.6885 | 1041168.5459 | 66.4199 | 160.3518 | 491.4885
0.69972 | 1049042.7938 | 74.9112 | 180.8515 | 497.1899
0.681 | 1020841.5967 | 73.1782 | 176.6677 | 475.7761
0.68468 | 1032236.9433 | 66.2179 | 159.8642 | 490.4841
0.69401 | 1065578.8722 | 78.4419 | 189.3754 | 493.7538
0.68812 | 1041426.7725 | 62.4408 | 150.7454 | 490.9298
0.6906 | 1052598.0848 | 56.2439 | 135.7849 | 478.49
0.68853 | 1036483.3983 | 62.1067 | 149.9389 | 493.6619
0.67782 | 994732.2955 | 68.1447 | 164.5159 | 510.4156
0.6954 | 1048601.159 | 87.9861 | 212.4173 | 506.7699
0.68367 | 1028126.6881 | 73.4136 | 177.2362 | 510.7177
0.69192 | 1038927.1302 | 78.362 | 189.1826 | 502.9875
0.68918 | 1041996.8673 | 70.2061 | 169.4924 | 496.544
0.68913 | 1049433.9101 | 79.474 | 191.8673 | 490.4061
0.6812 | 1019185.1343 | 87.3006 | 210.7623 | 472.3557
0.72395 | 1125135.297 | 72.1664 | 174.2251 | 484.5489
0.68177 | 1017821.9536 | 96.7427 | 233.5576 | 499.9433
0.68615 | 1029209.4635 | 74.9483 | 180.9412 | 504.6572
0.68261 | 1035774.3857 | 79.1895 | 191.1803 | 498.6775
0.68243 | 1009958.8938 | 66.8872 | 161.4799 | 474.8253
0.71202 | 1101180.6822 | 79.0131 | 190.7545 | 472.2206
0.67819 | 996206.4092 | 71.709 | 173.1208 | 485.9758
0.68603 | 1053801.0659 | 88.1226 | 212.7469 | 469.5484
0.68643 | 1056282.6276 | 85.1807 | 205.6444 | 475.0012
0.70319 | 1067857.4898 | 72.3463 | 174.6594 | 498.3184
0.70328 | 1061888.8905 | 86.6211 | 209.1218 | 460.5984
0.71442 | 1100624.4481 | 85.2783 | 205.8801 | 504.5634
0.68973 | 1033460.7312 | 74.867 | 180.7448 | 454.8773
0.69235 | 1060875.0996 | 63.218 | 152.6216 | 486.7653
0.67702 | 1006245.2215 | 63.3317 | 152.8963 | 485.8828
0.68833 | 1021300.8854 | 61.215 | 147.786 | 496.4866
0.68616 | 1030783.1255 | 61.9784 | 149.6291 | 493.7432
0.6949 | 1053589.6709 | 89.9308 | 217.1121 | 501.9908
0.69525 | 1041495.0681 | 65.7959 | 158.8453 | 512.8701
0.66411 | 985480.9223 | 75.8126 | 183.0278 | 450.671
0.70018 | 1049456.278 | 74.8917 | 180.8044 | 527.9907
0.71904 | 1135476.2495 | 78.9044 | 190.4921 | 498.3671
0.69003 | 1048380.9761 | 66.7317 | 161.1046 | 521.7626
0.704 | 1073252.1589 | 76.1733 | 183.8986 | 470.8502
0.69266 | 1049958.8108 | 75.7411 | 182.8551 | 472.0957
0.69108 | 1034853.3732 | 74.8219 | 180.636 | 488.2309
0.70313 | 1073266.3205 | 67.0431 | 161.8564 | 471.0658
0.71709 | 1101335.5677 | 80.1535 | 193.5076 | 493.5576
0.69079 | 1046756.1067 | 62.5178 | 150.9312 | 470.804
0.67036 | 987885.588 | 73.6335 | 177.7671 | 502.027
0.69449 | 1034330.5687 | 70.9527 | 171.295 | 497.4133
0.68762 | 1026396.029 | 67.4623 | 162.8684 | 503.0386
0.69272 | 1052396.7433 | 80.9206 | 195.3595 | 467.3613
0.6866 | 1018744.7594 | 65.1697 | 157.3336 | 467.0836
0.67874 | 1012552.1578 | 84.0941 | 203.0211 | 492.4438
0.66829 | 1000500.6533 | 82.8719 | 200.0704 | 505.1145
0.67155 | 982198.1454 | 67.914 | 163.9589 | 529.0881
0.68855 | 1022396.6524 | 68.948 | 166.4552 | 472.1742
0.66787 | 999304.475 | 63.7559 | 153.9203 | 488.8944
0.67739 | 1005623.3323 | 82.1775 | 198.3939 | 493.8077
0.68372 | 1020083.3123 | 74.5478 | 179.9743 | 501.2459
0.68981 | 1039615.0952 | 57.885 | 139.7467 | 479.8782
0.7206 | 1110448.1956 | 75.2278 | 181.616 | 507.2969
0.65898 | 963634.2784 | 60.3314 | 145.6528 | 471.959
0.68754 | 1034718.6281 | 75.1975 | 181.5429 | 495.6529
0.6954 | 1054887.2522 | 74.0532 | 178.7801 | 505.4791
0.70194 | 1048792.8714 | 61.6442 | 148.8222 | 440.1951
0.67442 | 1017191.5057 | 80.3369 | 193.9503 | 501.8971
0.70739 | 1080777.8218 | 69.0519 | 166.706 | 476.2103
0.67121 | 987288.307 | 71.8057 | 173.3542 | 489.0486
0.67242 | 1037786.5943 | 77.1158 | 186.174 | 488.0242
0.68078 | 1018615.2333 | 85.8376 | 207.2302 | 478.2579
0.6754 | 1011857.6049 | 73.1357 | 176.5652 | 480.6612
0.70401 | 1059250.6926 | 83.2864 | 201.0711 | 516.7202
0.67944 | 1021616.4913 | 69.8189 | 168.5578 | 498.46
0.68929 | 1055032.2691 | 83.1422 | 200.7231 | 503.4642
0.66775 | 992872.9806 | 67.2007 | 162.2369 | 478.9933
0.69957 | 1044765.0698 | 76.299 | 184.202 | 493.9519
0.69737 | 1041757.2417 | 61.6316 | 148.7919 | 474.3435
0.72157 | 1127941.014 | 94.2027 | 227.4255 | 488.861
0.67893 | 995350.6076 | 84.1062 | 203.0504 | 493.2621
0.69339 | 1049301.0809 | 72.0943 | 174.0512 | 503.4762
0.69774 | 1047073.5752 | 73.2582 | 176.8611 | 498.2735
0.6751 | 991505.6283 | 73.3845 | 177.1659 | 475.7193
0.72154 | 1107082.5869 | 69.3454 | 167.4146 | 510.4529
0.6838 | 1022957.0122 | 76.4294 | 184.5169 | 464.7769
0.69909 | 1069011.7869 | 78.1234 | 188.6065 | 478.3883
0.69157 | 1022965.4165 | 72.2101 | 174.3306 | 505.8714
0.67928 | 1025339.9394 | 64.318 | 155.2774 | 468.1951
0.69213 | 1044049.6574 | 85.0328 | 205.2874 | 467.9131
0.66571 | 974442.0074 | 62.3716 | 150.5785 | 495.5008
0.69113 | 1046991.3413 | 75.6934 | 182.7399 | 469.5551
0.68983 | 1037946.0642 | 62.0772 | 149.8676 | 488.3348
0.68412 | 1041204.907 | 64.8512 | 156.5647 | 508.1311
0.67876 | 1012269.0129 | 55.0326 | 132.8604 | 474.7657
0.70401 | 1069087.093 | 70.5136 | 170.235 | 501.7019
0.681 | 1003690.0241 | 73.643 | 177.79 | 484.7142
0.65707 | 959484.2185 | 71.2856 | 172.0987 | 504.4126
0.70875 | 1090183.7095 | 78.5847 | 189.7203 | 498.7954
0.70222 | 1067630.9712 | 77.6749 | 187.5239 | 480.9902
0.71186 | 1093915.6936 | 65.4792 | 158.0807 | 522.4847
0.7218 | 1115672.0185 | 61.6081 | 148.7352 | 474.3787
0.70283 | 1068368.5892 | 75.9531 | 183.3671 | 466.3389
0.70041 | 1067503.7644 | 77.8663 | 187.986 | 480.5141
0.67646 | 1010613.603 | 72.9316 | 176.0725 | 487.0181
0.6732 | 994638.5197 | 66.7758 | 161.211 | 483.0249
0.68421 | 1029949.8295 | 80.1798 | 193.5712 | 491.9746
0.67585 | 999021.2002 | 60.9854 | 147.2318 | 500.2443
0.68118 | 1010971.8439 | 74.1568 | 179.0303 | 460.5857
0.69877 | 1059533.0825 | 64.6095 | 155.9812 | 504.9138
0.68647 | 1021874.1542 | 76.0511 | 183.6037 | 479.6662
0.70496 | 1072326.7159 | 72.4673 | 174.9516 | 504.5948
0.6917 | 1036260.63 | 78.1856 | 188.7568 | 500.5902
0.68707 | 1030234.0156 | 59.7661 | 144.2881 | 483.0028
0.68044 | 1017722.7549 | 73.0148 | 176.2734 | 485.1942
0.66537 | 981091.3765 | 62.8291 | 151.6829 | 474.5226
0.70759 | 1083280.4056 | 78.2152 | 188.8282 | 520.4693
0.692 | 1037282.5872 | 63.8714 | 154.1993 | 506.9907
0.66972 | 987834.2758 | 60.8544 | 146.9154 | 477.6267
0.69829 | 1063605.2153 | 77.8166 | 187.866 | 506.7468
0.69977 | 1075064.6597 | 61.2349 | 147.8342 | 485.9052
0.67878 | 1014219.8361 | 65.6277 | 158.4394 | 497.7905
0.68292 | 1018592.7775 | 79.4459 | 191.7994 | 505.2669
0.712 | 1093386.647 | 81.7962 | 197.4734 | 503.9783
0.70258 | 1052766.2584 | 81.3944 | 196.5035 | 503.0664
0.70821 | 1074767.447 | 70.4673 | 170.1231 | 468.2067
0.70082 | 1081384.7833 | 75.3269 | 181.8552 | 507.8332
0.67996 | 1024991.3836 | 78.6703 | 189.9268 | 504.951
0.67967 | 1025368.5941 | 64.2837 | 155.1947 | 500.4673
0.68514 | 1043252.5363 | 79.4669 | 191.8501 | 490.9903
0.70736 | 1084138.9547 | 63.2335 | 152.6592 | 490.6458
0.70067 | 1080091.7444 | 85.3304 | 206.0058 | 513.2609
0.68062 | 1028730.7777 | 75.5224 | 182.3273 | 526.6273
0.71137 | 1089004.9935 | 74.7321 | 180.4193 | 494.3956
0.71407 | 1105634.0785 | 75.5864 | 182.4818 | 517.3446
0.69362 | 1057372.5458 | 66.9561 | 161.6462 | 493.8934
0.6953 | 1033820.4958 | 69.0418 | 166.6816 | 514.2319
0.71309 | 1100297.5974 | 80.5262 | 194.4075 | 465.99
0.68945 | 1018623.7472 | 72.1244 | 174.1236 | 510.6667
0.69378 | 1040661.5968 | 73.2257 | 176.7826 | 506.9108
0.70731 | 1078178.9402 | 66.9521 | 161.6366 | 504.8948
0.69 | 1034228.7993 | 76.5768 | 184.8727 | 487.3162
0.71772 | 1112508.8704 | 73.8462 | 178.2804 | 457.2857
0.69732 | 1050259.2722 | 66.153 | 159.7074 | 483.8573
0.69504 | 1050262.7806 | 86.6885 | 209.2846 | 500.1356
0.68983 | 1038162.6073 | 73.4682 | 177.3679 | 525.2179
0.68758 | 1034437.2108 | 73.4442 | 177.31 | 494.4117
0.67404 | 1015977.3438 | 89.4047 | 215.8419 | 504.5628
0.6917 | 1039278.6912 | 69.767 | 168.4325 | 523.5723
0.70825 | 1093565.2223 | 87.0387 | 210.1299 | 483.2112
0.6958 | 1078570.3341 | 79.1947 | 191.193 | 490.6323
0.65256 | 947008.1395 | 72.5116 | 175.0586 | 488.2163
0.71151 | 1102008.6756 | 70.0007 | 168.9966 | 474.0165
0.66458 | 966323.904 | 57.9554 | 139.9167 | 502.5866
0.66874 | 971592.4119 | 69.2152 | 167.1004 | 503.5373
0.70947 | 1093300.6976 | 54.4443 | 131.4402 | 506.0312
0.67587 | 1005197.5509 | 59.2795 | 143.1135 | 521.4293
0.68044 | 1027499.3702 | 75.4521 | 182.1574 | 496.4856
0.67462 | 996985.1117 | 77.127 | 186.2009 | 483.7992
0.69891 | 1073749.2711 | 67.478 | 162.9063 | 509.7031
0.6818 | 1026037.9198 | 78.5384 | 189.6084 | 502.4057
0.69723 | 1065385.8237 | 87.6603 | 211.6307 | 497.3969
0.7045 | 1079991.0505 | 80.0046 | 193.1483 | 466.4269
0.68785 | 1036729.2937 | 80.2802 | 193.8134 | 493.0635
0.68174 | 1010828.378 | 75.751 | 182.8792 | 484.2604
0.6803 | 1016966.6895 | 65.6584 | 158.5134 | 507.94
0.70467 | 1064946.5659 | 70.1749 | 169.4171 | 470.6176
0.66405 | 976141.1017 | 84.3307 | 203.5923 | 487.0665
0.69796 | 1049882.2333 | 76.8495 | 185.5312 | 525.3548
0.67194 | 1005416.3469 | 79.6995 | 192.4116 | 489.0453
0.67858 | 1011933.8566 | 87.0144 | 210.0713 | 490.7761
0.68161 | 1012333.1356 | 67.1248 | 162.0537 | 464.3215
0.68686 | 1032608.1942 | 69.8075 | 168.5303 | 512.9792
0.6911 | 1026464.7969 | 75.8811 | 183.1932 | 490.473
0.65976 | 980808.2763 | 54.8577 | 132.4382 | 457.679
0.70434 | 1082885.04 | 76.82 | 185.46 | 508.7358
0.71059 | 1093903.7404 | 81.9973 | 197.959 | 459.5255
0.72196 | 1119260.3538 | 77.4625 | 187.011 | 463.4773
0.69333 | 1060535.6239 | 69.1328 | 166.9013 | 490.3444
0.66226 | 987642.9127 | 77.8294 | 187.8967 | 480.1046
0.67353 | 1000788.2964 | 63.3143 | 152.8544 | 511.8695
0.68328 | 1027798.4151 | 83.2655 | 201.0206 | 482.1189
0.70315 | 1085498.0011 | 63.6917 | 153.7654 | 504.3568
0.67814 | 1018300.2084 | 81.5668 | 196.9196 | 491.5272
0.72103 | 1124540.9036 | 76.1366 | 183.81 | 469.5983
0.67588 | 1005697.7685 | 70.8881 | 171.139 | 494.9999
0.70064 | 1054477.2228 | 73.4704 | 177.3733 | 487.6957
0.70963 | 1092150.249 | 67.5982 | 163.1965 | 464.3556
0.69373 | 1047938.9908 | 71.9649 | 173.7385 | 498.5906
0.67214 | 985157.9609 | 79.6167 | 192.2117 | 512.1004
0.68983 | 1056870.1037 | 81.606 | 197.0142 | 502.2857
0.69393 | 1051478.6684 | 66.5528 | 160.6726 | 532.5453
0.69021 | 1048284.6915 | 77.5937 | 187.3278 | 497.0397
0.71965 | 1114219.4114 | 84.102 | 203.0401 | 491.064
0.69179 | 1035894.4109 | 79.7595 | 192.5566 | 472.0382
0.68278 | 1030360.4902 | 71.8931 | 173.5654 | 502.8444
0.67854 | 1029110.7604 | 60.4514 | 145.9426 | 483.6129
0.69442 | 1056893.3017 | 85.7953 | 207.1281 | 501.0687
0.68378 | 1023218.4047 | 71.2898 | 172.1088 | 499.7155
0.67331 | 1002130.0423 | 69.1114 | 166.8497 | 486.4619
0.70457 | 1085204.3813 | 80.3818 | 194.0589 | 503.4179
0.69767 | 1057115.1931 | 68.4806 | 165.3267 | 511.4043
0.69116 | 1050958.4599 | 80.9928 | 195.534 | 516.2611
0.67528 | 1002866.0303 | 61.2721 | 147.924 | 475.5474
0.69857 | 1042708.2969 | 77.8981 | 188.0626 | 503.9277
0.70243 | 1062998.9184 | 88.3012 | 213.178 | 501.0167
0.68571 | 1029230.68 | 93.238 | 225.0965 | 500.0363
0.70212 | 1073020.794 | 86.1486 | 207.9811 | 523.0497
0.69786 | 1036238.148 | 80.6582 | 194.7262 | 517.0522
0.67168 | 1004764.6299 | 67.6082 | 163.2207 | 494.9471
0.69208 | 1049594.2844 | 88.491 | 213.6363 | 484.6557
0.69885 | 1047652.1096 | 84.506 | 204.0156 | 489.6933
0.65817 | 969094.4692 | 71.8793 | 173.5319 | 501.625
0.67751 | 1007235.8231 | 70.8255 | 170.9879 | 496.8012
0.67796 | 998319.7438 | 60.7527 | 146.67 | 489.2158
0.6911 | 1050856.8621 | 75.9763 | 183.4229 | 507.5191
0.69598 | 1064865.2043 | 78.1683 | 188.7151 | 528.5035
0.68493 | 1030286.5243 | 71.6618 | 173.0069 | 469.751
0.69873 | 1066887.3236 | 72.7424 | 175.6157 | 501.0286
0.6628 | 975285.9854 | 73.5234 | 177.5011 | 516.9671
0.68008 | 1022820.1511 | 57.7652 | 139.4574 | 491.3854
0.7063 | 1080800.8508 | 87.7873 | 211.9372 | 482.3681
0.69803 | 1040624.7179 | 67.3777 | 162.6642 | 479.1086
0.68139 | 1010782.7265 | 63.2091 | 152.6003 | 507.2662
0.68242 | 1014937.9985 | 50.1 | 120.9522 | 490.1833
0.66255 | 970620.108 | 79.5768 | 192.1153 | 504.0396
0.66389 | 974828.8457 | 79.7606 | 192.5591 | 477.3171
0.68611 | 1046207.5211 | 62.9107 | 151.88 | 487.8744
0.70254 | 1073066.0466 | 65.0837 | 157.126 | 475.8271
0.68113 | 1022634.9542 | 71.5265 | 172.6802 | 473.4862
0.68941 | 1038587.4478 | 66.1681 | 159.7439 | 484.5072
0.67754 | 1011571.9934 | 77.918 | 188.1107 | 461.619
0.6753 | 1008030.5631 | 59.4968 | 143.638 | 509.0419
0.67759 | 1020657.3094 | 73.5332 | 177.525 | 477.5585
0.69098 | 1054087.6371 | 70.8356 | 171.0122 | 460.6065
0.69589 | 1042253.3095 | 71.3526 | 172.2605 | 496.7568
0.6964 | 1050054.0643 | 71.8338 | 173.4222 | 485.5018
0.68532 | 1026563.8063 | 63.8278 | 154.094 | 500.7795
0.68513 | 1038441.6404 | 66.5718 | 160.7186 | 502.7004
0.66851 | 990335.0971 | 59.0466 | 142.5511 | 486.7338
0.67523 | 1002597.9684 | 63.2689 | 152.7445 | 486.4744
0.71535 | 1093324.5095 | 90.6544 | 218.8592 | 481.986
0.69025 | 1049242.3583 | 68.7014 | 165.8599 | 504.6316
0.68912 | 1050144.8661 | 70.7612 | 170.8325 | 485.7987
0.66703 | 976056.977 | 62.3511 | 150.5288 | 502.0528
0.6651 | 964715.3217 | 75.6983 | 182.7518 | 504.6153
0.66449 | 977535.8489 | 82.9273 | 200.2042 | 481.8862
0.69303 | 1048489.3192 | 80.5847 | 194.5486 | 499.1998
0.69675 | 1072634.6024 | 76.9226 | 185.7076 | 482.5433
0.70109 | 1063356.3211 | 80.7879 | 195.0393 | 509.0203
0.71139 | 1087777.0384 | 72.8464 | 175.8667 | 502.5573
0.70178 | 1063123.1379 | 82.2884 | 198.6619 | 491.6933
0.69535 | 1043645.709 | 79.9448 | 193.0039 | 495.7436
0.66824 | 985644.5738 | 47.1361 | 113.7965 | 500.2789
0.68215 | 1029496.1444 | 73.675 | 177.8671 | 489.6725
0.65959 | 974009.8979 | 66.7399 | 161.1243 | 483.4263
0.69625 | 1046287.8719 | 61.9997 | 149.6806 | 503.0004
0.68209 | 1031247.1622 | 82.1258 | 198.2693 | 505.1076
0.68758 | 1035542.7298 | 53.2094 | 128.4589 | 484.0767
0.69201 | 1028172.9748 | 77.5404 | 187.199 | 465.6628
0.68566 | 1037076.5068 | 91.15 | 220.0556 | 474.7798
0.68724 | 1044799.9532 | 74.2162 | 179.1738 | 469.4688
0.67516 | 999399.3332 | 68.7342 | 165.9391 | 479.9236
0.68348 | 1025096.0124 | 74.9399 | 180.9209 | 512.9367
0.71922 | 1113270.5615 | 82.4479 | 199.0469 | 514.6245
0.70391 | 1077385.3782 | 69.1748 | 167.0026 | 505.0689
0.68088 | 1016174.7414 | 67.4521 | 162.8439 | 488.2413
0.70177 | 1072916.594 | 76.3925 | 184.4278 | 487.0729
0.66226 | 953171.1189 | 81.132 | 195.8699 | 499.1763
0.68802 | 1030845.1674 | 67.7155 | 163.4796 | 460.5598
0.69602 | 1055075.7047 | 70.3882 | 169.9323 | 476.3195
0.69287 | 1053990.1229 | 93.0939 | 224.7487 | 470.4386
0.69925 | 1043290.5833 | 65.5324 | 158.2092 | 495.0177
0.66902 | 999267.3043 | 61.6228 | 148.7706 | 489.5051
0.69813 | 1061907.4904 | 76.0021 | 183.4853 | 489.1873
0.67013 | 1001120.2258 | 55.1279 | 133.0905 | 490.2216
0.67002 | 994820.038 | 65.2134 | 157.4391 | 488.3865
0.6752 | 1017254.4676 | 75.2786 | 181.7387 | 500.908
0.69331 | 1053547.1015 | 77.6344 | 187.4259 | 479.3707
0.72503 | 1113896.2407 | 75 | 181.066 | 512.3386
0.68099 | 1024620.0935 | 69.8652 | 168.6695 | 528.4389
0.69689 | 1052406.8532 | 75.4561 | 182.1672 | 473.4326
0.69419 | 1053340.8756 | 77.5214 | 187.1533 | 488.5129
0.72147 | 1125796.87 | 83.1157 | 200.6591 | 469.7511
0.6886 | 1039827.653 | 80.4329 | 194.1821 | 496.3537
0.69537 | 1041139.0384 | 76.4928 | 184.67 | 482.1346
0.7009 | 1083504.9158 | 85.571 | 206.5867 | 516.3422
0.69649 | 1065163.5345 | 83.0728 | 200.5555 | 459.9615
0.70507 | 1078858.2229 | 65.1998 | 157.4062 | 479.9234
0.69674 | 1055121.7294 | 86.8769 | 209.7394 | 505.2801
0.69045 | 1056020.7322 | 84.6634 | 204.3956 | 486.6609
0.6627 | 975600.8029 | 68.3585 | 165.032 | 520.3567
0.6993 | 1051444.979 | 55.4834 | 133.9488 | 502.2945
0.70539 | 1094288.7825 | 77.5143 | 187.1361 | 486.4224
0.71184 | 1091825.9537 | 73.0715 | 176.4103 | 486.4578
0.69489 | 1028343.1848 | 77.2318 | 186.4541 | 494.1305
0.7113 | 1081470.7287 | 74.2229 | 179.1898 | 478.4211
0.69837 | 1064558.9368 | 62.8903 | 151.8305 | 458.0784
0.67736 | 1003356.0446 | 66.9468 | 161.6238 | 510.2036
0.69656 | 1056300.3619 | 82.3729 | 198.8658 | 472.4308
0.6999 | 1069993.5852 | 58.7177 | 141.7571 | 458.7311
0.69111 | 1017947.5802 | 76.8691 | 185.5785 | 496.7588
0.68052 | 1018515.7208 | 67.643 | 163.3046 | 469.7926
0.68286 | 1034298.8744 | 72.1593 | 174.208 | 489.6109
0.71881 | 1101354.5859 | 74.7971 | 180.5763 | 496.8562
0.69728 | 1036883.9202 | 67.2081 | 162.2546 | 484.8924
0.69033 | 1049934.7349 | 77.7026 | 187.5908 | 506.4458
0.68401 | 1026744.8032 | 76.6888 | 185.1431 | 515.2724
0.69421 | 1047408.2433 | 62.2481 | 150.2801 | 470.2439
0.69591 | 1058048.4683 | 70.5973 | 170.437 | 456.2636
0.70753 | 1099964.4377 | 81.1346 | 195.8761 | 517.9372
0.69854 | 1067918.6189 | 79.7847 | 192.6173 | 465.3798
0.69186 | 1034598.2023 | 73.0006 | 176.239 | 479.3295
0.66898 | 998713.7883 | 77.9118 | 188.0958 | 483.3061
0.71225 | 1096564.0727 | 63.5188 | 153.3479 | 511.4717
0.69266 | 1027877.384 | 63.2319 | 152.6554 | 492.5024
0.72245 | 1114955.3125 | 78.8456 | 190.35 | 485.4084
0.71135 | 1084533.3984 | 77.892 | 188.0478 | 457.7017
0.70869 | 1086975.0771 | 81.6969 | 197.2338 | 503.1104
0.69915 | 1074949.3959 | 67.5306 | 163.0332 | 456.9251
0.68502 | 1036055.5535 | 73.8338 | 178.2505 | 458.8052
0.71556 | 1099416.2479 | 78.4958 | 189.5056 | 466.4881
0.68658 | 1036725.2296 | 66.3216 | 160.1144 | 547.295
0.70028 | 1061937.7146 | 74.0291 | 178.722 | 509.1673
0.69608 | 1053352.6027 | 75.6122 | 182.544 | 516.913
0.65933 | 968044.2479 | 64.4875 | 155.6866 | 505.5879
0.69272 | 1035740.7605 | 69.4269 | 167.6115 | 506.9746
0.70783 | 1083897.1627 | 67.2575 | 162.3739 | 491.7613
0.67468 | 1001037.1619 | 82.1306 | 198.2809 | 458.3205
0.69403 | 1043911.4202 | 67.9801 | 164.1185 | 469.4323
0.71324 | 1104267.5557 | 65.9713 | 159.2689 | 465.9579
0.69205 | 1050414.5933 | 58.297 | 140.7414 | 474.3034
0.69029 | 1032496.1526 | 59.6177 | 143.9299 | 473.1936
0.69022 | 1029217.4978 | 65.4944 | 158.1175 | 499.7521
0.68689 | 1032589.1934 | 75.1134 | 181.3398 | 452.6695
0.67092 | 985709.2966 | 54.2152 | 130.887 | 459.0044
0.70972 | 1086707.2661 | 69.2022 | 167.069 | 498.3206
0.68908 | 1012282.7815 | 58.553 | 141.3594 | 489.7624
0.6782 | 1009630.0291 | 70.9223 | 171.2215 | 488.4793
0.70694 | 1080087.7037 | 60.897 | 147.0182 | 480.6838
0.70367 | 1063391.8862 | 75.5093 | 182.2956 | 530.3296
0.71524 | 1112404.3435 | 79.4591 | 191.8312 | 465.5001
0.69872 | 1059076.0697 | 79.1043 | 190.9747 | 521.9423
0.67789 | 1011329.2765 | 65.5261 | 158.1941 | 477.716
0.70993 | 1098153.6312 | 84.0242 | 202.8523 | 480.6529
0.66796 | 990909.7706 | 75.7903 | 182.974 | 495.6631
0.70615 | 1067186.3919 | 60.2368 | 145.4244 | 484.4943
0.70216 | 1072621.8151 | 72.7762 | 175.6972 | 521.1043
0.67218 | 989767.5288 | 76.846 | 185.5226 | 468.4155
0.69597 | 1065945.9486 | 71.118 | 171.6941 | 491.2621
0.691 | 1060065.1506 | 71.2344 | 171.975 | 507.2876
0.68854 | 1048518.4901 | 76.0891 | 183.6953 | 452.4058
0.68246 | 1028803.7023 | 60.0864 | 145.0614 | 498.9771
0.69254 | 1046258.9828 | 88.1521 | 212.818 | 495.8116
0.6817 | 1027783.1298 | 66.2199 | 159.8689 | 514.3209
0.68736 | 1017676.1161 | 56.8658 | 137.2863 | 481.3533
0.68799 | 1029698.7608 | 50.417 | 121.7173 | 502.133
0.69612 | 1060793.6554 | 72.5982 | 175.2675 | 510.2547
0.66706 | 982032.5605 | 81.4919 | 196.7388 | 486.4568
0.66569 | 970935.8224 | 77.6883 | 187.5562 | 490.4101
0.67158 | 999169.4102 | 82.3564 | 198.826 | 467.6943
0.68299 | 1036563.2475 | 74.0894 | 178.8676 | 466.3032
0.69508 | 1066170.4169 | 73.0755 | 176.4199 | 502.9076
0.69447 | 1046000.1777 | 71.7554 | 173.2328 | 484.484
0.70398 | 1092027.3021 | 61.8291 | 149.2687 | 493.5673
0.6753 | 999987.869 | 81.9697 | 197.8923 | 493.5633
0.68182 | 1021739.2855 | 73.1551 | 176.6121 | 489.6465
0.69197 | 1055241.0641 | 72.6506 | 175.394 | 490.9428
0.70479 | 1056760.9346 | 73.1968 | 176.7126 | 479.8838
0.68594 | 1031448.5923 | 80.0942 | 193.3645 | 486.5012
0.70062 | 1058981.1467 | 66.3879 | 160.2745 | 490.3133
0.71293 | 1095373.2788 | 78.4962 | 189.5065 | 480.0762
0.69933 | 1073940.2029 | 78.456 | 189.4096 | 510.2307
0.68391 | 1021204.7797 | 81.1089 | 195.8142 | 503.7848
0.68278 | 1021933.9376 | 52.8795 | 127.6624 | 471.1223
0.69367 | 1049660.6595 | 66.4624 | 160.4544 | 518.4515
0.6637 | 981714.414 | 72.2079 | 174.3253 | 478.5992
0.69141 | 1053128.4266 | 82.7375 | 199.746 | 511.3788
0.67715 | 1011390.6798 | 87.1505 | 210.3998 | 497.5866
0.69547 | 1050875.4143 | 75.5009 | 182.2754 | 477.3687
0.70441 | 1079112.5496 | 67.2857 | 162.4421 | 491.1067
0.69956 | 1060754.7919 | 66.5999 | 160.7865 | 499.5805
0.67643 | 1022941.7099 | 75.5622 | 182.4233 | 491.8719
0.68483 | 1027056.8905 | 65.8709 | 159.0264 | 488.0165
0.68682 | 1043891.4322 | 74.1965 | 179.1262 | 490.3297
0.68381 | 1027341.8852 | 80.0234 | 193.1936 | 489.1748
0.69006 | 1027516.9021 | 70.9386 | 171.2608 | 500.738
0.68675 | 1024562.038 | 69.2749 | 167.2445 | 483.8639
0.68631 | 1024777.4039 | 57.3503 | 138.4559 | 474.5854
0.68196 | 1019470.3529 | 67.5632 | 163.1121 | 477.3832
0.70677 | 1080417.0909 | 78.0992 | 188.5481 | 483.005
0.66573 | 960588.1757 | 66.9999 | 161.7521 | 479.7682
0.69026 | 1031972.2626 | 63.7052 | 153.7979 | 496.6112
0.66856 | 995243.2557 | 61.4174 | 148.2748 | 470.1571
0.6846 | 1036171.2496 | 62.445 | 150.7557 | 485.7499
0.68632 | 1031825.1017 | 76.8085 | 185.4321 | 496.7276
0.69919 | 1060558.8928 | 75.7911 | 182.9759 | 464.1365
0.69623 | 1056721.6466 | 74.0938 | 178.8782 | 507.8576
0.71369 | 1111479.5683 | 76.0261 | 183.5431 | 525.4742
0.66425 | 973367.8931 | 62.4402 | 150.744 | 483.4029
0.67389 | 999696.6991 | 72.4185 | 174.8337 | 497.76
0.69929 | 1060627.5583 | 60.4534 | 145.9474 | 483.5928
0.69057 | 1035507.9032 | 84.5274 | 204.0671 | 509.6229
0.68277 | 1018042.8979 | 59.4305 | 143.478 | 449.623
0.70496 | 1078513.727 | 72.8853 | 175.9606 | 496.3303
0.70044 | 1064042.9359 | 45.7699 | 110.4982 | 492.3365
0.69428 | 1038821.4682 | 79.6982 | 192.4086 | 483.7569
0.66746 | 992493.9079 | 65.311 | 157.6747 | 490.7902
0.69866 | 1062560.6165 | 85.7319 | 206.975 | 479.5661
0.68034 | 1018843.4528 | 59.6893 | 144.1027 | 468.8262
0.66463 | 976026.518 | 78.2211 | 188.8424 | 474.3656
0.677 | 996293.503 | 75.1848 | 181.5121 | 480.5892
0.70162 | 1086947.3128 | 70.017 | 169.036 | 494.4317
0.69912 | 1052542.8394 | 66.1262 | 159.6428 | 484.1227
0.68687 | 1045230.9624 | 70.6598 | 170.5878 | 468.804
0.70061 | 1067281.8657 | 81.1411 | 195.8919 | 475.4944
0.71751 | 1106010.7166 | 83.7317 | 202.1463 | 497.7027
0.67811 | 1004091.4918 | 65.0358 | 157.0103 | 467.1084
0.70998 | 1083736.5687 | 70.8788 | 171.1165 | 497.954
0.69615 | 1063871.8417 | 69.0691 | 166.7476 | 499.3028
0.68287 | 1028569.9077 | 64.182 | 154.9489 | 487.9849
0.69109 | 1037761.5358 | 76.7176 | 185.2126 | 502.7849
0.69945 | 1068982.3608 | 70.9666 | 171.3286 | 479.4117
0.67046 | 983928.1715 | 64.4572 | 155.6135 | 492.1296
0.70737 | 1073575.7299 | 82.441 | 199.0301 | 474.8423
0.67834 | 1011537.3563 | 77.8225 | 187.8802 | 488.9369
0.68344 | 1019811.3843 | 77.874 | 188.0044 | 455.6367
0.68196 | 1018207.403 | 57.3457 | 138.4448 | 488.0701
0.709 | 1098290.6353 | 85.748 | 207.0139 | 485.4689
0.69112 | 1040808.8428 | 75.2475 | 181.6635 | 502.873
0.68987 | 1026900.1924 | 72.667 | 175.4337 | 492.243
0.68541 | 1023876.1416 | 52.7765 | 127.4138 | 472.8255
0.68873 | 1048727.6175 | 84.288 | 203.4893 | 495.2191
0.68132 | 1017664.3754 | 72.9165 | 176.0359 | 500.4056
0.69839 | 1037352.5947 | 79.1166 | 191.0045 | 474.6001
0.70085 | 1069925.6351 | 77.9941 | 188.2943 | 505.3849
0.70429 | 1063954.147 | 80.4657 | 194.2615 | 500.3165
0.67945 | 1026945.6856 | 77.2179 | 186.4205 | 501.7472
0.69741 | 1068164.7361 | 86.3379 | 208.4382 | 499.0421
0.6824 | 1012854.7749 | 65.1915 | 157.3862 | 506.0393
0.69487 | 1070820.3066 | 67.6851 | 163.4063 | 505.717
0.70727 | 1078547.2486 | 77.9079 | 188.0863 | 465.4814
0.66436 | 972424.8539 | 71.4537 | 172.5045 | 476.6431
0.6739 | 991571.6489 | 73.5202 | 177.4936 | 479.6221
0.68479 | 1038260.846 | 77.4606 | 187.0064 | 477.1602
0.69602 | 1057073.8967 | 74.1403 | 178.9905 | 497.4708
0.68192 | 1012371.7013 | 79.0761 | 190.9067 | 503.8887
0.65638 | 967281.65 | 68.4285 | 165.2011 | 471.074
0.69542 | 1052302.099 | 82.1963 | 198.4393 | 492.953
0.66451 | 974703.8693 | 60.5162 | 146.099 | 527.218
0.69314 | 1059951.9183 | 76.4089 | 184.4675 | 471.0677
0.70981 | 1090296.1115 | 85.0263 | 205.2717 | 474.1727
0.69054 | 1035605.8889 | 67.7796 | 163.6345 | 520.0149
0.71176 | 1096760.2584 | 71.5506 | 172.7383 | 484.6704
0.68208 | 1014842.4163 | 69.0878 | 166.7926 | 497.3948
0.70283 | 1085772.0291 | 77.1911 | 186.3557 | 492.2154
0.68341 | 1021112.2794 | 71.1816 | 171.8477 | 515.1029
0.67634 | 1009204.0906 | 61.5625 | 148.6251 | 520.99
0.66341 | 981171.826 | 81.8688 | 197.6488 | 490.3985
0.71134 | 1088356.9789 | 67.2014 | 162.2385 | 491.4199
0.69745 | 1062731.1206 | 74.3843 | 179.5796 | 531.6076
0.69719 | 1048108.2335 | 75.1164 | 181.3469 | 511.97
0.6788 | 1022421.5657 | 57.8511 | 139.665 | 512.4587
0.69059 | 1044467.2946 | 62.3983 | 150.6429 | 502.6753
0.70489 | 1073366.5468 | 88.8514 | 214.5062 | 492.3157
0.69493 | 1058766.0167 | 77.838 | 187.9175 | 504.4298
0.6871 | 1030580.3141 | 57.6085 | 139.0793 | 480.1089
0.70051 | 1048274.7342 | 79.7082 | 192.4327 | 499.0932
0.67336 | 980618.839 | 71.093 | 171.6337 | 507.8898
0.68377 | 1025769.9233 | 61.8926 | 149.422 | 474.1396
0.69349 | 1049647.4461 | 77.3275 | 186.685 | 513.7103
0.69722 | 1065488.7826 | 74.6027 | 180.1069 | 516.9533
0.7074 | 1088192.2481 | 67.5655 | 163.1176 | 501.9568
0.70049 | 1063205.148 | 82.4463 | 199.0429 | 506.5802
0.69582 | 1053990.8328 | 68.4596 | 165.2762 | 489.7489
0.68287 | 1022820.432 | 67.6954 | 163.4312 | 515.2804
0.67397 | 1010777.597 | 67.7817 | 163.6394 | 473.7247
0.68355 | 1031305.8983 | 78.7395 | 190.0939 | 522.6925
0.66585 | 968487.6009 | 74.3104 | 179.4013 | 467.3562
0.67065 | 1007154.5528 | 83.129 | 200.6912 | 472.6041
0.67844 | 1030057.6549 | 68.6188 | 165.6605 | 474.9036
0.66632 | 979279.4884 | 78.282 | 188.9895 | 467.929
0.6919 | 1050847.3398 | 75.888 | 183.2098 | 484.7252
0.69086 | 1069471.2493 | 71.8811 | 173.5363 | 495.1052
0.67869 | 1013210.6368 | 57.3396 | 138.4301 | 493.5748
0.709 | 1091531.6899 | 82.0028 | 197.9723 | 501.2987
0.70433 | 1084667.4141 | 69.6012 | 168.0323 | 496.5323
0.72153 | 1135503.9322 | 78.6343 | 189.84 | 506.5687
0.7005 | 1050917.0358 | 73.238 | 176.8121 | 481.9169
0.68993 | 1040947.8645 | 50.8246 | 122.7014 | 492.0338
0.68678 | 1046110.7259 | 93.2172 | 225.0462 | 511.5117
0.68224 | 1029093.6858 | 80.3453 | 193.9706 | 490.208
0.69203 | 1050166.9991 | 74.0814 | 178.8482 | 494.0771
0.70752 | 1087820.7364 | 78.55 | 189.6365 | 513.022
0.69667 | 1070652.4003 | 75.0229 | 181.1213 | 474.2774
0.68684 | 1039377.9192 | 78.9755 | 190.6636 | 485.9997
0.68107 | 1035762.7485 | 64.8013 | 156.4441 | 484.3868
0.71222 | 1096076.374 | 73.317 | 177.003 | 459.2806
0.70917 | 1069093.6909 | 70.0007 | 168.9967 | 460.8282
0.70087 | 1079912.7909 | 73.4059 | 177.2174 | 477.9406
0.69078 | 1040827.6829 | 71.7817 | 173.2964 | 499.7386
0.67871 | 1013904.2307 | 63.6548 | 153.6763 | 456.6172
0.68034 | 1019898.0279 | 88.9182 | 214.6674 | 491.9663
0.68345 | 1020392.4083 | 55.0933 | 133.0071 | 491.4593
0.67861 | 1001359.2947 | 52.1334 | 125.8612 | 534.3775
0.68787 | 1046148.7515 | 70.6191 | 170.4896 | 475.7261
0.69311 | 1044633.4861 | 75.3071 | 181.8074 | 497.8493
0.65472 | 966619.294 | 64.8455 | 156.5508 | 483.299
0.69049 | 1032931.5009 | 73.0278 | 176.3048 | 473.2396
0.7167 | 1115960.0464 | 81.2617 | 196.183 | 526.0166
0.67121 | 997938.0488 | 61.3405 | 148.0891 | 489.7584
0.71359 | 1102371.4491 | 91.6294 | 221.2131 | 462.315
0.69612 | 1057923.5701 | 70.5678 | 170.3657 | 487.9941
0.661 | 976395.5898 | 75.4663 | 182.1917 | 487.4808
0.68187 | 1020256.9896 | 61.7317 | 149.0334 | 505.3692
0.68599 | 1024116.8942 | 76.2711 | 184.1347 | 497.0092
0.68257 | 1020612.7404 | 81.2211 | 196.085 | 510.2903
0.69901 | 1071636.6295 | 63.2509 | 152.7011 | 463.7738
0.72323 | 1119917.5495 | 65.9821 | 159.2948 | 456.6808
0.67785 | 1022151.8594 | 68.1519 | 164.5332 | 487.0039
0.68506 | 1022275.1648 | 52.7214 | 127.2807 | 536.3654
0.68964 | 1042858.0756 | 91.129 | 220.0048 | 479.764
0.66579 | 989367.5364 | 74.2497 | 179.2547 | 503.5351
0.67673 | 1008528.6788 | 74.465 | 179.7744 | 506.7765
0.70891 | 1074780.083 | 86.9768 | 209.9805 | 502.6831
0.692 | 1052725.945 | 81.0309 | 195.6259 | 500.837
0.69517 | 1059522.6489 | 81.1814 | 195.9892 | 476.6841
0.70276 | 1066836.4894 | 70.4899 | 170.1777 | 495.3649
0.66827 | 994081.4554 | 69.0729 | 166.7567 | 497.206
0.69639 | 1048975.5117 | 68.0786 | 164.3562 | 476.2086
0.67421 | 1009819.6674 | 69.5481 | 167.904 | 504.5323
0.68137 | 1017763.3596 | 69.3591 | 167.4476 | 489.2757
0.70525 | 1079302.3521 | 65.2105 | 157.4321 | 490.4029
0.67284 | 990025.9766 | 68.3235 | 164.9476 | 496.8428
0.68452 | 1042839.1966 | 72.5789 | 175.221 | 501.6617
0.69437 | 1050702.3735 | 70.2632 | 169.6303 | 490.1188
0.70708 | 1090329.0755 | 67.729 | 163.5122 | 497.3773
0.70461 | 1062517.6942 | 70.0208 | 169.0452 | 481.1834
0.7026 | 1074782.5018 | 74.6769 | 180.286 | 490.1489
0.7162 | 1093395.0563 | 79.6743 | 192.3509
0.68218 | 1020765.7102 | 76.4467 | 184.5586
0.70355 | 1071962.5352 | 66.9218 | 161.5635
0.66243 | 984344.2927 | 57.0731 | 137.7867
0.68896 | 1046184.7411 | 70.5104 | 170.2271
0.69304 | 1031617.1216 | 63.4276 | 153.1279
0.72023 | 1108491.1072 | 73.3225 | 177.0163
0.71133 | 1097678.6559 | 91.5304 | 220.974
0.69627 | 1060854.9556 | 66.9212 | 161.562
0.6817 | 1026758.9909 | 69.9377 | 168.8445
0.7014 | 1074761.5633 | 79.9691 | 193.0625
0.70119 | 1076175.0822 | 74.9238 | 180.882
0.6709 | 1000190.5309 | 59.4212 | 143.4554
0.69063 | 1047998.0319 | 64.6593 | 156.1014
0.7166 | 1090678.3541 | 85.7967 | 207.1317
0.67963 | 1027410.2261 | 62.3691 | 150.5724
0.70712 | 1084361.0826 | 76.6227 | 184.9837
0.68535 | 1023548.9245 | 79.5488 | 192.0479
0.69277 | 1053853.9681 | 68.7519 | 165.9818
0.68792 | 1036838.7404 | 78.6585 | 189.8985
0.70369 | 1075191.1407 | 67.1097 | 162.017
0.6963 | 1053458.4848 | 68.378 | 165.0792
0.69743 | 1059376.3009 | 70.5325 | 170.2805
0.70195 | 1071189.4705 | 94.7946 | 228.8544
0.6975 | 1050781.6313 | 66.1683 | 159.7445
0.71475 | 1113369.1448 | 66.3745 | 160.2423
0.68163 | 1012548.6171 | 83.738 | 202.1614
0.68313 | 1014206.1617 | 63.2616 | 152.727
0.69555 | 1049769.3036 | 81.9832 | 197.925
0.69181 | 1051751.0066 | 77.6701 | 187.5121
0.69634 | 1053611.0478 | 75.6021 | 182.5196
0.67927 | 1014803.6016 | 77.9672 | 188.2295
0.69738 | 1061868.9534 | 79.4649 | 191.8453
0.67312 | 1008778.0418 | 61.8304 | 149.2717
0.71218 | 1084336.2678 | 99.9895 | 241.3961
0.67672 | 1005249.4537 | 83.7369 | 202.1587
0.72331 | 1118222.7733 | 85.5768 | 206.6007
0.69679 | 1054998.7599 | 75.653 | 182.6426
0.68171 | 1011606.9666 | 72.4408 | 174.8877
0.67602 | 1008813.1782 | 72.8495 | 175.8742
0.70887 | 1085482.5809 | 56.5063 | 136.4182
0.68276 | 1023266.6885 | 73.1633 | 176.6319
0.68611 | 1025680.7013 | 74.6957 | 180.3314
0.68315 | 1020106.3151 | 76.2891 | 184.1782
0.6586 | 951877.1187 | 67.111 | 162.0203
0.6556 | 947884.5594 | 70.5812 | 170.3981
0.66996 | 998813.0003 | 74.4925 | 179.8408
0.70086 | 1074221.0714 | 73.0198 | 176.2854
0.69593 | 1058535.5826 | 87.2599 | 210.6642
0.68866 | 1034286.1782 | 76.7301 | 185.243
0.69485 | 1056452.4092 | 48.5878 | 117.3014
0.70496 | 1070168.0918 | 56.0179 | 135.2392
0.67536 | 1008507.2549 | 76.2914 | 184.1837
0.71114 | 1084401.1817 | 72.6466 | 175.3844
0.71759 | 1120324.1896 | 84.8062 | 204.7402
0.68321 | 1013811.8282 | 65.1756 | 157.3478
0.67526 | 1022806.7642 | 55.8939 | 134.9399
0.71337 | 1093958.6328 | 73.0575 | 176.3765
0.70916 | 1098960.5781 | 80.8743 | 195.2478
0.69221 | 1042681.1717 | 87.8925 | 212.1913
0.68813 | 1030998.7608 | 82.9354 | 200.2239
0.6946 | 1055188.0117 | 74.1595 | 179.0369
0.69965 | 1060678.5243 | 80.7497 | 194.9469
0.69168 | 1046055.8377 | 60.2793 | 145.5272
0.67978 | 1020949.4594 | 75.0434 | 181.1708
0.70696 | 1097919.9799 | 87.5449 | 211.352
0.68985 | 1037422.6736 | 71.6438 | 172.9634
0.69683 | 1077662.6237 | 88.7222 | 214.1944
0.65259 | 965357.2128 | 64.9726 | 156.8578
0.68709 | 1036965.0346 | 62.266 | 150.3235
0.66989 | 988867.9256 | 72.0628 | 173.975
0.6816 | 1022132.4155 | 83.0107 | 200.4055
0.6596 | 978754.1501 | 68.9022 | 166.3446
0.6874 | 1049078.2547 | 81.2899 | 196.2513
0.69133 | 1039076.9462 | 66.0152 | 159.3748
0.69333 | 1050741.5639 | 79.0917 | 190.9444
0.70228 | 1066811.4332 | 74.1573 | 179.0315
0.71258 | 1089880.7493 | 67.1779 | 162.1817
0.68358 | 1034793.564 | 79.0025 | 190.7288
0.67591 | 1009900.4971 | 76.3339 | 184.2864
0.71394 | 1092930.4501 | 74.7321 | 180.4192
0.6898 | 1049733.1243 | 66.8399 | 161.3657
0.70103 | 1056837.9468 | 64.237 | 155.0819
0.67181 | 989093.5588 | 84.4622 | 203.9099
0.70503 | 1081151.6089 | 78.3447 | 189.1409
0.69185 | 1040810.056 | 80.1306 | 193.4525
0.69234 | 1037963.6043 | 76.5936 | 184.9132
0.69703 | 1058291.3824 | 80.8709 | 195.2397
0.71599 | 1116662.2124 | 66.4554 | 160.4375
0.7119 | 1086990.8596 | 74.5381 | 179.9509
0.67435 | 1007323.526 | 76.3153 | 184.2414
0.69332 | 1032953.1284 | 61.4675 | 148.3956
0.69887 | 1061178.1343 | 64.7005 | 156.2008
0.68855 | 1048289.2063 | 76.0714 | 183.6525
0.66079 | 974310.4302 | 80.0923 | 193.3599
0.69382 | 1062350.094 | 78.2263 | 188.855
0.70551 | 1054372.1896 | 57.4001 | 138.5761
0.71428 | 1078700.1505 | 88.1195 | 212.7394
0.69563 | 1053387.7746 | 64.5246 | 155.7762
0.71483 | 1095281.6215 | 69.2496 | 167.1834
0.71033 | 1091112.5251 | 89.629 | 216.3836
0.6989 | 1073274.6899 | 80.1441 | 193.4851
0.68291 | 1038482.3712 | 85.7211 | 206.9491
0.69009 | 1061923.269 | 81.9651 | 197.8813
0.69119 | 1038898.6474 | 76.7899 | 185.3871
0.66181 | 955096.0482 | 65.2245 | 157.4659
0.70225 | 1055538.5437 | 69.7948 | 168.4995
0.72195 | 1101092.108 | 87.4968 | 211.236
0.68487 | 1015665.9604 | 80.5559 | 194.4793
0.6908 | 1051576.578 | 70.6843 | 170.6471
0.68216 | 1025284.1116 | 69.7909 | 168.4901
0.66861 | 1002660.2416 | 67.7762 | 163.6261
0.70607 | 1086344.7968 | 71.4431 | 172.4789
0.69397 | 1049522.3553 | 82.1499 | 198.3274
0.71272 | 1093156.4458 | 54.1885 | 130.8226
0.69164 | 1058927.964 | 72.9393 | 176.091
0.71295 | 1075251.5307 | 85.0163 | 205.2474
0.69336 | 1037349.1526 | 73.036 | 176.3246
0.68984 | 1046917.9417 | 72.3002 | 174.5482
0.69235 | 1042728.3775 | 69.9328 | 168.8327
0.68956 | 1047994.7305 | 62.7032 | 151.379
0.70398 | 1083454.4532 | 67.5982 | 163.1964
0.7022 | 1077660.9589 | 63.3174 | 152.8616
0.71007 | 1097200.8826 | 65.8416 | 158.9556
0.6909 | 1044148.4114 | 66.3789 | 160.2529
0.67874 | 1005293.7337 | 68.7743 | 166.0358
0.68925 | 1041981.4334 | 53.4112 | 128.9459
0.69192 | 1052690.8472 | 87.2813 | 210.7157
0.68677 | 1027396.4032 | 63.7445 | 153.8928
0.67082 | 998516.3619 | 69.1168 | 166.8628
0.70405 | 1068001.0536 | 82.3995 | 198.9301
0.69156 | 1049188.3544 | 62.5928 | 151.1125
0.72372 | 1118534.8038 | 75.449 | 182.1501
0.70195 | 1057007.3583 | 81.7956 | 197.472
0.70377 | 1077063.2179 | 74.7105 | 180.3671
0.69777 | 1056516.8486 | 66.5037 | 160.5541
0.70989 | 1089132.3235 | 80.0688 | 193.3032
0.70687 | 1098222.5141 | 79.3742 | 191.6262
0.6671 | 979955.5482 | 68.2107 | 164.6751
0.70316 | 1082614.9595 | 61.3677 | 148.1547
0.67635 | 1024501.4571 | 53.5166 | 129.2005
0.66594 | 963264.56 | 77.0458 | 186.0049
0.67436 | 1000028.2119 | 68.8586 | 166.2395
0.67997 | 1023242.8615 | 64.1222 | 154.8048
0.67679 | 1017387.8175 | 73.3853 | 177.1678
0.70032 | 1057649.9423 | 85.1326 | 205.5283
0.67575 | 997492.218 | 52.1515 | 125.9048
0.6663 | 975376.6209 | 58.3476 | 140.8635
0.69562 | 1053250.4636 | 72.1331 | 174.1446
0.71607 | 1091553.1457 | 77.8785 | 188.0154
0.68187 | 1004668.0651 | 71.927 | 173.6472
0.65568 | 955587.3889 | 80.6732 | 194.7624
0.70307 | 1078083.9536 | 61.1182 | 147.5523
0.72599 | 1129210.3318 | 72.2241 | 174.3645
0.69619 | 1044299.2387 | 69.6518 | 168.1544
0.69431 | 1066648.4665 | 71.3358 | 172.2197
0.68608 | 1034951.0906 | 73.6412 | 177.7856
0.7046 | 1064773.2084 | 80.4917 | 194.3242
0.70008 | 1042510.09 | 69.1466 | 166.9346
0.70884 | 1076118.3397 | 79.2296 | 191.2771
0.71693 | 1115935.8489 | 66.3257 | 160.1245
0.69332 | 1043323.9242 | 75.1908 | 181.5268
0.685 | 1013032.4417 | 68.6706 | 165.7856
0.70556 | 1091397.9907 | 75.5132 | 182.3051
0.65677 | 965111.7796 | 60.6264 | 146.3652
0.68481 | 1022505.4529 | 85.6586 | 206.7981
0.69695 | 1060659.1328 | 84.2416 | 203.3773
0.66967 | 994427.3394 | 79.0953 | 190.953
0.67902 | 1005092.9981 | 81.9368 | 197.813
0.68304 | 1017296.4697 | 63.4623 | 153.2116
0.68607 | 1037001.2366 | 73.7876 | 178.139
0.70009 | 1056451.6289 | 70.0698 | 169.1634
0.68507 | 1027383.329 | 83.0135 | 200.4123
0.68601 | 1037246.3959 | 66.4099 | 160.3276
0.66992 | 997201.631 | 65.38 | 157.8412
0.68374 | 1020635.1465 | 76.2618 | 184.1123
0.69333 | 1053708.8504 | 62.7024 | 151.3771
0.69357 | 1054279.7841 | 79.3306 | 191.521
0.69576 | 1074323.4656 | 69.7869 | 168.4806
0.68669 | 1033901.0498 | 76.9219 | 185.7058
0.69113 | 1062316.3371 | 83.4939 | 201.572
0.7095 | 1083252.0925 | 76.2875 | 184.1743
0.67727 | 1011484.4343 | 65.5969 | 158.3648
0.69941 | 1059285.8796 | 60.3673 | 145.7396
0.69353 | 1050821.4172 | 60.392 | 145.7993
0.66674 | 988833.5621 | 61.1882 | 147.7213
0.67679 | 1004069.2229 | 74.1284 | 178.9618
0.68691 | 1018559.7873 | 68.8536 | 166.2272
0.68688 | 1034276.987 | 68.054 | 164.2969
0.66057 | 982255.2905 | 65.311 | 157.6748
0.69567 | 1047316.2003 | 86.3669 | 208.5082
0.7017 | 1067560.417 | 64.8268 | 156.5058
0.67067 | 996781.5324 | 69.0128 | 166.6117
0.6803 | 1024396.5859 | 73.509 | 177.4663
0.7058 | 1073836.876 | 45.84 | 110.6675
0.68468 | 1020341.2459 | 84.0058 | 202.808
0.71454 | 1092232.6152 | 63.0654 | 152.2534
0.6762 | 1003169.8766 | 75.7929 | 182.9803
0.68949 | 1029853.658 | 74.6323 | 180.1783
0.68522 | 1026464.0097 | 79.7912 | 192.633
0.68487 | 1023638.8318 | 76.3605 | 184.3506
0.6842 | 1021518.2913 | 77.9176 | 188.1097
0.69163 | 1042452.0656 | 71.8056 | 173.3541
0.6886 | 1040123.9927 | 73.7382 | 178.0199
0.69703 | 1057481.8806 | 77.5087 | 187.1226
0.68573 | 1035112.5828 | 75.4243 | 182.0904
0.6933 | 1040734.8734 | 73.2807 | 176.9152
0.68667 | 1039014.2767 | 66.4597 | 160.448
0.71244 | 1084521.8469 | 68.7999 | 166.0976
0.6996 | 1063259.5305 | 70.5707 | 170.3727
0.71944 | 1121680.6518 | 55.083 | 132.9822
0.69895 | 1040731.3847 | 70.9368 | 171.2565
0.70633 | 1079221.8439 | 51.4694 | 124.2581
0.71214 | 1106858.8672 | 77.6611 | 187.4904
0.68525 | 1035532.5577 | 77.205 | 186.3894
0.70309 | 1078048.8823 | 76.6129 | 184.9599
0.69958 | 1063203.0984 | 77.6439 | 187.4491
0.68503 | 1027883.5808 | 35.7405 | 86.2852
0.67816 | 1012772.8768 | 59.5246 | 143.7051
0.69386 | 1058843.7487 | 77.9629 | 188.2191
0.67494 | 1004327.9653 | 79.2588 | 191.3476
0.67198 | 994046.6523 | 77.2017 | 186.3813
0.69152 | 1044828.3114 | 77.1914 | 186.3565
0.69243 | 1044088.6696 | 77.0352 | 185.9794
0.70264 | 1063055.6026 | 78.9584 | 190.6225
0.68486 | 1039507.7213 | 79.214 | 191.2394
0.6774 | 1013193.2553 | 78.0589 | 188.451
0.67246 | 997578.9422 | 61.204 | 147.7594
0.7047 | 1074589.2718 | 60.7796 | 146.735
0.69203 | 1041149.3339 | 69.8592 | 168.6551
0.67869 | 1021540.7688 | 73.4779 | 177.3914
0.71479 | 1102340.6908 | 80.484 | 194.3056
0.73178 | 1146769.8959 | 72.7663 | 175.6735
0.68621 | 1048480.6993 | 69.0556 | 166.7149
0.6669 | 992204.2274 | 77.6453 | 187.4522
0.7062 | 1071785.1933 | 68.5926 | 165.5973
0.70086 | 1083435.1943 | 82.3395 | 198.7852
0.70368 | 1071665.9868 | 66.8906 | 161.4882
0.70603 | 1068995.891 | 70.3057 | 169.7329
0.68469 | 1029145.1463 | 71.6079 | 172.8768
0.72536 | 1117580.7519 | 69.6342 | 168.1119
0.68639 | 1031259.0332 | 50.6109 | 122.1855
0.69366 | 1056018.0834 | 68.8052 | 166.1104
0.71192 | 1090474.8901 | 71.1251 | 171.7111
0.68435 | 1020740.3833 | 69.2586 | 167.2051
0.68859 | 1026569.6197 | 79.408 | 191.7079
0.70645 | 1064903.3543 | 60.0824 | 145.0518
0.72048 | 1113554.3669 | 73.6321 | 177.7637
0.68381 | 1024417.0583 | 57.5396 | 138.913
0.6748 | 999556.8605 | 83.4992 | 201.5849
0.719 | 1118158.8707 | 82.8352 | 199.9819
0.70462 | 1079091.0331 | 90.8373 | 219.3007
0.72323 | 1122690.9927 | 85.9289 | 207.4506
0.6784 | 1017479.014 | 69.6432 | 168.1336
0.6934 | 1049009.4147 | 77.2903 | 186.5953
0.69757 | 1057100.6798 | 70.3665 | 169.8797
0.68313 | 1016764.106 | 82.3899 | 198.9068
0.70341 | 1059288.2166 | 71.7432 | 173.2034
0.68166 | 1017153.0217 | 69.288 | 167.276
0.70525 | 1087656.7487 | 85.3692 | 206.0994
0.6934 | 1041838.9218 | 80.0305 | 193.2108
0.67923 | 1010438.0361 | 86.802 | 209.5586
0.71019 | 1101159.5819 | 61.8006 | 149.2
0.69394 | 1079913.7975 | 70.3337 | 169.8006
0.70265 | 1052632.6166 | 67.7065 | 163.4581
0.66671 | 982258.2569 | 70.189 | 169.4512
0.68602 | 1034134.0718 | 73.5887 | 177.6588
0.71623 | 1085121.5879 | 68.473 | 165.3084
0.6681 | 972099.2954 | 77.5015 | 187.1052
0.6667 | 987513.8845 | 65.0576 | 157.063
0.68952 | 1034786.9417 | 76.6714 | 185.101
0.66097 | 962604.1257 | 79.1091 | 190.9864
0.69811 | 1067411.2757 | 76.399 | 184.4434
0.70301 | 1063304.9682 | 70.2314 | 169.5537
0.68952 | 1045856.5404 | 66.6203 | 160.8356
0.65924 | 983225.254 | 63.5519 | 153.4279
0.71311 | 1109562.2943 | 81.2023 | 196.0397
0.70721 | 1074643.6998 | 72.3924 | 174.7707
0.69993 | 1059550.4221 | 60.3374 | 145.6672
0.71245 | 1074072.4025 | 89.4205 | 215.8802
0.66823 | 994718.0013 | 76.0507 | 183.6027
0.69356 | 1040109.7778 | 81.4616 | 196.6656
0.67372 | 999441.1849 | 70.6734 | 170.6207
0.70338 | 1058075.8289 | 86.2981 | 208.3421
0.68404 | 1022460.0233 | 65.6515 | 158.4966
0.68649 | 1033301.264 | 66.1115 | 159.6072
0.68116 | 1020488.1319 | 82.9297 | 200.21
0.69331 | 1054145.2419 | 84.1644 | 203.1908
0.67318 | 1013002.9999 | 73.4607 | 177.3497
0.70325 | 1084930.8822 | 65.7364 | 158.7018
0.71022 | 1101851.7446 | 72.0388 | 173.9171
0.69651 | 1065050.6121 | 79.8027 | 192.6607
0.6694 | 991497.6737 | 80.4571 | 194.2407
0.69068 | 1041738.2267 | 70.0353 | 169.0803
0.69318 | 1059839.3507 | 60.9418 | 147.1264
0.69739 | 1038565.1775 | 76.083 | 183.6806
0.67522 | 1011910.0296 | 78.2115 | 188.8193
0.69136 | 1036821.7537 | 71.4464 | 172.4868
0.7084 | 1079648.9453 | 76.1517 | 183.8465
0.69214 | 1028746.4495 | 89.5711 | 216.2437
0.68364 | 1025769.9724 | 67.8137 | 163.7168
0.67008 | 997965.7941 | 74.4153 | 179.6545
0.70475 | 1061536.9559 | 75.2625 | 181.6997
0.69739 | 1073085.6275 | 75.1163 | 181.3468
0.68398 | 1031590.3687 | 81.9255 | 197.7856
0.69722 | 1063840.7094 | 76.2406 | 184.061
0.70861 | 1097409.4425 | 75.6021 | 182.5197
0.71176 | 1076338.2554 | 64.4361 | 155.5625
0.6585 | 964433.564 | 74.8254 | 180.6444
0.6935 | 1031853.9665 | 63.3998 | 153.0607
0.67866 | 1020987.9076 | 57.7823 | 139.4989
0.69003 | 1049611.1609 | 74.2559 | 179.2696
0.67215 | 1001462.0538 | 74.4465 | 179.7298
0.72406 | 1115800.7739 | 99.862 | 241.0881
0.66636 | 983152.9995 | 75.3779 | 181.9784
0.69086 | 1035186.9123 | 81.3923 | 196.4984
0.69533 | 1065982.5989 | 71.1865 | 171.8594
0.66696 | 991135.3402 | 72.5443 | 175.1375
0.68035 | 1022778.781 | 87.7176 | 211.7691
0.69412 | 1039728.9851 | 71.3616 | 172.2822
0.70608 | 1094728.0237 | 82.5829 | 199.3729
0.68887 | 1042700.7315 | 73.8919 | 178.3907
0.68608 | 1033645.0111 | 66.1729 | 159.7555
0.68301 | 1030089.1924 | 75.3576 | 181.9294
0.69719 | 1070629.3165 | 67.448 | 162.8339
0.69945 | 1067090.5883 | 51.3919 | 124.0711
0.70695 | 1082113.7497 | 69.1801 | 167.0155
0.69896 | 1063738.0317 | 66.8448 | 161.3777
0.67934 | 1015106.3712 | 69.3845 | 167.5089
0.66075 | 975838.9424 | 81.7511 | 197.3647
0.69319 | 1048134.1833 | 72.0351 | 173.908
0.68253 | 1028970.4255 | 81.8911 | 197.7027
0.70537 | 1086106.7747 | 75.167 | 181.4692
0.70528 | 1084901.8682 | 54.347 | 131.2052
0.67373 | 1003281.2481 | 65.4064 | 157.9051
0.72827 | 1120072.5482 | 86.6979 | 209.3073
0.68877 | 1014716.4552 | 73.6122 | 177.7156
0.67856 | 1021670.5413 | 84.8744 | 204.9049
0.67473 | 1000771.7781 | 68.8718 | 166.2712
0.68307 | 1025626.1754 | 78.3608 | 189.1798
0.70848 | 1095154.1107 | 80.6717 | 194.7588
0.66872 | 994479.209 | 86.774 | 209.491
0.69471 | 1059990.4139 | 56.3039 | 135.9297
0.66868 | 998754.1903 | 85.687 | 206.8666
0.67994 | 1001109.4868 | 70.1441 | 169.3428
0.68241 | 1028115.2627 | 81.8375 | 197.5731
0.67398 | 1005887.7165 | 72.8583 | 175.8955
0.67339 | 983018.3453 | 72.941 | 176.0952
0.68066 | 1022176.2748 | 69.7318 | 168.3474
0.67302 | 1013038.9064 | 80.1748 | 193.5591
0.69797 | 1075843.437 | 69.6637 | 168.183
0.67897 | 1011744.1411 | 72.7431 | 175.6173
0.702 | 1073133.8198 | 76.3827 | 184.4042
0.68482 | 1031245.039 | 78.5889 | 189.7304
0.69683 | 1064085.1379 | 74.5457 | 179.9691
0.692 | 1057762.0898 | 78.0927 | 188.5324
0.66847 | 980233.0425 | 71.5847 | 172.8208
0.68619 | 1040711.4037 | 77.4844 | 187.064
0.68968 | 1027317.4809 | 76.1001 | 183.7218
0.67944 | 1009013.3185 | 75.8708 | 183.1684
0.68973 | 1030830.1992 | 79.606 | 192.1859
0.6799 | 1011855.7221 | 72.6521 | 175.3977
0.67787 | 1008769.6315 | 65.362 | 157.7979
0.69533 | 1058372.6091 | 77.7634 | 187.7375
0.69978 | 1083365.4889 | 63.8722 | 154.2012
0.68874 | 1032425.5216 | 65.8393 | 158.9501
0.6953 | 1054442.7458 | 80.7778 | 195.0149
0.67844 | 1003867.0863 | 70.4926 | 170.1843
0.69506 | 1053186.774 | 73.9706 | 178.5807
0.66582 | 990159.0485 | 89.6189 | 216.3592
0.6678 | 986232.3631 | 66.7461 | 161.1394
0.66658 | 994471.7529 | 55.0533 | 132.9103
0.67794 | 1018171.5518 | 63.9953 | 154.4983
0.69293 | 1066103.019 | 69.5351 | 167.8725
0.6916 | 1050381.7196 | 79.2617 | 191.3548
0.68416 | 1013515.6737 | 77.3118 | 186.6472
0.68989 | 1042639.5442 | 75.5147 | 182.3085
0.69144 | 1044205.6886 | 68.2537 | 164.7791
0.67146 | 987395.0064 | 71.7866 | 173.3081
0.70939 | 1084124.8403 | 56.3947 | 136.1488
0.67711 | 1016707.803 | 76.713 | 185.2015
0.67881 | 1021620.9131 | 75.1998 | 181.5483
0.68762 | 1030104.9676 | 55.2805 | 133.4588
0.70669 | 1070618.5603 | 80.6946 | 194.8141
0.69071 | 1050334.5124 | 79.3608 | 191.5938
0.68638 | 1032133.4884 | 69.8246 | 168.5716
0.6973 | 1084745.0978 | 77.7295 | 187.6557
0.70103 | 1069932.1804 | 76.26 | 184.1079
0.67576 | 1005928.2535 | 63.4294 | 153.1322
0.69881 | 1071068.1422 | 67.3195 | 162.5237
0.70371 | 1092920.1131 | 79.3457 | 191.5576
0.68652 | 1029864.6244 | 84.5627 | 204.1523
0.71315 | 1088013.7498 | 76.231 | 184.0378
0.68518 | 1046994.5259 | 76.351 | 184.3276
0.68553 | 1033089.2671 | 76.9912 | 185.8731
0.67783 | 1032868.8484 | 72.5942 | 175.258
0.69152 | 1059942.889 | 61.0932 | 147.492
0.69733 | 1052556.2464 | 71.9095 | 173.6048
0.6911 | 1042904.6596 | 83.5334 | 201.6675
0.68841 | 1015062.8281 | 78.5387 | 189.6092
0.68936 | 1033655.2012 | 84.5245 | 204.0603
0.70363 | 1062595.9857 | 64.3489 | 155.3521
0.69632 | 1051842.6556 | 73.742 | 178.0289
0.70038 | 1051141.5311 | 74.9901 | 181.0421
0.69059 | 1034896.0364 | 65.8197 | 158.9029
0.71456 | 1097664.8613 | 67.4708 | 162.8888
0.70053 | 1067877.1336 | 79.5556 | 192.0641
0.70176 | 1056409.1659 | 76.9089 | 185.6744
0.70721 | 1067361.0762 | 77.1067 | 186.152
0.6705 | 1004483.4238 | 67.7082 | 163.4621
0.67182 | 990422.7521 | 71.9686 | 173.7475
0.67075 | 996917.5021 | 54.8452 | 132.4079
0.70312 | 1067853.6348 | 64.7141 | 156.2337
0.69186 | 1046628.34 | 73.8491 | 178.2875
0.6847 | 1027707.1037 | 69.6958 | 168.2605
0.69082 | 1056085.892 | 89.2008 | 215.3498
0.69182 | 1043882.8937 | 73.6603 | 177.8317
0.69311 | 1047670.6054 | 62.3628 | 150.5572
0.71194 | 1086436.9534 | 56.3124 | 135.9502
0.66083 | 954316.1357 | 75.3661 | 181.9497
0.68501 | 1031730.6127 | 66.1538 | 159.7094
0.67474 | 1006398.7798 | 65.5116 | 158.1589
0.66319 | 976807.859 | 70.1698 | 169.405
0.66834 | 996760.5986 | 67.4043 | 162.7284
0.66847 | 987816.7295 | 73.8362 | 178.2564
0.68524 | 1025575.9087 | 73.0634 | 176.3907
0.68692 | 1035301.6852 | 64.2213 | 155.0439
0.7198 | 1115222.1741 | 82.5255 | 199.2343
0.68116 | 1006327.6538 | 70.3835 | 169.9207
0.70036 | 1078465.383 | 76.9412 | 185.7524
0.67293 | 992875.1721 | 49.7016 | 119.9904
0.68754 | 1034064.399 | 75.1497 | 181.4275
0.67496 | 998652.3529 | 66.87 | 161.4383
0.69128 | 1057127.1034 | 74.6531 | 180.2286
0.70533 | 1074723.3409 | 68.8653 | 166.2554
0.68364 | 1039561.7662 | 94.4586 | 228.0431
0.70866 | 1081054.6902 | 68.9451 | 166.4483
0.67945 | 1021671.0437 | 83.235 | 200.9471
0.68712 | 1032660.9087 | 84.718 | 204.5274
0.69216 | 1031991.0615 | 71.0303 | 171.4824
0.6842 | 1040556.7293 | 70.6192 | 170.4898
0.7107 | 1085475.2009 | 70.9209 | 171.2181
0.69626 | 1053448.796 | 91.1798 | 220.1275
0.66702 | 975633.879 | 73.0843 | 176.4411
0.70353 | 1082731.1977 | 75.1299 | 181.3796
0.69661 | 1065262.625 | 58.7345 | 141.7976
0.69121 | 1048426.0931 | 82.3917 | 198.9111
0.71068 | 1097882.4864 | 63.7357 | 153.8716
0.68742 | 1032093.8774 | 59.8429 | 144.4736
0.70656 | 1084943.4423 | 69.3992 | 167.5446
0.68118 | 1013986.0595 | 72.6686 | 175.4375
0.69468 | 1060440.2073 | 84.4995 | 203.9998
0.66373 | 984445.3368 | 79.0027 | 190.7293
0.68723 | 1031719.6859 | 65.9393 | 159.1914
0.69386 | 1037921.4183 | 69.4595 | 167.69
0.68876 | 1048617.1568 | 79.4549 | 191.821
0.69095 | 1037356.3111 | 73.292 | 176.9426
0.69994 | 1065674.9479 | 66.8725 | 161.4445
0.70983 | 1082028.3002 | 72.6214 | 175.3236
0.67149 | 985559.0516 | 73.3368 | 177.0508
0.7002 | 1069518.6642 | 89.1253 | 215.1676
0.69569 | 1056131.8983 | 77.5587 | 187.2434
0.68207 | 1006756.4817 | 54.4672 | 131.4955
0.6809 | 1012829.7028 | 75.9947 | 183.4675
0.68835 | 1055341.6394 | 69.291 | 167.2832
0.70792 | 1092158.759 | 83.8033 | 202.319
0.67604 | 999929.98 | 68.6804 | 165.8091
0.70018 | 1062997.9688 | 66.6811 | 160.9823
0.69925 | 1064908.5 | 72.5528 | 175.158
0.67839 | 1008668.7048 | 67.1993 | 162.2334
0.67305 | 1011809.5222 | 67.9431 | 164.0291
0.68229 | 1026678.2588 | 75.1866 | 181.5165
0.70557 | 1054529.5662 | 77.9991 | 188.3064
0.67941 | 1022030.1782 | 80.1256 | 193.4403
0.68383 | 1034183.6641 | 72.0484 | 173.9401
0.69598 | 1061037.1134 | 60.7649 | 146.6994
0.66768 | 991693.9935 | 72.5078 | 175.0492
0.66757 | 997686.2229 | 71.1989 | 171.8893
0.70496 | 1085863.472 | 67.6544 | 163.3322
0.70178 | 1059801.7256 | 83.5943 | 201.8145
0.68882 | 1054892.9941 | 89.3595 | 215.7328
0.70239 | 1054982.8268 | 59.8671 | 144.5319
0.70884 | 1093401.3115 | 82.1616 | 198.3558
0.70627 | 1061898.6409 | 76.8147 | 185.447
0.70917 | 1091117.8828 | 69.5369 | 167.8769
0.68017 | 1012308.9317 | 74.3651 | 179.5333
0.67821 | 1020153.3703 | 64.0938 | 154.7361
0.70113 | 1067798.8256 | 60.7032 | 146.5505
0.69666 | 1062955.0713 | 93.3756 | 225.4287
0.7018 | 1058825.1424 | 75.3851 | 181.9958
0.68591 | 1041319.1642 | 68.614 | 165.649
0.67986 | 1025143.5053 | 71.8358 | 173.4268
0.67691 | 1008773.409 | 85.494 | 206.4008
0.70061 | 1059190.1518 | 72.0232 | 173.8793
0.70883 | 1100549.3175 | 85.9449 | 207.4894
0.69921 | 1047582.7848 | 81.3842 | 196.4787
0.67246 | 998260.3348 | 68.931 | 166.4142
0.67708 | 1006372.6647 | 75.9364 | 183.3267
0.69965 | 1069625.5346 | 78.1058 | 188.5641
0.68318 | 1021299.5667 | 78.0177 | 188.3513
0.66906 | 997681.1663 | 69.2407 | 167.162
0.68499 | 1017989.1664 | 59.2352 | 143.0064
0.704 | 1088217.6115 | 67.8929 | 163.908
0.72127 | 1119533.6172 | 77.5564 | 187.2377
0.68907 | 1033973.3479 | 78.3103 | 189.0579
0.68906 | 1042639.8979 | 78.9678 | 190.6452
0.6955 | 1063372.2085 | 76.3948 | 184.4335
0.68216 | 1019701.3741 | 61.6723 | 148.8901
0.70165 | 1069414.048 | 56.2432 | 135.7831
0.66292 | 967250.4921 | 76.1465 | 183.8338
0.67348 | 983893.4489 | 71.3177 | 172.1761
0.70342 | 1069466.6591 | 80.9821 | 195.5081
0.67646 | 1007786.4471 | 63.4806 | 153.2557
0.6726 | 994680.7653 | 70.4542 | 170.0915
0.69142 | 1041379.4355 | 84.9021 | 204.9719
0.66503 | 969134.639 | 64.0062 | 154.5247
0.69966 | 1055734.5554 | 60.1434 | 145.1989
0.68495 | 1017619.2403 | 75.9023 | 183.2443
0.66454 | 978687.1296 | 73.5867 | 177.654
0.68272 | 1027166.5826 | 86.3118 | 208.3751
0.68968 | 1041343.4845 | 83.579 | 201.7776
0.68711 | 1038690.2865 | 78.8919 | 190.4619
0.70416 | 1066911.54 | 55.911 | 134.9811
0.67838 | 1011996.2811 | 72.825 | 175.815
0.71974 | 1111660.3115 | 85.4972 | 206.4084
0.66914 | 987768.1254 | 70.326 | 169.7821
0.70145 | 1055208.5675 | 86.3008 | 208.3485
0.70103 | 1057373.0745 | 77.0586 | 186.036
0.68364 | 1025546.3656 | 61.8912 | 149.4186
0.697 | 1051163.1477 | 73.6228 | 177.7412
0.67067 | 986568.8309 | 75.1534 | 181.4364
0.70031 | 1064536.938 | 60.5885 | 146.2736
0.68983 | 1045177.8607 | 70.353 | 169.8473
0.67452 | 1010395.4727 | 78.1753 | 188.7318
0.66116 | 954733.5451 | 61.5308 | 148.5484
0.68798 | 1030971.6368 | 69.5836 | 167.9897
0.66176 | 972757.4444 | 75.0986 | 181.304
0.71742 | 1115007.7116 | 81.1107 | 195.8185
0.68392 | 1032951.3797 | 73.7307 | 178.0017
0.69899 | 1066811.728 | 87.5564 | 211.3799
0.69543 | 1048114.477 | 84.9588 | 205.1086
0.68455 | 1028793.9679 | 73.3663 | 177.1219
0.67516 | 1004300.575 | 63.3832 | 153.0206
0.68471 | 1016407.6342 | 77.4698 | 187.0286
0.67266 | 990781.3097 | 66.5698 | 160.7137
0.67652 | 1009865.4604 | 81.0844 | 195.7552
0.68572 | 1037501.6697 | 75.5992 | 182.5127
0.72198 | 1116394.9385 | 73.667 | 177.8478
0.67412 | 1019335.2958 | 73.817 | 178.21
0.69696 | 1034990.709 | 72.4943 | 175.0166
0.69436 | 1051334.5865 | 85.351 | 206.0554
0.67429 | 1002348.5005 | 65.2913 | 157.6271
0.70505 | 1088537.0085 | 87.032 | 210.1138
0.68379 | 1018891.6397 | 57.2305 | 138.1666
0.68128 | 1014090.0146 | 77.6832 | 187.544
0.69542 | 1049455.9166 | 85.0687 | 205.3741
0.69154 | 1053567.6074 | 72.5814 | 175.227
0.71194 | 1107714.458 | 64.1923 | 154.9739
0.68838 | 1033193.6759 | 79.9455 | 193.0056
0.69201 | 1037294.9726 | 81.4233 | 196.5733
0.71548 | 1096649.4934 | 71.1723 | 171.8251
0.72069 | 1105143.6936 | 74.8591 | 180.7259
0.70925 | 1101980.3098 | 82.7585 | 199.7966
0.70804 | 1086136.9007 | 67.3769 | 162.6621
0.672 | 1004071.529 | 74.6576 | 180.2395
0.6947 | 1049961.7806 | 72.0034 | 173.8316
0.68136 | 1039406.9968 | 67.0027 | 161.7587
0.68215 | 1027659.3133 | 80.2438 | 193.7258
0.65806 | 964639.1328 | 68.067 | 164.3283
0.67259 | 1002107.4642 | 83.9878 | 202.7646
0.68599 | 1021840.3354 | 74.4292 | 179.688
0.70717 | 1079242.7041 | 74.5021 | 179.8639
0.67524 | 1011992.9068 | 71.3995 | 172.3737
0.6714 | 994690.308 | 70.58 | 170.3953
0.68077 | 1023042.6781 | 80.0443 | 193.2439
0.69215 | 1043277.7806 | 77.0067 | 185.9106
0.71124 | 1093198.1762 | 80.2147 | 193.6553
0.68802 | 1046878.9543 | 72.4891 | 175.0041
0.67234 | 998430.8082 | 67.116 | 162.0323
0.66101 | 971038.9484 | 60.6082 | 146.3211
0.67574 | 1004566.957 | 77.5838 | 187.3039
0.70708 | 1079872.1536 | 72.7583 | 175.6541
0.67018 | 989376.1671 | 72.0682 | 173.988
0.6963 | 1056770.9852 | 72.3035 | 174.556
0.69184 | 1057132.5185 | 72.7946 | 175.7418
0.66389 | 976426.9829 | 75.1405 | 181.4052
0.67955 | 1024026.1468 | 75.2952 | 181.7787
0.69069 | 1034315.7452 | 87.3788 | 210.9511
0.69432 | 1055819.8453 | 64.7953 | 156.4298
0.68927 | 1022150.5531 | 74.0973 | 178.8867
0.69796 | 1053817.59 | 60.1802 | 145.2877
0.67946 | 1008191.6704 | 79.0352 | 190.8079
0.69762 | 1055309.2165 | 75.585 | 182.4784
0.67611 | 1015415.1145 | 71.5487 | 172.7339
0.67982 | 1025979.749 | 80.3865 | 194.0702
0.67767 | 993581.4416 | 69.8492 | 168.6309
0.68255 | 1018904.3173 | 67.4673 | 162.8804
0.70845 | 1084394.8486 | 74.723 | 180.3973
0.6756 | 1018694.2405 | 60.0993 | 145.0925
0.70701 | 1090121.8861 | 85.3496 | 206.0522
0.65792 | 978156.6825 | 77.5911 | 187.3216
0.73319 | 1174695.6783 | 83.6151 | 201.8646
0.68283 | 1001572.8004 | 79.6765 | 192.3561
0.69695 | 1070758.6049 | 81.8329 | 197.5622
0.68633 | 1026359.3385 | 71.2432 | 171.9962
0.68729 | 1039396.4949 | 64.4402 | 155.5725
0.69825 | 1060467.0373 | 70.5224 | 170.256
0.67926 | 1007689.5449 | 83.626 | 201.8911
0.7108 | 1087211.2772 | 87.7484 | 211.8434
0.704 | 1070315.6996 | 79.3434 | 191.5519
0.71362 | 1103196.2193 | 67.8999 | 163.9249
0.67997 | 1008409.7607 | 65.1371 | 157.2548
0.69945 | 1062906.2112 | 82.5465 | 199.2848
0.69979 | 1075294.878 | 65.4752 | 158.0712
0.70377 | 1070513.5394 | 76.0038 | 183.4895
0.69763 | 1070016.771 | 71.2205 | 171.9414
0.68739 | 1053574.5879 | 71.1705 | 171.8207
0.68161 | 1030946.3845 | 77.2741 | 186.5561
0.69696 | 1049315.2082 | 86.3724 | 208.5215
0.69212 | 1039762.7804 | 78.4022 | 189.2798
0.71528 | 1115420.88 | 76.1264 | 183.7854
0.69125 | 1037406.8313 | 71.4972 | 172.6096
0.68438 | 1027367.7055 | 83.0966 | 200.6128
0.66915 | 993209.5481 | 74.1637 | 179.047
0.68999 | 1047612.5036 | 74.1189 | 178.9388
0.68973 | 1042059.1843 | 77.1866 | 186.345
0.67603 | 995452.7907 | 66.5501 | 160.6661
0.70311 | 1069878.869 | 60.65 | 146.4221
0.69032 | 1052904.0084 | 82.8509 | 200.0199
0.70988 | 1090945.1515 | 76.7885 | 185.3838
0.69268 | 1064200.9742 | 67.48 | 162.911
0.67685 | 1010453.3644 | 67.4711 | 162.8896
0.67998 | 1003734.1163 | 88.9561 | 214.759
0.66427 | 969759.2017 | 72.9453 | 176.1054
0.6709 | 998618.3285 | 74.373 | 179.5524
0.68268 | 1034195.8054 | 62.5006 | 150.8899
0.67567 | 1006306.1878 | 86.9026 | 209.8013
0.69531 | 1055504.7394 | 84.03 | 202.8664
0.69019 | 1044839.368 | 73.9017 | 178.4145
0.67435 | 1006383.2796 | 69.2971 | 167.2981
0.689 | 1032229.9347 | 76.5402 | 184.7843
0.68997 | 1024771.9725 | 79.1114 | 190.9919
0.69925 | 1070198.1382 | 81.5212 | 196.8096
0.67755 | 1009766.6706 | 88.3974 | 213.4101
0.68591 | 1034259.8713 | 75.3422 | 181.8922
0.68085 | 1016259.9845 | 69.3092 | 167.3273
0.70251 | 1076316.4917 | 74.7933 | 180.567
0.66542 | 971826.9459 | 77.7918 | 187.8061
0.69261 | 1043794.015 | 68.112 | 164.4369
0.65942 | 956983.076 | 65.5902 | 158.3488
0.69302 | 1045027.3579 | 78.6968 | 189.991
0.71126 | 1081686.0095 | 72.4515 | 174.9135
0.67989 | 1022405.9259 | 80.0983 | 193.3745
0.69362 | 1063754.0207 | 65.3484 | 157.7651
0.72561 | 1118686.9485 | 71.588 | 172.8287
0.70139 | 1069566.3923 | 68.92 | 166.3876
0.69664 | 1055707.9378 | 77.422 | 186.9133
0.67897 | 1027154.8518 | 78.9512 | 190.6049
0.67847 | 1012771.3146 | 71.1142 | 171.6848
0.68778 | 1045954.8626 | 74.2398 | 179.2307
0.6622 | 977979.4738 | 62.9297 | 151.9258
0.69589 | 1058292.0397 | 77.6952 | 187.5728
0.68323 | 1018839.5718 | 91.2595 | 220.32
0.67703 | 1011131.1227 | 77.9421 | 188.169
0.67635 | 997982.539 | 64.809 | 156.4628
0.71565 | 1102020.3239 | 76.9427 | 185.7562
0.69448 | 1034929.5461 | 67.3515 | 162.6009
0.69466 | 1065179.8558 | 58.8942 | 142.1831
0.68195 | 1022185.4456 | 65.8351 | 158.94
0.70108 | 1081237.3872 | 63.6289 | 153.6137
0.69533 | 1055281.3742 | 66.659 | 160.9291
0.6836 | 1026604.9988 | 79.9906 | 193.1144
0.66766 | 990819.7609 | 67.0713 | 161.9244
0.68528 | 1028343.3333 | 78.3152 | 189.0695
0.68432 | 1024802.6453 | 85.9119 | 207.4096
0.68806 | 1032862.4516 | 75.0054 | 181.079
0.69936 | 1055583.366 | 73.683 | 177.8865
0.69449 | 1048997.9463 | 78.8278 | 190.3072
0.66199 | 980634.2614 | 61.1166 | 147.5486
0.69281 | 1043949.0107 | 71.7863 | 173.3075
0.67585 | 1023889.1548 | 73.9563 | 178.5463
0.68347 | 1008577.2218 | 78.6012 | 189.76
0.67493 | 1013142.2657 | 79.9555 | 193.0298
0.66037 | 965576.0279 | 79.4253 | 191.7495
0.68656 | 1036636.2587 | 77.1403 | 186.2331
0.70294 | 1065176.8669 | 73.9739 | 178.5888
0.68849 | 1022583.3352 | 79.1111 | 190.9912
0.70604 | 1089618.9808 | 83.8301 | 202.3837
0.67937 | 1026192.5037 | 73.7391 | 178.0219
0.68328 | 1024275.9486 | 90.3389 | 218.0973
0.67313 | 995163.7902 | 80.3071 | 193.8784
0.6946 | 1051893.4225 | 73.9283 | 178.4786
0.68367 | 1024522.646 | 59.7303 | 144.2017
0.71157 | 1092894.8662 | 67.4591 | 162.8606
0.68986 | 1041728.7785 | 78.8069 | 190.2567
0.69462 | 1057481.8259 | 64.8695 | 156.6089
0.71141 | 1099350.9633 | 73.6197 | 177.7337
0.66679 | 978004.5542 | 73.2867 | 176.9298
0.68963 | 1037690.1648 | 81.7347 | 197.3251
0.69587 | 1045779.924 | 81.6336 | 197.0809
0.66144 | 973853.9103 | 76.8378 | 185.5029
0.70141 | 1062025.8469 | 72.1338 | 174.1464
0.69498 | 1032688.5557 | 63.3157 | 152.8577
0.67661 | 1002878.6781 | 75.7255 | 182.8176
0.68313 | 1027912.4889 | 66.0903 | 159.5562
0.67504 | 1008315.7104 | 74.6351 | 180.1851
0.67663 | 1017185.9721 | 89.1672 | 215.2686
0.68833 | 1057503.7819 | 74.9997 | 181.0652
0.68491 | 1027987.9402 | 76.6803 | 185.1225
0.70219 | 1069637.1622 | 77.8243 | 187.8844
0.68389 | 1026605.579 | 79.957 | 193.0332
0.66574 | 977386.1162 | 69.942 | 168.855
0.68344 | 1014036.5036 | 72.4312 | 174.8644
0.66312 | 963864.1745 | 63.8416 | 154.1272
0.66174 | 969909.2488 | 73.2355 | 176.8062
0.71772 | 1125241.7748 | 70.8576 | 171.0654
0.68987 | 1036304.3612 | 73.3323 | 177.0398
0.679 | 1019778.2924 | 78.3802 | 189.2267
0.68083 | 1028780.9643 | 55.7603 | 134.6174
0.68442 | 1033174.488 | 75.4274 | 182.0979
0.69491 | 1060733.142 | 67.0443 | 161.8594
0.7002 | 1055682.6382 | 68.1404 | 164.5054
0.69341 | 1046866.1577 | 57.9863 | 139.9913
0.68092 | 1004793.1968 | 64.8134 | 156.4733
0.67635 | 1001235.7155 | 89.2114 | 215.3753
0.72862 | 1126911.1364 | 98.8614 | 238.6726
0.6928 | 1052290.4475 | 79.2683 | 191.3706
0.69502 | 1041916.647 | 74.9054 | 180.8376
0.70411 | 1064816.4872 | 78.9235 | 190.5383
0.68639 | 1031088.9077 | 83.1528 | 200.7487
0.68614 | 1014671.8044 | 72.191 | 174.2846
0.69152 | 1064455.0484 | 59.1218 | 142.7325
0.68167 | 1017405.9166 | 76.2519 | 184.0884
0.70051 | 1078375.6754 | 91.3085 | 220.4383
0.6718 | 985292.4123 | 66.2863 | 160.0294
0.70383 | 1078607.3683 | 73.0085 | 176.2581
0.68322 | 1028602.1583 | 69.9073 | 168.7711
0.6671 | 975473.5971 | 69.6628 | 168.181
0.68717 | 1040773.2622 | 79.0432 | 190.827
0.68169 | 1014161.2351 | 69.0813 | 166.7769
0.6996 | 1060641.3157 | 85.4617 | 206.3227
0.66901 | 980057.6139 | 78.0291 | 188.3789
0.69211 | 1053738.1397 | 58.4769 | 141.1757
0.70438 | 1079813.48 | 81.6542 | 197.1307
0.71623 | 1109482.0202 | 80.8423 | 195.1705
0.68856 | 1038916.5035 | 73.6571 | 177.8239
0.70134 | 1071778.5493 | 74.6018 | 180.1046
0.68666 | 1045015.9076 | 81.0631 | 195.7035
0.68492 | 1014648.1675 | 70.5758 | 170.385
0.70323 | 1068260.0513 | 64.7283 | 156.2679
0.6877 | 1033611.4221 | 69.7901 | 168.4882
0.68429 | 1053652.1372 | 69.5912 | 168.0079
0.69302 | 1062040.965 | 78.2937 | 189.0178
0.69796 | 1039940.268 | 65.4426 | 157.9924
0.68396 | 1036494.3502 | 77.4224 | 186.9142
0.68025 | 1019042.0024 | 71.368 | 172.2976
0.6908 | 1038243.9007 | 81.6401 | 197.0965
0.67854 | 1008370.4274 | 72.5471 | 175.1443
0.69758 | 1064046.737 | 85.9839 | 207.5836
0.69035 | 1054182.847 | 84.5284 | 204.0697
0.68541 | 1022162.3043 | 65.5013 | 158.1341
0.70757 | 1079646.2359 | 77.9548 | 188.1994
0.65849 | 961107.2874 | 64.1074 | 154.769
0.68285 | 1010901.3365 | 78.8539 | 190.3702
0.70558 | 1078426.3595 | 76.3695 | 184.3722
0.66989 | 979386.583 | 66.076 | 159.5215
0.69604 | 1064946.5387 | 69.1493 | 166.9412
0.66667 | 987133.4837 | 70.5412 | 170.3016
0.68285 | 1014721.5022 | 79.1513 | 191.0882
0.71178 | 1084615.6288 | 81.165 | 195.9496
0.70141 | 1071932.5253 | 58.6975 | 141.7083
0.70627 | 1084554.6506 | 71.3043 | 172.1437
0.69575 | 1049970.9594 | 71.8953 | 173.5706
0.68014 | 1007455.3301 | 74.0124 | 178.6816
0.66772 | 993311.5205 | 63.3879 | 153.032
0.70028 | 1078724.7615 | 71.5276 | 172.6829
0.7017 | 1083259.8615 | 85.4542 | 206.3046
0.68952 | 1024779.6862 | 72.553 | 175.1585
0.70925 | 1085578.056 | 79.88 | 192.8473
0.6782 | 1009986.5238 | 72.6994 | 175.5119
0.69237 | 1047301.1097 | 72.0627 | 173.9746
0.71578 | 1114775.0591 | 77.8185 | 187.8705
0.66913 | 996424.037 | 67.3018 | 162.4808
0.69819 | 1052970.6185 | 68.161 | 164.5553
0.70828 | 1075894.3451 | 91.5417 | 221.0012
0.65904 | 963856.9198 | 79.4435 | 191.7936
0.67489 | 1011256.4531 | 78.1992 | 188.7895
0.69037 | 1032995.6912 | 63.7582 | 153.9258
0.68555 | 1019636.1352 | 57.1257 | 137.9136
0.68426 | 1016886.6419 | 65.2729 | 157.5828
0.6971 | 1054573.1512 | 85.2769 | 205.8766
0.67715 | 986624.461 | 77.6085 | 187.3635
0.6601 | 960002.8129 | 70.9959 | 171.3993
0.69031 | 1045713.6989 | 88.3905 | 213.3934
0.68284 | 1031058.7074 | 64.3401 | 155.3307
0.67566 | 1024287.8166 | 75.7484 | 182.8727
0.67993 | 1014755.8128 | 65.5514 | 158.2552
0.69916 | 1058360.2775 | 70.2905 | 169.6963
0.66714 | 1000342.5116 | 84.5249 | 204.0611
0.68461 | 998027.7934 | 79.1874 | 191.1754
0.72188 | 1121978.9196 | 64.1239 | 154.8087
0.69202 | 1046230.8264 | 68.3287 | 164.9601
0.6692 | 1000384.0232 | 74.6938 | 180.3269
0.69158 | 1030379.5096 | 59.3847 | 143.3673
0.70694 | 1084417.8876 | 74.884 | 180.7861
0.68734 | 1022797.9351 | 74.3523 | 179.5023
0.68545 | 1013121.9534 | 66.8012 | 161.2724
0.73056 | 1135373.1658 | 70.239 | 169.5719
0.66669 | 997092.2393 | 67.5935 | 163.1852
0.7168 | 1093585.5147 | 86.002 | 207.6272
0.68042 | 1012953.1364 | 70.9238 | 171.2252
0.66889 | 995420.265 | 74.3576 | 179.5152
0.67933 | 1006731.3865 | 66.9775 | 161.6981
0.67799 | 1029989.5795 | 77.6134 | 187.3754
0.68789 | 1035252.4004 | 62.0912 | 149.9013
0.7004 | 1051556.3128 | 61.9419 | 149.5411
0.70783 | 1077145.9487 | 79.4181 | 191.7323
0.71759 | 1095335.4207 | 79.661 | 192.3186
0.72711 | 1127449.8092 | 70.0795 | 169.1869
0.68125 | 1016185.1704 | 69.263 | 167.2158
0.69196 | 1053279.1417 | 87.3249 | 210.8209
0.66042 | 988370.3306 | 61.7446 | 149.0646
0.68852 | 1034573.8606 | 65.9977 | 159.3325
0.69387 | 1056412.1473 | 72.5324 | 175.1088
0.69236 | 1045423.628 | 68.2425 | 164.7519
0.69361 | 1040958.041 | 76.6745 | 185.1087
0.72507 | 1123371.8561 | 79.1539 | 191.0943
0.66586 | 993993.5001 | 72.8746 | 175.9348
0.6781 | 1006748.052 | 74.1826 | 179.0926
0.69626 | 1072616.5097 | 73.3919 | 177.1837
0.70493 | 1079029.6295 | 79.8843 | 192.8577
0.66621 | 983167.1164 | 56.6814 | 136.8409
0.69059 | 1053636.8701 | 70.8445 | 171.0336
0.68819 | 1038313.3564 | 77.0814 | 186.0909
0.67586 | 982524.3235 | 67.5028 | 162.9661
0.68689 | 1025926.106 | 61.6214 | 148.7671
0.69233 | 1047896.5151 | 70.282 | 169.6757
0.6986 | 1064829.9991 | 77.8725 | 188.001
0.68382 | 1027447.1041 | 64.0229 | 154.565
0.68852 | 1036142.3576 | 78.5374 | 189.6059
0.69089 | 1043966.2667 | 78.8587 | 190.3818
0.69942 | 1077175.1278 | 76.9533 | 185.7817
0.69215 | 1042494.351 | 68.0718 | 164.3398
0.68456 | 1026285.6393 | 79.8542 | 192.785
0.67016 | 999938.7243 | 65.5353 | 158.2161
0.68583 | 1039470.7089 | 83.8981 | 202.5479
0.6897 | 1029221.2105 | 76.2399 | 184.0593
0.68268 | 1004973.4745 | 58.6753 | 141.6546
0.73856 | 1172949.9831 | 84.5924 | 204.2242
0.68253 | 1025873.1453 | 73.3743 | 177.1413
0.67067 | 992773.3066 | 70.5206 | 170.2517
0.6928 | 1055706.1625 | 72.5957 | 175.2615
0.70047 | 1064986.2859 | 75.9685 | 183.4041
0.6853 | 1032371.9299 | 78.6193 | 189.8037
0.6958 | 1050786.2893 | 83.5958 | 201.8182
0.70783 | 1078182.0816 | 68.843 | 166.2017
0.66552 | 977958.6565 | 73.424 | 177.2612
0.69128 | 1029964.961 | 76.1195 | 183.7687
0.65715 | 952398.5755 | 74.0428 | 178.7551
0.70653 | 1082104.9073 | 74.7672 | 180.5039
0.70128 | 1072988.0124 | 72.8917 | 175.9761
0.70791 | 1097637.1628 | 80.92 | 195.3581
0.67707 | 1007669.9949 | 67.2638 | 162.3892
0.71562 | 1097711.5611 | 67.3733 | 162.6535
0.69227 | 1035074.9436 | 72.991 | 176.2158
0.70154 | 1080416.3414 | 69.6041 | 168.0391
0.68724 | 1035237.839 | 57.8569 | 139.679
0.68287 | 1021630.1478 | 59.631 | 143.9621
0.6893 | 1043579.4167 | 77.3479 | 186.7343
0.67717 | 1013643.9917 | 87.3429 | 210.8644
0.66958 | 990530.8515 | 70.4764 | 170.1451
0.70569 | 1091452.247 | 65.903 | 159.1039
0.69535 | 1046970.1054 | 80.1996 | 193.619
0.67884 | 1016626.6764 | 67.6843 | 163.4043
0.70591 | 1087831.0384 | 95.7263 | 231.1037
0.66318 | 977552.5685 | 73.6518 | 177.8111
0.69805 | 1045054.9642 | 67.6622 | 163.351
0.66307 | 969756.9316 | 71.3962 | 172.3656
0.70938 | 1091073.6717 | 71.4192 | 172.4213
0.67242 | 1008693.6524 | 78.2876 | 189.003
0.69948 | 1063861.6823 | 62.7496 | 151.491
0.68701 | 1042806.3966 | 82.4517 | 199.056
0.68077 | 1020470.9542 | 78.2306 | 188.8654
0.68084 | 1005454.2552 | 80.2181 | 193.6636
0.68328 | 1026266.4017 | 83.1415 | 200.7213
0.69115 | 1041445.8628 | 68.1025 | 164.4139
0.68947 | 1034204.204 | 73.2258 | 176.7827
0.64909 | 952318.8901 | 79.6959 | 192.4029
0.6761 | 1001488.7287 | 68.7815 | 166.0533
0.70385 | 1073808.4417 | 60.8083 | 146.8043
0.69222 | 1032677.7035 | 67.4247 | 162.7776
0.68008 | 1026388.0597 | 68.2103 | 164.6742
0.72434 | 1119098.0607 | 83.9237 | 202.6097
0.67178 | 999162.029 | 70.857 | 171.064
0.69986 | 1073016.8149 | 78.3298 | 189.1048
0.69877 | 1060717.1779 | 87.0847 | 210.241
0.70579 | 1080995.0343 | 69.222 | 167.1166
0.67698 | 1004136.7211 | 77.4086 | 186.881
0.68093 | 1017867.7396 | 55.1611 | 133.1706
0.67611 | 1002120.529 | 61.1511 | 147.6318
0.66651 | 988972.3518 | 63.2253 | 152.6395
0.68556 | 1032537.6645 | 69.9352 | 168.8385
0.69777 | 1055359.624 | 72.0551 | 173.9565
0.68944 | 1028177.406 | 72.8018 | 175.7592
0.70087 | 1059164.8451 | 68.3005 | 164.892
0.68018 | 1021983.4631 | 59.8878 | 144.5818
0.67318 | 996090.1825 | 75.5962 | 182.5053
0.70532 | 1065142.9763 | 55.7266 | 134.536
0.68209 | 1034424.9799 | 80.7789 | 195.0175
0.69071 | 1043013.1257 | 85.1225 | 205.5038
0.68912 | 1041195.9291 | 77.1912 | 186.356
0.67071 | 1010762.5397 | 75.2997 | 181.7896
0.68126 | 1024456.2473 | 76.8269 | 185.4766
0.6934 | 1048459.6519 | 77.25 | 186.498
0.69247 | 1043648.627 | 84.9581 | 205.1069
0.68622 | 1037464.6401 | 88.2157 | 212.9715
0.68964 | 1037397.6246 | 71.6634 | 173.0107
0.69068 | 1033501.6126 | 74.7674 | 180.5045
0.67976 | 1011577.7355 | 79.2658 | 191.3645
0.70474 | 1082758.3612 | 46.2903 | 111.7547
0.68503 | 1020442.6364 | 72.7198 | 175.5611
0.67934 | 1020323.4809 | 76.7331 | 185.25
0.66334 | 972920.8863 | 76.5449 | 184.7957
0.69221 | 1040688.9156 | 69.8104 | 168.5373
0.6829 | 1023190.2086 | 67.7169 | 163.483
0.67728 | 1005266.3118 | 65.1925 | 157.3885
0.68385 | 1023356.9554 | 74.773 | 180.5179
0.71414 | 1108536.0922 | 87.5019 | 211.2483
0.71506 | 1102893.4149 | 85.3774 | 206.1193
0.69178 | 1045355.6803 | 70.167 | 169.398
0.68898 | 1038772.7011 | 68.1412 | 164.5073
0.685 | 1027432.3363 | 79.8923 | 192.877
0.65507 | 951927.5038 | 75.8383 | 183.09
0.7111 | 1080264.0766 | 74.1605 | 179.0393
0.68503 | 1013597.8069 | 70.491 | 170.1803
0.71075 | 1079945.5925 | 64.7441 | 156.3061
0.6797 | 1027801.6364 | 80.5432 | 194.4486
0.70185 | 1074996.9502 | 76.5311 | 184.7623
0.72979 | 1134734.6952 | 72.1999 | 174.306
0.68225 | 1018685.7405 | 84.1688 | 203.2015
0.68216 | 1000643.9557 | 74.9413 | 180.9244
0.6988 | 1067548.7763 | 86.1165 | 207.9036
0.68219 | 1026341.3992 | 73.4446 | 177.3111
0.69005 | 1030207.5977 | 71.1244 | 171.7094
0.69415 | 1074022.3321 | 87.2195 | 210.5664
0.71165 | 1081585.9767 | 87.1909 | 210.4974
0.68325 | 1014181.1683 | 85.155 | 205.5824
0.65781 | 986056.5562 | 79.8438 | 192.76
0.69225 | 1049590.0475 | 72.2329 | 174.3857
0.71271 | 1108966.648 | 68.3416 | 164.9913
0.72568 | 1118007.1037 | 71.5385 | 172.7093
0.68656 | 1037580.286 | 85.1912 | 205.6699
0.71236 | 1092114.7871 | 71.1883 | 171.8637
0.71427 | 1099266.6057 | 80.3183 | 193.9055
0.70284 | 1067066.5877 | 85.5836 | 206.617
0.67993 | 1018496.5761 | 87.4041 | 211.0122
0.68231 | 1020558.4655 | 85.8085 | 207.1601
0.6799 | 1015460.7824 | 72.9106 | 176.0219
0.69503 | 1059144.3967 | 74.3918 | 179.5977
0.69662 | 1054401.0283 | 91.2148 | 220.2121
0.67232 | 1006014.7384 | 65.7223 | 158.6676
0.69145 | 1053026.6819 | 66.8899 | 161.4864
0.70867 | 1085251.8524 | 89.3558 | 215.724
0.70033 | 1066965.4727 | 77.8353 | 187.9111
0.67803 | 1016326.5325 | 73.1744 | 176.6586
0.68353 | 1026098.123 | 63.9661 | 154.4278
0.68861 | 1049001.6253 | 74.3668 | 179.5374
0.71362 | 1094084.7853 | 87.9544 | 212.3406
0.67979 | 1029884.9247 | 75.699 | 182.7536
0.69533 | 1051829.9273 | 72.0662 | 173.9832
0.70956 | 1085016.0872 | 74.6483 | 180.2169
0.67557 | 1019779.3863 | 83.558 | 201.7268
0.69611 | 1048590.8756 | 83.3837 | 201.3061
0.69331 | 1033800.3111 | 71.0376 | 171.4999
0.67607 | 994986.2552 | 73.5358 | 177.5312
0.68454 | 1030164.0335 | 68.0144 | 164.2012
0.68033 | 1019236.2021 | 71.3732 | 172.3102
0.68851 | 1041328.2559 | 66.4296 | 160.3752
0.67325 | 1015991.4344 | 73.6693 | 177.8534
0.71061 | 1074780.8529 | 75.5726 | 182.4485
0.71696 | 1091032.4434 | 70.6612 | 170.5911
0.71053 | 1079651.8173 | 68.6748 | 165.7957
0.69317 | 1057211.6421 | 64.9886 | 156.8963
0.68277 | 1017091.6987 | 63.4745 | 153.241
0.6916 | 1064475.25 | 78.6346 | 189.8407
0.68101 | 1020159.4934 | 65.7957 | 158.8449
0.68886 | 1035355.8655 | 64.0435 | 154.6148
0.66284 | 972104.2296 | 93.7437 | 226.3174
0.69229 | 1030952.3107 | 67.3051 | 162.4888
0.70026 | 1061404.3249 | 83.9029 | 202.5595
0.68514 | 1021515.7869 | 74.4974 | 179.8526
0.69077 | 1037102.2275 | 76.7039 | 185.1796
0.70857 | 1076549.8586 | 75.9149 | 183.2747
0.69188 | 1046679.7639 | 66.4453 | 160.4131
0.69481 | 1044999.0812 | 81.6978 | 197.2359
0.69258 | 1044337.0907 | 80.3174 | 193.9034
0.70422 | 1077648.5291 | 71.0574 | 171.5477
0.66532 | 978524.6176 | 76.7563 | 185.3061
0.70623 | 1096661.6459 | 75.8052 | 183.01
0.6685 | 1000767.8246 | 82.8673 | 200.0594
0.69254 | 1039813.5094 | 79.9745 | 193.0754
0.69908 | 1075048.2159 | 59.0641 | 142.5933
0.67018 | 983418.9047 | 79.7087 | 192.4339
0.72591 | 1132297.4514 | 91.015 | 219.7296
0.70452 | 1070599.8887 | 79.3068 | 191.4635
0.69193 | 1049751.9975 | 64.6462 | 156.0696
0.70035 | 1068372.7599 | 81.254 | 196.1645
0.69802 | 1038876.326 | 76.0372 | 183.5701
0.68343 | 1020311.1174 | 69.3231 | 167.3609
0.70972 | 1083576.3672 | 84.5138 | 204.0343
0.68679 | 1033371.3285 | 75.1814 | 181.504
0.6796 | 1019115.5368 | 81.4573 | 196.6554
0.71484 | 1074344.3148 | 76.7831 | 185.3709
0.6863 | 1037711.1011 | 68.3377 | 164.9818
0.69072 | 1049133.3811 | 88.6266 | 213.9636
0.70487 | 1075062.049 | 76.208 | 183.9825
0.70074 | 1074876.278 | 74.882 | 180.7811
0.69672 | 1078097.7669 | 73.7183 | 177.9718
0.69454 | 1060112.733 | 65.039 | 157.018
0.69284 | 1061054.7091 | 78.2417 | 188.8922
0.67272 | 995055.7754 | 70.4682 | 170.1253
0.68238 | 1033983.655 | 85.4165 | 206.2138
0.68808 | 1027420.2325 | 87.1259 | 210.3406
0.69087 | 1035983.9871 | 78.2588 | 188.9334
0.68984 | 1033068.9461 | 67.6821 | 163.3991
0.68955 | 1022881.425 | 75.7858 | 182.9632
0.70788 | 1078240.881 | 56.2083 | 135.6989
0.69625 | 1042088.2524 | 79.7217 | 192.4652
0.69998 | 1073747.6453 | 74.8842 | 180.7864
0.68434 | 1003171.0552 | 82.6336 | 199.4952
0.69438 | 1064379.6178 | 71.1763 | 171.8347
0.68324 | 1027936.6578 | 67.9966 | 164.1584
0.68457 | 1018386.6842 | 92.409 | 223.095
0.68545 | 1034063.1817 | 73.6455 | 177.7961
0.68935 | 1035593.5378 | 81.6027 | 197.0064
0.69362 | 1039111.4836 | 82.9262 | 200.2015
0.68521 | 1015490.6029 | 72.3483 | 174.6641
0.6993 | 1071071.267 | 72.9678 | 176.1598
0.68705 | 1029005.5419 | 73.1139 | 176.5127
0.71422 | 1101879.8746 | 83.8031 | 202.3187
0.69724 | 1063864.8778 | 89.0308 | 214.9393
0.70198 | 1054038.5454 | 69.7222 | 168.3243
0.71058 | 1088749.849 | 79.963 | 193.0478
0.69149 | 1033584.4596 | 67.5546 | 163.0912
0.69817 | 1069281.8528 | 76.5258 | 184.7495
0.67366 | 1002903.1531 | 79.2971 | 191.4402
0.69549 | 1048755.6898 | 83.471 | 201.5168
0.68313 | 1020844.8016 | 71.1921 | 171.873
0.67434 | 1011035.8658 | 69.6185 | 168.0739
0.68543 | 1024833.5581 | 67.5753 | 163.1413
0.70031 | 1050907.6514 | 63.9431 | 154.3722
0.69072 | 1041961.8422 | 70.047 | 169.1085
0.70248 | 1095984.1024 | 84.7482 | 204.6003
0.68435 | 1018952.7408 | 70.9334 | 171.2484
0.66298 | 969709.5822 | 67.9174 | 163.9672
0.678 | 1011899.1553 | 77.8896 | 188.042
0.69718 | 1055742.0827 | 66.5113 | 160.5725
0.66759 | 993390.2203 | 67.2525 | 162.3619
0.69724 | 1039393.8933 | 84.0326 | 202.8727
0.68899 | 1026598.0498 | 61.4737 | 148.4107
0.68425 | 1019873.3242 | 71.0086 | 171.4299
0.70012 | 1060216.4676 | 75.6651 | 182.6718
0.68051 | 1000896.1894 | 66.4742 | 160.483
0.67857 | 1015687.4211 | 67.6487 | 163.3185
0.69202 | 1041732.8231 | 66.1181 | 159.6231
0.68421 | 1022347.694 | 69.3824 | 167.5039
0.69587 | 1056327.1286 | 84.0285 | 202.8627
0.67025 | 978240.2578 | 69.5326 | 167.8665
0.68318 | 1022601.8292 | 71.3093 | 172.1558
0.69715 | 1044168.4972 | 71.3548 | 172.2658
0.71373 | 1104529.5709 | 90.2991 | 218.0013
0.70479 | 1088407.6082 | 65.8036 | 158.8639
0.70251 | 1077940.0433 | 83.3492 | 201.2226
0.69961 | 1059063.6146 | 70.2915 | 169.6987
0.70138 | 1073948.4395 | 87.842 | 212.0693
0.6807 | 1014810.4822 | 69.2261 | 167.1267
0.6815 | 1019882.5824 | 61.7425 | 149.0595
0.70755 | 1089083.7868 | 75.1414 | 181.4074
0.70137 | 1044196.9815 | 67.287 | 162.4452
0.70003 | 1049733.6287 | 78.2596 | 188.9353
0.70788 | 1083663.6223 | 78.2105 | 188.8169
0.69315 | 1040837.0022 | 76.7165 | 185.2099
0.68274 | 1021758.8926 | 65.7321 | 158.6913
0.69557 | 1055383.5466 | 77.5175 | 187.1437
0.67903 | 993768.9865 | 67.9574 | 164.0637
0.67352 | 994342.2687 | 94.3678 | 227.8239
0.70186 | 1066892.9588 | 79.3314 | 191.5229
0.70522 | 1069096.0514 | 75.7477 | 182.871
0.67758 | 1011516.6203 | 81.0907 | 195.7703
0.7151 | 1111468.8972 | 77.662 | 187.4927
0.68653 | 1018299.3454 | 72.408 | 174.8085
0.6855 | 1035615.2055 | 68.2788 | 164.8395
0.7117 | 1098769.7353 | 78.8755 | 190.4223
0.6977 | 1065634.3878 | 70.9213 | 171.2191
0.69994 | 1076257.1381 | 59.0581 | 142.5789
0.69489 | 1047309.6312 | 80.2783 | 193.8089
0.69743 | 1068202.486 | 61.0479 | 147.3827
0.69691 | 1068713.1075 | 78.3455 | 189.1429
0.69065 | 1029394.3445 | 63.5098 | 153.3262
0.68925 | 1039542.0757 | 71.8524 | 173.4671
0.70982 | 1091172.8555 | 83.5034 | 201.595
0.67909 | 1024358.495 | 62.2334 | 150.2448
0.66824 | 1008959.3663 | 77.1381 | 186.2278
0.68819 | 1041721.9602 | 75.6349 | 182.5987
0.69294 | 1066667.2045 | 83.6164 | 201.8679
0.69707 | 1076810.658 | 67.849 | 163.8021
0.66839 | 980275.4226 | 78.1197 | 188.5976
0.67417 | 1005821.6364 | 71.0771 | 171.5954
0.69591 | 1054495.292 | 72.1888 | 174.2792
0.68504 | 1031994.9542 | 69.9516 | 168.8781
0.66764 | 990729.9057 | 70.6957 | 170.6746
0.69847 | 1069563.405 | 74.7537 | 180.4713
0.71572 | 1102266.0265 | 63.5397 | 153.3985
0.69196 | 1052232.4722 | 72.4451 | 174.8978
0.66486 | 980707.2853 | 75.3597 | 181.9344
0.69679 | 1051253.5206 | 73.6514 | 177.8102
0.67872 | 1011039.936 | 65.4278 | 157.9566
0.69412 | 1060553.1686 | 70.9471 | 171.2815
0.67624 | 997870.0278 | 70.2995 | 169.7181
0.71125 | 1093770.7649 | 78.5657 | 189.6743
0.70953 | 1077626.275 | 84.0719 | 202.9674
0.70145 | 1080148.6511 | 75.1731 | 181.484
0.67396 | 1023963.6693 | 77.6218 | 187.3957
0.695 | 1066501.5226 | 63.3104 | 152.8447
0.68443 | 1021813.4339 | 85.6476 | 206.7716
0.69338 | 1044040.9658 | 69.9175 | 168.7957
0.6834 | 1033812.2616 | 70.2437 | 169.5833
0.69252 | 1055496.2411 | 66.9432 | 161.6151
0.66815 | 1000443.3857 | 61.7763 | 149.1412
0.70647 | 1082059.2723 | 81.5828 | 196.9583
0.68292 | 1013105.9766 | 67.0475 | 161.867
0.70836 | 1087103.2916 | 69.3607 | 167.4515
0.69603 | 1054435.7386 | 66.5392 | 160.6398
0.69027 | 1062556.0242 | 89.8187 | 216.8416
0.69457 | 1058124.6746 | 79.5051 | 191.9423
0.68724 | 1018127.5647 | 83.5114 | 201.6143
0.66787 | 991798.1606 | 80.666 | 194.7451
0.68461 | 1037536.3085 | 83.891 | 202.5307
0.68794 | 1044256.0996 | 74.1017 | 178.8972
0.71408 | 1110309.5878 | 68.2935 | 164.8751
0.68663 | 1034430.7452 | 69.9547 | 168.8856
0.70224 | 1069648.8213 | 68.8421 | 166.1994
0.69644 | 1046128.0754 | 78.3928 | 189.2571
0.68039 | 1031835.9767 | 58.2471 | 140.6208
0.70847 | 1092575.8707 | 65.712 | 158.6427
0.69716 | 1053762.7274 | 77.2219 | 186.4302
0.68845 | 1037947.3245 | 64.242 | 155.094
0.69034 | 1031032.4434 | 59.5291 | 143.7159
0.67447 | 1003087.7612 | 61.3592 | 148.1343
0.67814 | 1022797.2212 | 77.0544 | 186.0258
0.71527 | 1111478.3387 | 66.2133 | 159.8531
0.68918 | 1048703.4195 | 64.8388 | 156.5347
0.69837 | 1071198.3309 | 69.3234 | 167.3616
0.70349 | 1078967.2068 | 80.0431 | 193.2412
0.68622 | 1030041.5425 | 65.9286 | 159.1658
0.70537 | 1076103.5324 | 58.5951 | 141.4611
0.67289 | 1012218.0132 | 83.2687 | 201.0284
0.66156 | 975416.1096 | 80.0894 | 193.3529
0.68499 | 1032705.8577 | 66.2177 | 159.8636
0.66193 | 976242.8551 | 66.6568 | 160.9239
0.6893 | 1036586.1818 | 66.0511 | 159.4614
0.6909 | 1046926.6737 | 79.082 | 190.9208
0.70211 | 1065322.4357 | 68.5762 | 165.5575
0.69191 | 1041414.2092 | 84.8453 | 204.8348
0.6796 | 1021353.2317 | 80.0501 | 193.2581
0.68472 | 1032424.3636 | 75.3389 | 181.8842
0.69879 | 1026652.0735 | 51.9063 | 125.313
0.70007 | 1073287.6599 | 84.8646 | 204.8813
0.69913 | 1072339.7247 | 80.0478 | 193.2526
0.67786 | 1008264.788 | 79.2271 | 191.2712
0.69802 | 1061224.2412 | 62.7252 | 151.4321
0.6886 | 1043472.4301 | 77.418 | 186.9037
0.7254 | 1123530.3917 | 82.4992 | 199.1706
0.71128 | 1069748.3279 | 86.2814 | 208.3018
0.70095 | 1053913.0974 | 75.1265 | 181.3715
0.69656 | 1052719.0923 | 73.7781 | 178.1161
0.68414 | 1023962.1479 | 71.3047 | 172.1447
0.71278 | 1102160.7198 | 78.999 | 190.7205
0.6539 | 961546.8313 | 82.8923 | 200.1198
0.68086 | 1016349.9892 | 74.5864 | 180.0675
0.69823 | 1056200.0091 | 68.367 | 165.0525
0.68957 | 1029081.0046 | 55.0698 | 132.9503
0.68853 | 1030974.6391 | 74.8173 | 180.6249
0.71608 | 1100252.158 | 83.4614 | 201.4936
0.67211 | 997801.7768 | 59.1727 | 142.8556
0.74323 | 1170421.4227 | 75.0235 | 181.1228
0.67807 | 1008467.4476 | 83.8384 | 202.4037
0.68357 | 1021177.4286 | 72.928 | 176.0638
0.68708 | 1034005.0376 | 66.0992 | 159.5776
0.69721 | 1064664.4105 | 70.9071 | 171.1849
0.66764 | 989034.5669 | 70.5746 | 170.3822
0.69066 | 1034563.2824 | 63.8279 | 154.0942
0.67817 | 1006758.5418 | 75.8431 | 183.1015
0.67985 | 1022803.3746 | 75.1719 | 181.4809
0.67333 | 1008829.9652 | 76.5456 | 184.7975
0.67979 | 1003851.7462 | 73.7854 | 178.1336
0.70578 | 1067235.4205 | 55.8001 | 134.7133
0.70576 | 1096030.0785 | 68.2096 | 164.6725
0.70689 | 1072443.4345 | 66.6359 | 160.8733
0.69262 | 1040386.6871 | 61.9724 | 149.6147
0.6596 | 963484.111 | 76.2324 | 184.0413
0.68573 | 1022916.2675 | 63.5631 | 153.455
0.71547 | 1098176.624 | 78.4169 | 189.3152
0.70191 | 1063867.6609 | 66.5713 | 160.7173
0.71207 | 1098211.9913 | 65.7176 | 158.6563
0.69883 | 1048218.6383 | 63.9139 | 154.3019
0.68613 | 1038555.5308 | 73.8915 | 178.3897
0.67987 | 993357.3595 | 74.6464 | 180.2122
0.70394 | 1062649.2139 | 80.4067 | 194.119
0.69498 | 1067026.2496 | 65.4209 | 157.9401
0.70133 | 1064849.5579 | 75.865 | 183.1543
0.6968 | 1074011.857 | 80.44 | 194.1993
0.72426 | 1122457.3097 | 71.9631 | 173.7344
0.70687 | 1066559.1288 | 62.2509 | 150.287
0.67948 | 998059.7785 | 62.5886 | 151.1023
0.69642 | 1056327.3536 | 62.5793 | 151.0798
0.70527 | 1062419.3824 | 83.3864 | 201.3127
0.69425 | 1044286.0512 | 69.8463 | 168.6239
0.67712 | 1013700.0738 | 73.0831 | 176.4382
0.66282 | 974745.474 | 76.7485 | 185.2872
0.68256 | 1019022.5728 | 69.3236 | 167.3621
0.66622 | 982158.8782 | 75.3175 | 181.8325
0.68966 | 1045307.7436 | 62.5329 | 150.9677
0.70023 | 1061843.899 | 74.0271 | 178.7172
0.68022 | 1023446.2681 | 62.0751 | 149.8626
0.68168 | 1019450.0707 | 67.59 | 163.1766
0.68331 | 1033338.9619 | 68.3476 | 165.0057
0.70519 | 1079322.8732 | 79.8761 | 192.838
0.72944 | 1149426.8731 | 80.5052 | 194.3567
0.70292 | 1066101.669 | 79.1527 | 191.0914
0.68971 | 1046189.0784 | 60.4588 | 145.9605
0.66129 | 956283.5924 | 88.4973 | 213.6513
0.70106 | 1071081.2597 | 60.8953 | 147.0142
0.6976 | 1068253.9847 | 81.566 | 196.9176
0.68961 | 1048488.0518 | 79.3478 | 191.5626
0.69169 | 1050461.7334 | 75.8829 | 183.1976
0.69773 | 1079469.8421 | 73.7057 | 177.9412
0.71596 | 1099623.4459 | 80.155 | 193.5114
0.70757 | 1093372.0295 | 80.6442 | 194.6923
0.68564 | 1018580.8623 | 63.3022 | 152.825
0.67788 | 1004305.2598 | 78.9137 | 190.5146
0.68263 | 1028432.8773 | 62.1573 | 150.0611
0.69062 | 1051272.4986 | 74.3108 | 179.4022
0.69081 | 1046252.602 | 77.3764 | 186.8031
0.69887 | 1045176.6236 | 79.6584 | 192.3123
0.69393 | 1042431.1076 | 71.566 | 172.7757
0.6743 | 989433.7684 | 66.3178 | 160.1054
0.67333 | 997905.7406 | 72.349 | 174.6658
0.68794 | 1022756.1981 | 67.1917 | 162.2151
0.69595 | 1049482.8298 | 73.6157 | 177.7241
0.66284 | 967653.3713 | 69.5487 | 167.9055
0.67063 | 1003614.077 | 78.1008 | 188.5521
0.67564 | 999771.2946 | 66.8773 | 161.4561
0.65957 | 983895.2715 | 65.5191 | 158.1772
0.68593 | 1026396.2568 | 73.0477 | 176.3528
0.70053 | 1072120.0856 | 61.0008 | 147.269
0.68833 | 1038527.4503 | 68.1896 | 164.6242
0.70874 | 1088662.2918 | 71.3256 | 172.1953
0.67172 | 1003070.7465 | 72.8979 | 175.9911
0.69586 | 1054640.1762 | 68.4978 | 165.3682
0.71249 | 1091215.5307 | 76.5261 | 184.7504
0.69273 | 1053928.8639 | 82.6432 | 199.5184
0.70264 | 1062642.5024 | 73.8296 | 178.2404
0.68041 | 1013108.9039 | 67.2947 | 162.4637
0.69925 | 1072231.4751 | 61.1185 | 147.5532
0.68394 | 1023883.0882 | 63.7891 | 154.0006
0.6936 | 1053672.8524 | 75.9056 | 183.2523
0.68168 | 1022173.8296 | 76.6079 | 184.9477
0.68976 | 1034581.9368 | 59.7839 | 144.331
0.69791 | 1059190.2861 | 66.6586 | 160.928
0.70639 | 1071675.4879 | 70.446 | 170.0716
0.68823 | 1046005.6242 | 70.5732 | 170.3789
0.72046 | 1114376.1031 | 88.7349 | 214.2251
0.68297 | 1033549.2733 | 66.4683 | 160.4686
0.66575 | 978608.5552 | 58.5885 | 141.4453
0.66151 | 974643.9789 | 68.4143 | 165.1668
0.68141 | 1009594.4913 | 63.1328 | 152.4161
0.68094 | 1026849.3766 | 64.2557 | 155.127
0.70689 | 1057963.2463 | 71.3012 | 172.1362
0.68827 | 1034586.7361 | 69.7986 | 168.5087
0.67269 | 993038.0026 | 60.9046 | 147.0367
0.703 | 1062242.5152 | 75.4993 | 182.2714
0.69107 | 1069570.0987 | 75.5544 | 182.4044
0.71614 | 1093771.0011 | 63.0125 | 152.1257
0.68499 | 1047282.0241 | 83.0501 | 200.5006
0.67519 | 1018570.2405 | 68.1804 | 164.6021
0.69908 | 1050311.338 | 86.1912 | 208.0838
0.6964 | 1055845.1501 | 85.4659 | 206.3329
0.69019 | 1039403.6009 | 71.7449 | 173.2076
0.68352 | 1018091.8114 | 67.3182 | 162.5205
0.70627 | 1070121.311 | 80.7054 | 194.8401
0.68659 | 1033450.2474 | 69.8921 | 168.7345
0.67044 | 1006642.204 | 65.9475 | 159.2114
0.70675 | 1083505.6239 | 74.8408 | 180.6816
0.69696 | 1043608.578 | 73.0976 | 176.4733
0.69274 | 1063043.7303 | 83.3565 | 201.2405
0.70091 | 1060628.9054 | 78.972 | 190.6552
0.68792 | 1044482.166 | 82.4747 | 199.1116
0.69365 | 1051374.4561 | 79.102 | 190.9692
0.72073 | 1124988.3973 | 79.7456 | 192.523
0.6842 | 1034861.6464 | 73.8368 | 178.2579
0.71721 | 1106950.0231 | 93.0108 | 224.5481
0.67794 | 1005766.6826 | 85.1097 | 205.473
0.69018 | 1037506.0559 | 81.6554 | 197.1336
0.69186 | 1056908.1559 | 79.7251 | 192.4733
0.71006 | 1091208.65 | 76.787 | 185.3803
0.68363 | 1033931.3408 | 68.3547 | 165.0229
0.68873 | 1047087.0251 | 82.5827 | 199.3722
0.69017 | 1042549.5084 | 53.3684 | 128.8427
0.68678 | 1039154.7245 | 68.3847 | 165.0953
0.68531 | 1018770.301 | 70.8214 | 170.9779
0.68521 | 1034133.6157 | 75.4097 | 182.0551
0.67238 | 987984.3966 | 73.4197 | 177.2509
0.68323 | 1020792.7567 | 65.6225 | 158.4267
0.69008 | 1030928.2937 | 61.6358 | 148.8019
0.68904 | 1041612.4476 | 66.3965 | 160.2952
0.6993 | 1078650.9638 | 71.6234 | 172.9141
0.7101 | 1097229.3548 | 57.0187 | 137.6554
0.72527 | 1128157.0123 | 79.4699 | 191.8574
0.6811 | 1025471.9428 | 68.0467 | 164.2792
0.6919 | 1043407.4071 | 70.0244 | 169.0538
0.67924 | 1023044.0389 | 69.9378 | 168.8447
0.67427 | 990032.7242 | 71.6785 | 173.0472
0.69876 | 1041767.2452 | 80.9846 | 195.5141
0.71115 | 1087417.9112 | 65.156 | 157.3004
0.67909 | 1009314.4995 | 69.0135 | 166.6133
0.69295 | 1051262.3907 | 77.507 | 187.1184
0.70487 | 1058876.9055 | 59.8764 | 144.5545
0.70059 | 1066444.1278 | 81.0802 | 195.745
0.68381 | 1023552.5346 | 85.8719 | 207.313
0.6809 | 1026964.1657 | 75.3614 | 181.9385
0.6738 | 1002278.3392 | 85.3361 | 206.0195
0.70939 | 1077599.2424 | 72.368 | 174.7118
0.70912 | 1078581.4172 | 72.4625 | 174.9399
0.6831 | 1041660.1151 | 62.9135 | 151.8866
0.68703 | 1032731.4824 | 57.7348 | 139.3842
0.67872 | 1005175.7222 | 66.9339 | 161.5926
0.67269 | 1006887.5187 | 73.0699 | 176.4063
0.68308 | 1043603.4726 | 74.5821 | 180.0571
0.67422 | 1002146.7098 | 65.285 | 157.6119
0.69297 | 1044265.0829 | 75.2961 | 181.7808
0.69186 | 1044655.3626 | 66.3189 | 160.1079
0.69483 | 1051019.1432 | 70.2379 | 169.5694
0.70824 | 1081396.3988 | 67.0637 | 161.9062
0.71122 | 1098857.007 | 78.9344 | 190.5645
0.68897 | 1037222.9015 | 73.4927 | 177.4271
0.67986 | 1007650.2762 | 77.3507 | 186.7412
0.70397 | 1061645.1747 | 60.1093 | 145.1168
0.67394 | 1006638.4398 | 71.1587 | 171.7923
0.66712 | 998698.556 | 87.3513 | 210.8848
0.69771 | 1053449.9528 | 60.1804 | 145.2884
0.71849 | 1102119.095 | 73.0782 | 176.4263
0.70014 | 1046469.3775 | 57.4072 | 138.5932
0.7052 | 1078496.5453 | 65.1666 | 157.3261
0.67936 | 1007767.3694 | 70.6709 | 170.6146
0.68277 | 1040599.2784 | 64.9421 | 156.7842
0.71229 | 1108419.8702 | 75.9664 | 183.399
0.68093 | 1030282.9517 | 85.2913 | 205.9113
0.6782 | 1005413.4802 | 65.1135 | 157.198
0.69383 | 1036556.4081 | 61.6371 | 148.8052
0.69701 | 1058865.0316 | 80.9177 | 195.3525
0.68332 | 1031755.3781 | 71.9331 | 173.6618
0.66917 | 985346.8168 | 81.5093 | 196.7808
0.67321 | 1001695.6167 | 69.0915 | 166.8016
0.70868 | 1086694.5307 | 76.618 | 184.9722
0.69047 | 1034165.0525 | 84.4298 | 203.8315
0.70752 | 1074553.5575 | 66.7074 | 161.0458
0.69984 | 1057283.7081 | 80.4702 | 194.2723
0.67951 | 992191.4474 | 72.9588 | 176.1381
0.68672 | 1037786.6441 | 80.4951 | 194.3324
0.70399 | 1073959.3866 | 67.0045 | 161.7632
0.69476 | 1052170.3905 | 72.1488 | 174.1826
0.67896 | 1018822.6675 | 63.9394 | 154.3635
0.70246 | 1057634.5574 | 86.375 | 208.5277
0.68928 | 1032910.7087 | 59.7289 | 144.1982
0.69652 | 1059540.1286 | 60.6316 | 146.3777
0.69498 | 1056457.9511 | 84.1458 | 203.1458
0.6537 | 957632.4897 | 66.2089 | 159.8425
0.66982 | 995569.7109 | 74.2659 | 179.2936
0.6872 | 1038973.6209 | 68.331 | 164.9657
0.68072 | 1025929.2732 | 76.9981 | 185.8897
0.69259 | 1034937.1543 | 65.6327 | 158.4513
0.70923 | 1094221.0986 | 73.5014 | 177.448
0.70369 | 1082306.1091 | 75.0829 | 181.2663
0.6719 | 993997.4778 | 68.6938 | 165.8416
0.67232 | 999434.8082 | 67.1849 | 162.1988
0.68918 | 1039610.2619 | 64.2769 | 155.1781
0.70419 | 1067303.2586 | 77.4337 | 186.9414
0.71522 | 1097500.4858 | 67.3977 | 162.7124
0.66452 | 964577.6296 | 65.603 | 158.3796
0.70869 | 1080188.3408 | 70.7748 | 170.8655
0.69966 | 1067919.7245 | 72.716 | 175.552
0.69891 | 1061442.7745 | 77.5902 | 187.3193
0.70307 | 1069237.2248 | 72.4523 | 174.9153
0.69096 | 1042078.0972 | 81.9231 | 197.7798
0.6627 | 973220.9521 | 75.0281 | 181.1338
0.69975 | 1041312.3624 | 80.8346 | 195.1519
0.69194 | 1026994.3966 | 77.4598 | 187.0045
0.68169 | 1021563.1852 | 63.8552 | 154.1601
0.68044 | 1011215.152 | 88.1476 | 212.8072
0.68748 | 1041108.3723 | 80.6634 | 194.7386
0.70111 | 1074971.0181 | 76.61 | 184.953
0.68228 | 1020438.977 | 65.6909 | 158.5918
0.66665 | 987789.1007 | 65.0805 | 157.1182
0.71067 | 1103251.8325 | 77.5681 | 187.266
0.69674 | 1065691.9211 | 72.8685 | 175.9202
0.69303 | 1023404.6063 | 75.444 | 182.1379
0.68196 | 1018829.877 | 69.3711 | 167.4767
0.69739 | 1055439.7578 | 73.8808 | 178.3641
0.68136 | 1010812.0131 | 65.1592 | 157.3083
0.67855 | 998380.9098 | 62.695 | 151.359
0.68348 | 1031986.356 | 62.2945 | 150.3923
0.67887 | 1014948.1248 | 79.6337 | 192.2527
0.69419 | 1046182.6 | 73.5703 | 177.6143
0.70901 | 1082174.0152 | 87.563 | 211.3958
0.70653 | 1071561.529 | 79.1517 | 191.089
0.66967 | 974989.1074 | 75.9823 | 183.4376
0.6776 | 998662.2267 | 71.9481 | 173.698
0.70715 | 1073981.1911 | 80.8955 | 195.299
0.70279 | 1082360.2252 | 86.8156 | 209.5914
0.68222 | 1019368.0205 | 77.9793 | 188.2588
0.65299 | 949754.3408 | 60.0074 | 144.8706
0.69083 | 1036497.8207 | 76.6525 | 185.0554
0.7026 | 1067477.3813 | 64.7393 | 156.2944
0.67638 | 1012368.7492 | 61.7518 | 149.082
0.67391 | 994396.209 | 75.3852 | 181.9959
0.68071 | 1020672.2116 | 79.8658 | 192.813
0.68611 | 1023085.1548 | 63.538 | 153.3943
0.70287 | 1078738.3376 | 63.3205 | 152.8691
0.67937 | 1008709.4841 | 82.1895 | 198.423
0.67798 | 1022236.1014 | 64.2469 | 155.1058
0.69889 | 1049365.1463 | 57.6687 | 139.2245
0.71029 | 1085528.7816 | 79.7585 | 192.5541
0.69611 | 1055960.3306 | 82.4004 | 198.9323
0.69073 | 1040191.3583 | 73.5153 | 177.4816
0.68243 | 1014944.7667 | 72.6035 | 175.2804
0.66974 | 999308.4709 | 69.2836 | 167.2654
0.67956 | 1008812.4686 | 67.0929 | 161.9766
0.68935 | 1046262.0438 | 79.2095 | 191.2285
0.69451 | 1045558.7282 | 81.7757 | 197.4241
0.70212 | 1064719.1974 | 70.4147 | 169.996
0.69796 | 1089378.7185 | 70.5411 | 170.3012
0.72043 | 1108266.069 | 69.5263 | 167.8513
0.68336 | 1017986.7718 | 59.3759 | 143.3461
0.69356 | 1053728.53 | 68.6657 | 165.7736
0.67623 | 1004159.9766 | 57.0341 | 137.6926
0.68167 | 1004386.0623 | 78.4809 | 189.4696
0.69686 | 1063036.8766 | 68.9444 | 166.4465
0.66533 | 971238.4903 | 74.5791 | 180.0499
0.7049 | 1081947.1992 | 74.9297 | 180.8964
0.69102 | 1036648.0233 | 79.9063 | 192.9109
0.70843 | 1082420.3898 | 87.6543 | 211.6161
0.68343 | 1021420.4312 | 69.3984 | 167.5425
0.67395 | 999819.4195 | 61.2059 | 147.7642
0.68062 | 1020193.02 | 78.2255 | 188.8532
0.69306 | 1056703.296 | 56.2634 | 135.8319
0.70855 | 1082337.0137 | 67.0923 | 161.9752
0.69447 | 1055899.0416 | 60.6017 | 146.3054
0.68325 | 1016463.0535 | 82.3504 | 198.8115
0.70301 | 1074532.6082 | 51.4883 | 124.3038
0.70865 | 1098250.169 | 75.6556 | 182.6488
0.69618 | 1045196.0686 | 71.5242 | 172.6747
0.70566 | 1083615.9526 | 78.453 | 189.4023
0.68791 | 1028434.4598 | 67.5683 | 163.1243
0.69452 | 1041440.9028 | 81.4532 | 196.6455
0.68922 | 1053208.6329 | 79.75 | 192.5336
0.66119 | 967895.4572 | 69.7399 | 168.3669
0.68265 | 1026316.7193 | 69.8105 | 168.5374
0.71755 | 1099224.2138 | 81.5963 | 196.9909
0.72154 | 1125823.7897 | 81.4295 | 196.5882
0.66023 | 976992.1299 | 70.4969 | 170.1947
0.68939 | 1038530.3898 | 69.9502 | 168.8747
0.6917 | 1055466.9132 | 58.6072 | 141.4904
0.68657 | 1032971.3745 | 62.0736 | 149.859
0.70361 | 1081246.3917 | 61.2507 | 147.8724
0.69073 | 1046608.5275 | 55.9137 | 134.9877
0.68044 | 1001940.2215 | 62.7437 | 151.4768
0.69513 | 1049592.9856 | 82.0185 | 198.0101
0.70034 | 1056105.4452 | 80.3222 | 193.915
0.70312 | 1063236.8849 | 82.1412 | 198.3064
0.67728 | 1012801.8889 | 74.7194 | 180.3885
0.69085 | 1024763.8362 | 74.1172 | 178.9349
0.66416 | 973359.7757 | 62.4482 | 150.7633
0.66378 | 982848.4409 | 75.6474 | 182.6291
0.68726 | 1053724.4972 | 69.2952 | 167.2933
0.6825 | 1020098.3628 | 72.9648 | 176.1526
0.69232 | 1056770.9529 | 87.5685 | 211.409
0.69835 | 1058291.6822 | 82.3165 | 198.7296
0.70028 | 1059466.7411 | 68.377 | 165.0767
0.69449 | 1069544.3909 | 80.3719 | 194.035
0.67517 | 1006249.2339 | 61.4886 | 148.4465
0.699 | 1058235.1423 | 58.1472 | 140.3798
0.68985 | 1057859.494 | 58.8975 | 142.191
0.68361 | 1032089.3166 | 66.3115 | 160.0902
0.68944 | 1031125.7308 | 76.4095 | 184.4688
0.69133 | 1034022.9695 | 69.4515 | 167.6708
0.6753 | 997510.234 | 54.5081 | 131.5943
0.69548 | 1062646.6583 | 92.0788 | 222.2979
0.70019 | 1064481.7557 | 74.2154 | 179.1719
0.72855 | 1135936.1769 | 74.2804 | 179.3288
0.68323 | 1018183.1065 | 72.7497 | 175.6333
0.68202 | 1020906.3464 | 73.0736 | 176.4153
0.67674 | 1001156.5017 | 70.4791 | 170.1516
0.69728 | 1045051.1734 | 74.887 | 180.7932
0.69222 | 1042481.7075 | 83.3642 | 201.259
0.67818 | 1005737.2707 | 68.8612 | 166.2456
0.6976 | 1053520.1892 | 73.724 | 177.9854
0.68382 | 1044452.2664 | 51.4514 | 124.2146
0.74052 | 1170057.4862 | 85.828 | 207.2071
0.67495 | 998785.1659 | 68.7134 | 165.8887
0.67917 | 1025005.1661 | 68.2097 | 164.6727
0.69546 | 1048363.2101 | 95.0078 | 229.3692
0.67544 | 1009336.4495 | 74.6663 | 180.2604
0.6889 | 1044076.5971 | 64.5789 | 155.9072
0.67661 | 1007138.2859 | 67.2742 | 162.4142
0.68404 | 1031016.5802 | 69.9879 | 168.9658
0.69287 | 1059693.5719 | 92.168 | 222.5133
0.70467 | 1076686.4121 | 72.7912 | 175.7336
0.68203 | 1011204.3074 | 67.0087 | 161.7732
0.67716 | 1010759.1227 | 74.582 | 180.057
0.66981 | 996478.6313 | 70.875 | 171.1073
0.69137 | 1046635.1656 | 60.3395 | 145.6723
0.6804 | 1036131.7097 | 82.7195 | 199.7024
0.68895 | 1037022.2379 | 55.5446 | 134.0966
0.69195 | 1047676.9599 | 83.8121 | 202.3403
0.70687 | 1079772.3355 | 81.3637 | 196.4293
0.70656 | 1074033.0316 | 62.7151 | 151.4076
0.67853 | 1012640.3904 | 77.8858 | 188.0329
0.70061 | 1066221.0356 | 79.4222 | 191.7422
0.69304 | 1049111.064 | 75.6376 | 182.6054
0.71428 | 1091500.5876 | 76.3537 | 184.3342
0.71278 | 1089604.6168 | 60.9302 | 147.0984
0.70286 | 1063231.5727 | 60.294 | 145.5627
0.67409 | 1018503.82 | 78.853 | 190.3679
0.67987 | 1010444.8002 | 61.3179 | 148.0345
0.687 | 1032835.5958 | 67.9468 | 164.0382
0.69455 | 1065300.1342 | 74.8756 | 180.7656
0.66624 | 968930.4273 | 63.877 | 154.2127
0.67153 | 989709.034 | 78.5022 | 189.521
0.69262 | 1050276.1204 | 80.7692 | 194.9941
0.6677 | 988248.4654 | 80.3975 | 194.0967
0.68164 | 1028145.4895 | 60.1198 | 145.142
0.67436 | 1006808.003 | 71.1263 | 171.714
0.67732 | 1003869.548 | 80.5078 | 194.3629
0.70912 | 1107757.5573 | 72.5371 | 175.1201
0.70592 | 1075871.2599 | 77.8633 | 187.9787
0.69582 | 1050698.8868 | 75.4502 | 182.1529
0.69865 | 1067260.1547 | 70.9334 | 171.2484
0.68143 | 1025425.8214 | 68.0105 | 164.1919
0.68316 | 1019977.4075 | 76.2776 | 184.1504
0.69246 | 1057964.8477 | 65.1021 | 157.1705
0.6986 | 1057332.4256 | 82.2127 | 198.4791
0.67694 | 1008268.1778 | 74.4227 | 179.6723
0.69824 | 1056870.3484 | 64.496 | 155.707
0.67262 | 1003180.9909 | 76.1973 | 183.9566
0.68971 | 1029146.547 | 71.0865 | 171.6179
0.68698 | 1031403.6425 | 65.9423 | 159.1989
0.70433 | 1071171.7049 | 87.2483 | 210.6361
0.67666 | 1019352.7628 | 78.6185 | 189.8019
0.70735 | 1075538.4624 | 75.9748 | 183.4194
0.68125 | 1015309.5921 | 80.9561 | 195.4453
0.68324 | 1020115.1072 | 72.8868 | 175.9643
0.68555 | 1016913.8565 | 70.6679 | 170.6074
0.68796 | 1016054.6516 | 64.8689 | 156.6073
0.67703 | 1008371.564 | 63.4638 | 153.2153
0.68295 | 1025602.9769 | 71.5429 | 172.7199
0.71003 | 1105451.0119 | 74.1716 | 179.0661
0.68299 | 1030600.2686 | 78.7502 | 190.1198
0.70031 | 1069302.7671 | 72.7419 | 175.6146
0.70786 | 1087582.4718 | 76.7702 | 185.3396
0.68883 | 1086968.5185 | 84.0509 | 202.9169
0.67767 | 1022698.4256 | 92.3565 | 222.9684
0.69212 | 1050125.4648 | 70.4529 | 170.0884
0.68017 | 1006487.4039 | 81.911 | 197.7506
0.69645 | 1066667.0113 | 81.5794 | 196.9501
0.69491 | 1046072.4879 | 60.0003 | 144.8534
0.7021 | 1059468.5304 | 70.0696 | 169.1629
0.72263 | 1106588.8706 | 67.2595 | 162.3788
0.68723 | 1040442.4348 | 77.3476 | 186.7337
0.65555 | 955734.1445 | 71.1709 | 171.8217
0.70794 | 1086015.6146 | 74.8881 | 180.7959
0.67716 | 1018591.3109 | 78.5028 | 189.5225
0.67526 | 1009605.5746 | 69.4254 | 167.6077
0.68636 | 1038874.4282 | 70.283 | 169.6782
0.71229 | 1077313.784 | 82.1572 | 198.3449
0.66124 | 969501.4579 | 85.5507 | 206.5376
0.69014 | 1014835.7071 | 76.7415 | 185.2704
0.66537 | 990458.0627 | 65.3819 | 157.8457
0.71288 | 1095723.7485 | 65.4937 | 158.1157
0.68957 | 1027874.2167 | 65.4703 | 158.0592
0.68006 | 1025993.3019 | 71.8749 | 173.5213
0.68961 | 1025112.9831 | 67.1839 | 162.1962
0.70473 | 1076916.757 | 84.5884 | 204.2144
0.69389 | 1048511.4169 | 77.1996 | 186.3763
0.69856 | 1045007.9142 | 71.2682 | 172.0567
0.67686 | 1014077.9578 | 67.7128 | 163.4732
0.6813 | 1023290.3251 | 69.7429 | 168.3744
0.69544 | 1048759.8236 | 73.619 | 177.732
0.69971 | 1080623.9766 | 78.9023 | 190.487
0.69471 | 1036826.7002 | 61.5753 | 148.6559
0.69603 | 1036055.3982 | 80.8566 | 195.205
0.68589 | 1032051.4548 | 64.1154 | 154.7884
0.70003 | 1057032.1112 | 68.2449 | 164.7578
0.68216 | 1020631.2962 | 91.5624 | 221.0513
0.71391 | 1113809.8277 | 68.5256 | 165.4353
0.70678 | 1070279.921 | 62.4765 | 150.8316
0.66814 | 996445.7943 | 74.518 | 179.9023
0.69919 | 1051918.7532 | 71.6739 | 173.0362
0.70382 | 1071503.814 | 62.8105 | 151.638
0.66943 | 1001808.7786 | 50.4111 | 121.7032
0.68562 | 1036187.8738 | 62.3531 | 150.5336
0.69514 | 1045446.6187 | 73.4183 | 177.2474
0.7101 | 1095045.3892 | 69.6965 | 168.2623
0.67869 | 1008041.4833 | 77.514 | 187.1354
0.67441 | 998660.8869 | 80.4361 | 194.19
0.6892 | 1036844.2725 | 64.1044 | 154.7616
0.67134 | 999162.3615 | 69.1081 | 166.8417
0.69035 | 1042570.9748 | 64.351 | 155.357
0.713 | 1076867.5485 | 68.2606 | 164.7957
0.69776 | 1058590.9618 | 68.5675 | 165.5365
0.68282 | 1021501.0925 | 79.6845 | 192.3753
0.70472 | 1049280.9534 | 69.137 | 166.9116
0.68184 | 1022577.2084 | 73.2132 | 176.7523
0.6784 | 1012311.1817 | 76.2362 | 184.0505
0.67542 | 999622.1875 | 78.8317 | 190.3166
0.68964 | 1042364.7686 | 68.3116 | 164.9187
0.67944 | 1018535.7278 | 65.8428 | 158.9587
0.68236 | 1024429.1852 | 89.0485 | 214.982
0.68991 | 1049407.6813 | 70.63 | 170.5158
0.71536 | 1098897.8415 | 78.4807 | 189.4691
0.66969 | 984480.1263 | 75.1037 | 181.3163
0.68695 | 1046618.2569 | 74.3764 | 179.5605
0.69386 | 1042622.4822 | 67.913 | 163.9565
0.674 | 998530.5459 | 63.1274 | 152.403
0.69844 | 1043026.2085 | 72.9021 | 176.0011
0.68961 | 1046544.1311 | 74.0845 | 178.8559
0.69942 | 1052210.4011 | 78.512 | 189.5447
0.69823 | 1055687.1112 | 77.2441 | 186.4836
0.69476 | 1047914.9561 | 83.4923 | 201.5684
0.66606 | 984420.9143 | 86.2387 | 208.1987
0.71428 | 1103226.3035 | 80.6065 | 194.6014
0.68696 | 1044780.1054 | 62.7471 | 151.485
0.69245 | 1045331.4091 | 71.9191 | 173.628
0.6671 | 974597.7174 | 72.998 | 176.2328
0.67852 | 1012952.1737 | 74.5941 | 180.0862
0.69862 | 1056326.4256 | 66.4114 | 160.3313
0.66976 | 995008.3242 | 74.2305 | 179.2082
0.67894 | 1021733.9167 | 69.8785 | 168.7017
0.68795 | 1025770.5591 | 69.2362 | 167.1509
0.70761 | 1100941.2257 | 82.8722 | 200.0713
0.68663 | 1033279.401 | 71.6928 | 173.0817
0.66788 | 976525.0056 | 71.3773 | 172.32
0.73127 | 1127465.4519 | 77.0369 | 185.9836
0.69687 | 1059376.4214 | 75.9583 | 183.3796
0.66828 | 994814.634 | 65.3828 | 157.848
0.71825 | 1141733.7848 | 79.4265 | 191.7526
0.69204 | 1057968.8583 | 70.9869 | 171.3776
0.68215 | 1021415.7439 | 76.499 | 184.6848
0.69621 | 1045807.1844 | 64.0075 | 154.5278
0.70845 | 1085938.1115 | 93.8066 | 226.4692
0.70468 | 1094457.1177 | 81.0556 | 195.6856
0.68256 | 1042269.6586 | 72.747 | 175.6268
0.69753 | 1054159.8001 | 68.303 | 164.898
0.68226 | 1019215.6791 | 71.1694 | 171.8181
0.67234 | 1010387.4071 | 78.4277 | 189.3412
0.67083 | 1005931.0338 | 76.6099 | 184.9527
0.67994 | 1010756.9786 | 75.9468 | 183.3518
0.6721 | 1002916.0792 | 82.6012 | 199.4168
0.69982 | 1060894.5124 | 72.2206 | 174.3559
0.68791 | 1046391.205 | 77.7123 | 187.6142
0.68813 | 1036234.0949 | 71.7405 | 173.197
0.69297 | 1045010.433 | 75.3502 | 181.9116
0.70392 | 1083968.6997 | 75.1741 | 181.4863
0.70391 | 1076721.7933 | 79.8288 | 192.7237
0.69332 | 1056714.8553 | 51.3421 | 123.9509
0.69739 | 1049777.6939 | 66.116 | 159.6182
0.67834 | 1004255.6052 | 66.9104 | 161.5361
0.68404 | 1038165.1862 | 87.0683 | 210.2016
0.70031 | 1065922.8625 | 74.3029 | 179.3831
0.69542 | 1055866.7054 | 70.5014 | 170.2054
0.68812 | 1045081.1377 | 72.3537 | 174.6773
0.69486 | 1065594.7326 | 73.0584 | 176.3787
0.667 | 986295.9559 | 80.3233 | 193.9175
0.70099 | 1068239.3394 | 87.4242 | 211.0607
0.68653 | 1049251.3538 | 83.0191 | 200.4259
0.68731 | 1047486.3705 | 82.7847 | 199.86
0.68806 | 1036555.594 | 66.1679 | 159.7434
0.69545 | 1052132.3618 | 68.3753 | 165.0726
0.67944 | 1007860.1721 | 74.1179 | 178.9364
0.67252 | 1010847.5974 | 76.7912 | 185.3903
0.69464 | 1064387.237 | 67.3008 | 162.4786
0.67288 | 1006434.3381 | 78.0591 | 188.4513
0.69903 | 1062514.5205 | 66.7611 | 161.1756
0.68558 | 1049923.9924 | 69.2138 | 167.097
0.6904 | 1040216.7853 | 69.9617 | 168.9025
0.70876 | 1078710.751 | 82.6727 | 199.5894
0.69834 | 1065487.0915 | 79.6257 | 192.2336
0.71185 | 1089590.2401 | 74.7219 | 180.3945
0.71217 | 1111329.1545 | 83.119 | 200.6671
0.68739 | 1037402.9917 | 80.3862 | 194.0695
0.70182 | 1082246.0523 | 73.4646 | 177.3591
0.66889 | 989212.7234 | 77.1812 | 186.332
0.69491 | 1058660.1202 | 66.9131 | 161.5425
0.67716 | 1017102.3016 | 73.3243 | 177.0205
0.67787 | 1019764.3378 | 64.1015 | 154.7546
0.6969 | 1067228.4122 | 72.9564 | 176.1324
0.70325 | 1049861.1972 | 72.8777 | 175.9423
0.69781 | 1055225.107 | 62.3889 | 150.6201
0.68405 | 1034050.9379 | 72.0161 | 173.8621
0.68147 | 1038701.1177 | 77.8706 | 187.9962
0.70277 | 1053725.1719 | 74.9184 | 180.869
0.68739 | 1029277.0547 | 59.6821 | 144.0854
0.70593 | 1078232.8401 | 63.016 | 152.134
0.70273 | 1066316.2962 | 70.4589 | 170.1028
0.66847 | 988705.0124 | 80.3424 | 193.9638
0.68897 | 1053769.4887 | 83.0807 | 200.5746
0.69736 | 1047562.1269 | 69.7523 | 168.3969
0.70876 | 1087577.9826 | 80.0825 | 193.3363
0.71648 | 1096254.9708 | 76.4515 | 184.5702
0.6919 | 1056679.6942 | 90.5853 | 218.6924
0.67014 | 997711.8507 | 61.4247 | 148.2923
0.6914 | 1039015.472 | 73.361 | 177.1092
0.6794 | 1031320.3489 | 75.8476 | 183.1124
0.68907 | 1034225.2821 | 69.6935 | 168.2549
0.6927 | 1045179.8552 | 60.1399 | 145.1905
0.68424 | 1019400.8788 | 59.3816 | 143.3599
0.68252 | 1027644.8999 | 71.731 | 173.1738
0.68799 | 1025633.2508 | 76.1626 | 183.8729
0.69332 | 1044226.9758 | 76.947 | 185.7665
0.69265 | 1041218.3753 | 66.0251 | 159.3987
0.67208 | 998183.9188 | 80.6759 | 194.7689
0.71658 | 1094944.8017 | 70.9516 | 171.2924
0.67361 | 998582.7783 | 86.4988 | 208.8265
0.69203 | 1056974.9526 | 65.0813 | 157.1201
0.69523 | 1042591.9621 | 79.8771 | 192.8405
0.70388 | 1082202.6651 | 85.1067 | 205.4658
0.65972 | 973694.1317 | 78.6147 | 189.7927
0.69574 | 1053194.1731 | 80.6306 | 194.6595
0.69475 | 1033514.5731 | 69.0758 | 166.7638
0.68778 | 1040997.707 | 77.3265 | 186.6828
0.66248 | 982734.5146 | 74.1211 | 178.9442
0.65879 | 978972.0608 | 70.7944 | 170.9127
0.67047 | 992725.5258 | 90.6129 | 218.759
0.68961 | 1024140.6309 | 67.1199 | 162.0417
0.68635 | 1037217.0982 | 71.9952 | 173.8119
0.69808 | 1046337.3499 | 82.9704 | 200.3083
0.66689 | 987382.6961 | 66.916 | 161.5494
0.66412 | 983625.9918 | 63.1103 | 152.3618
0.66773 | 996370.8588 | 69.288 | 167.2761
0.70075 | 1063050.9747 | 79.1707 | 191.1351
0.69849 | 1052934.2825 | 81.8835 | 197.6842
0.66541 | 980473.7323 | 62.7859 | 151.5786
0.68706 | 1021046.4873 | 74.2548 | 179.267
0.67954 | 1011068.4425 | 74.3438 | 179.4818
0.67423 | 995464.7591 | 75.4306 | 182.1056
0.66883 | 997054.0853 | 76.1614 | 183.8698
0.69228 | 1060426.2198 | 87.8764 | 212.1525
0.67254 | 1002104.9084 | 69.1878 | 167.0342
0.70148 | 1077791.0394 | 69.4408 | 167.645
0.68869 | 1038149.841 | 75.3208 | 181.8406
0.6888 | 1037675.0183 | 71.7705 | 173.2692
0.70628 | 1067097.148 | 73.7565 | 178.0639
0.71335 | 1103154.8364 | 61.6384 | 148.8083
0.72231 | 1130693.5714 | 66.7018 | 161.0323
0.68355 | 1022014.5062 | 84.2826 | 203.4762
0.67627 | 1004389.579 | 59.1971 | 142.9144
0.68939 | 1020885.1424 | 73.9927 | 178.6342
0.72247 | 1121152.7318 | 80.9874 | 195.5208
0.68123 | 1028029.7646 | 53.3162 | 128.7166
0.69153 | 1047183.5321 | 66.5109 | 160.5714
0.6792 | 1012866.9288 | 69.2453 | 167.173
0.68473 | 1017172.4551 | 72.5189 | 175.0761
0.71482 | 1098493.4145 | 71.3568 | 172.2706
0.68097 | 1011812.8779 | 75.9102 | 183.2635
0.68356 | 1015253.876 | 65.2348 | 157.4907
0.69724 | 1061541.4545 | 58.4758 | 141.1731
0.70055 | 1062650.6268 | 48.4753 | 117.0298
0.72468 | 1138049.7979 | 72.9207 | 176.046
0.69407 | 1042075.3052 | 69.7166 | 168.3107
0.66675 | 976262.4274 | 83.9994 | 202.7925
0.69971 | 1056037.0663 | 60.1556 | 145.2285
0.67058 | 1005945.5187 | 73.2314 | 176.7962
0.66776 | 997923.0666 | 74.6713 | 180.2725
0.69341 | 1065794.467 | 78.526 | 189.5786
0.69684 | 1043251.2877 | 74.8054 | 180.5962
0.69481 | 1056202.5908 | 53.7067 | 129.6595
0.69839 | 1063520.6381 | 76.0889 | 183.6949
0.68115 | 1022753.3908 | 69.721 | 168.3214
0.68115 | 1017212.535 | 77.9364 | 188.1551
0.70418 | 1062261.5088 | 66.2573 | 159.9592
0.69249 | 1042043.9269 | 78.5323 | 189.5937
0.69725 | 1039202.9971 | 74.4717 | 179.7906
0.70981 | 1096204.2286 | 78.9003 | 190.4821
0.71352 | 1098643.8766 | 67.6287 | 163.2701
0.704 | 1067469.2939 | 77.466 | 187.0196
0.68848 | 1037423.5349 | 78.6396 | 189.8528
0.68522 | 1018642.0752 | 72.5713 | 175.2027
0.69901 | 1074691.1216 | 71.3898 | 172.3503
0.69186 | 1049159.4893 | 66.6048 | 160.7981
0.69544 | 1059049.7229 | 68.8263 | 166.1614
0.72201 | 1134345.3907 | 81.9925 | 197.9473
0.69968 | 1053159.0287 | 71.709 | 173.1207
0.7055 | 1077481.8279 | 78.0264 | 188.3724
0.69089 | 1044561.8928 | 78.9036 | 190.4902
0.70411 | 1085218.3536 | 78.7421 | 190.1003
0.69874 | 1057349.6037 | 77.3099 | 186.6427
0.69105 | 1049980.0297 | 75.9187 | 183.2838
0.69174 | 1040136.5916 | 73.4039 | 177.2127
0.69085 | 1054862.9947 | 67.0648 | 161.9086
0.67416 | 1004954.0543 | 80.9734 | 195.487
0.68124 | 1041023.1095 | 89.166 | 215.2658
0.6828 | 1024692.9838 | 71.4953 | 172.6049
0.69157 | 1048340.878 | 85.4513 | 206.2976
0.73056 | 1155425.4146 | 99.6785 | 240.6453
0.70981 | 1083885.5373 | 74.5444 | 179.9661
0.71171 | 1088527.7127 | 79.8196 | 192.7016
0.69201 | 1054599.8036 | 77.8585 | 187.967
0.66669 | 1000938.4309 | 67.8544 | 163.8149
0.68893 | 1036204.83 | 59.9497 | 144.7314
0.6813 | 1011294.717 | 70.7155 | 170.7223
0.70442 | 1073713.043 | 81.1887 | 196.0069
0.69604 | 1036260.7882 | 69.1308 | 166.8966
0.67434 | 1004253.664 | 79.3234 | 191.5036
0.68303 | 1008890.9365 | 75.4629 | 182.1836
0.70114 | 1062000.2614 | 77.3844 | 186.8224
0.69268 | 1047362.982 | 68.3379 | 164.9823
0.67743 | 1010591.9363 | 71.07 | 171.5782
0.69398 | 1072322.0647 | 67.8672 | 163.846
0.69603 | 1047337.4284 | 68.0075 | 164.1847
0.71632 | 1108668.2775 | 63.8548 | 154.159
0.68332 | 1013610.0601 | 63.2065 | 152.5939
0.68084 | 1018901.5796 | 65.731 | 158.6886
0.71315 | 1105956.1938 | 77.4878 | 187.072
0.69664 | 1046757.2007 | 71.5839 | 172.8188
0.67393 | 1009700.2132 | 67.7128 | 163.4731
0.70512 | 1082051.2446 | 83.3565 | 201.2404
0.6871 | 1033071.9491 | 73.3978 | 177.1979
0.69741 | 1058145.5842 | 85.5287 | 206.4845
0.67564 | 1009018.611 | 74.6213 | 180.1518
0.72348 | 1104208.4773 | 56.3203 | 135.9691
0.72014 | 1103703.2385 | 84.8373 | 204.8153
0.70167 | 1065909.3301 | 82.5178 | 199.2157
0.69076 | 1042541.3405 | 75.4127 | 182.0623
0.68722 | 1044954.1213 | 74.2289 | 179.2043
0.7126 | 1101579.5783 | 73.6457 | 177.7964
0.68869 | 1025705.9225 | 72.7064 | 175.5288
0.6813 | 1022754.6577 | 78.3445 | 189.1402
0.69717 | 1035699.1132 | 72.522 | 175.0836
0.67579 | 1007710.9638 | 65.4947 | 158.1182
0.69363 | 1017824.9736 | 58.4879 | 141.2023
0.70128 | 1055385.0479 | 64.1259 | 154.8135
0.67494 | 993358.2781 | 66.5178 | 160.5881
0.6668 | 992153.5892 | 69.0907 | 166.7997
0.6835 | 1035489.2411 | 83.8613 | 202.4591
0.70525 | 1071894.6143 | 68.6913 | 165.8355
0.67957 | 1006238.8517 | 68.7548 | 165.9887
0.70944 | 1086912.4655 | 67.2915 | 162.4562
0.69687 | 1076561.1047 | 75.4149 | 182.0678
0.68059 | 1012126.689 | 74.3062 | 179.3909
0.70291 | 1076855.0974 | 74.7881 | 180.5544
0.67256 | 1000313.4188 | 85.3255 | 205.994
0.70728 | 1070372.5498 | 94.0085 | 226.9566
0.69416 | 1054458.9495 | 80.9119 | 195.3385
0.6935 | 1059533.1599 | 81.3778 | 196.4633
0.68176 | 1022643.5964 | 69.6512 | 168.1528
0.68656 | 1019256.6278 | 74.1791 | 179.0843
0.69969 | 1055219.8705 | 70.0283 | 169.0632
0.68725 | 1024396.2479 | 79.3944 | 191.6749
0.68906 | 1037050.957 | 66.1034 | 159.5877
0.68474 | 1016568.0928 | 65.207 | 157.4237
0.68644 | 1021966.5376 | 57.1817 | 138.0489
0.70376 | 1074248.19 | 76.271 | 184.1344
0.67557 | 997316.9758 | 75.4211 | 182.0827
0.71778 | 1101876.5064 | 51.7116 | 124.8428
0.68989 | 1058726.772 | 74.9169 | 180.8655
0.68644 | 1024941.8827 | 69.332 | 167.3822
0.68732 | 1049312.5646 | 81.3178 | 196.3186
0.69376 | 1048189.4907 | 58.0307 | 140.0985
0.69347 | 1056411.8039 | 60.7191 | 146.5888
0.69466 | 1061139.9211 | 76.3271 | 184.27
0.67338 | 1005313.1681 | 74.5965 | 180.092
0.69683 | 1044623.8021 | 79.8491 | 192.7727
0.69477 | 1046483.3991 | 75.261 | 181.696
0.7014 | 1067321.7003 | 87.2984 | 210.757
0.68482 | 1022139.4973 | 69.3066 | 167.3208
0.71199 | 1087007.6982 | 72.7339 | 175.5953
0.66505 | 984161.6173 | 87.7751 | 211.9077
0.69958 | 1077780.017 | 63.5615 | 153.451
0.68672 | 1054624.4542 | 81.8011 | 197.4852
0.68452 | 1032631.515 | 53.2973 | 128.6711
0.67407 | 1015130.9556 | 73.5179 | 177.488
0.69196 | 1059427.6047 | 65.2392 | 157.5013
0.69954 | 1061447.1974 | 75.8161 | 183.0363
0.68705 | 1042364.8569 | 80.8115 | 195.0963
0.68973 | 1056207.5239 | 58.8158 | 141.9939
0.67136 | 995039.1263 | 65.0983 | 157.1612
0.70052 | 1077322.2384 | 71.411 | 172.4014
0.71748 | 1118323.9371 | 70.4453 | 170.07
0.69778 | 1072851.0317 | 71.3938 | 172.3598
0.71129 | 1088199.4874 | 87.768 | 211.8908
0.68198 | 1026624.9325 | 62.1068 | 149.9391
0.69949 | 1061605.7164 | 72.3034 | 174.5559
0.6738 | 1010015.4316 | 71.4121 | 172.4041
0.69639 | 1051148.4688 | 68.219 | 164.6952
0.70841 | 1090025.5732 | 69.5426 | 167.8906
0.67807 | 1012659.9458 | 76.2991 | 184.2022
0.6977 | 1077905.7612 | 77.1945 | 186.3641
0.7172 | 1118991.3266 | 83.1052 | 200.6336
0.68542 | 1033962.208 | 71.8665 | 173.501
0.70126 | 1067829.3044 | 76.0852 | 183.6859
0.6859 | 1029902.0397 | 83.1169 | 200.662
0.7103 | 1109662.7132 | 84.8606 | 204.8717
0.68063 | 1017915.9864 | 84.3053 | 203.531
0.68367 | 1041151.9477 | 68.8279 | 166.1652
0.70296 | 1064119.893 | 91.2251 | 220.2369
0.6767 | 1012156.3558 | 79.6693 | 192.3387
0.70089 | 1079142.7515 | 82.0104 | 197.9907
0.68815 | 1027801.3831 | 76.3062 | 184.2194
0.68199 | 1024170.7111 | 89.075 | 215.046
0.6899 | 1044398.2803 | 72.1205 | 174.1142
0.67853 | 1020933.6539 | 78.5241 | 189.574
0.68901 | 1035897.8836 | 76.8112 | 185.4385
0.6831 | 1035013.2002 | 79.0742 | 190.9021
0.68191 | 1017709.5825 | 77.8882 | 188.0388
0.68432 | 1042627.889 | 66.9195 | 161.5579
0.67894 | 1014820.4031 | 69.4058 | 167.5604
0.67549 | 999858.4083 | 70.5176 | 170.2446
0.67996 | 1033621.3718 | 79.3322 | 191.5249
0.67621 | 1021760.373 | 78.8574 | 190.3787
0.69226 | 1042050.8137 | 80.1792 | 193.5698
0.67875 | 996053.43 | 78.1645 | 188.7058
0.68436 | 1036658.3776 | 75.6881 | 182.7272
0.68779 | 1015596.1927 | 62.9194 | 151.9009
0.70913 | 1088201.2501 | 76.6755 | 185.1111
0.68172 | 1025138.3454 | 83.2399 | 200.9588
0.68779 | 1042408.4541 | 77.4309 | 186.9347
0.68403 | 1024834.8296 | 82.9343 | 200.2212
0.68408 | 1020073.6015 | 96.6306 | 233.287
0.68326 | 1016179.2973 | 69.0597 | 166.7249
0.68169 | 1024576.1176 | 78.4391 | 189.3688
0.69214 | 1052475.15 | 74.2315 | 179.2108
0.67131 | 985270.3973 | 69.9216 | 168.8056
0.67356 | 1010159.4105 | 67.7808 | 163.6373
0.6904 | 1049663.6624 | 68.7144 | 165.8913
0.68978 | 1052507.3798 | 67.0189 | 161.7979
0.72482 | 1144752.5202 | 78.9152 | 190.5182
0.7052 | 1080221.5914 | 65.2705 | 157.5769
0.69696 | 1073125.807 | 73.3537 | 177.0915
0.69448 | 1046986.3788 | 73.5959 | 177.6763
0.67227 | 1003665.5242 | 71.1249 | 171.7106
0.69723 | 1049520.891 | 81.7665 | 197.4019
0.70201 | 1083868.3034 | 73.5547 | 177.5767
0.68344 | 1026212.497 | 61.9498 | 149.56
0.68833 | 1018382.9123 | 90.5418 | 218.5872
0.69514 | 1050981.0623 | 82.1265 | 198.2708
0.68022 | 1002943.369 | 75.3279 | 181.8576
0.67028 | 1008532.3744 | 78.5251 | 189.5764
0.71413 | 1102294.9077 | 65.7712 | 158.7857
0.68039 | 1025195.0657 | 78.6364 | 189.8451
0.71966 | 1119811.444 | 68.3435 | 164.9959
0.69361 | 1047286.4257 | 73.7099 | 177.9515
0.66202 | 969890.6215 | 64.3477 | 155.3491
0.68383 | 1024408.9864 | 78.1811 | 188.7459
0.68484 | 1033322.3917 | 68.193 | 164.6324
0.67952 | 1007622.9736 | 67.0178 | 161.7953
0.68709 | 1035481.3775 | 62.1768 | 150.1081
0.71592 | 1097055.7403 | 79.9715 | 193.0684
0.67869 | 1008482.8867 | 67.1577 | 162.1331
0.65165 | 955092.5371 | 71.5488 | 172.734
0.67928 | 1028436.8364 | 70.3018 | 169.7235
0.67853 | 1014101.4968 | 77.9987 | 188.3054
0.70787 | 1092782.3752 | 67.8054 | 163.6967
0.67213 | 991155.1752 | 83.1271 | 200.6866
0.71885 | 1103329.8517 | 71.27 | 172.061
0.66539 | 980765.4874 | 60.812 | 146.8131
0.6696 | 990537.9663 | 66.8575 | 161.4082
0.6907 | 1063390.1824 | 63.5193 | 153.3491
0.69566 | 1052772.8486 | 72.0966 | 174.0567
0.68415 | 1023613.3055 | 67.621 | 163.2515
0.67773 | 1014565.3532 | 82.866 | 200.0561
0.68471 | 1021938.6845 | 72.8603 | 175.9003
0.71578 | 1097912.0058 | 69.3242 | 167.3633
0.69938 | 1067454.5231 | 72.7128 | 175.5442
0.70541 | 1065415.4657 | 66.3372 | 160.1523
0.70484 | 1082257.1026 | 67.441 | 162.817
0.68161 | 1015296.79 | 67.3211 | 162.5275
0.69362 | 1056008.9852 | 62.8123 | 151.6423
0.68413 | 1037565.1184 | 66.6128 | 160.8176
0.69429 | 1043736.0147 | 84.7422 | 204.5859
0.67559 | 1002243.183 | 75.7914 | 182.9766
0.69831 | 1070633.7018 | 86.4258 | 208.6504
0.6693 | 991286.7888 | 66.9384 | 161.6037
0.69954 | 1080118.1552 | 83.8719 | 202.4846
0.68389 | 1027786.7933 | 69.7835 | 168.4723
0.69223 | 1057481.7777 | 65.0942 | 157.1512
0.70886 | 1076668.6427 | 80.9019 | 195.3145
0.66612 | 974057.4059 | 58.2845 | 140.7112
0.67269 | 1010694.9994 | 71.5784 | 172.8055
0.68928 | 1030779.2714 | 79.6543 | 192.3025
0.67716 | 1022004.3939 | 75.4112 | 182.0587
0.68706 | 1026540.4951 | 72.9382 | 176.0883
0.68935 | 1026194.3983 | 69.1235 | 166.879
0.70422 | 1074280.6979 | 62.9375 | 151.9447
0.69226 | 1035131.3073 | 91.5089 | 220.922
0.68481 | 1032948.4332 | 66.2505 | 159.9429
0.69886 | 1042350.7464 | 67.4097 | 162.7414
0.68323 | 1041298.2453 | 68.8817 | 166.2952
0.68334 | 1024478.8107 | 82.9891 | 200.3535
0.69334 | 1054933.3003 | 79.4059 | 191.7027
0.68487 | 1034854.0367 | 71.9862 | 173.79
0.70488 | 1079725.4426 | 87.876 | 212.1514
0.69517 | 1057554.444 | 60.6361 | 146.3885
0.71861 | 1109610.549 | 85.4734 | 206.3511
0.68026 | 1023285.6451 | 59.9568 | 144.7486
0.69626 | 1035906.7617 | 70.8806 | 171.1209
0.67648 | 1022088.1554 | 60.2218 | 145.3884
0.70278 | 1068981.0169 | 97.2273 | 234.7275
0.67628 | 1014559.1209 | 83.0437 | 200.4852
0.67182 | 1005141.7394 | 74.7574 | 180.4803
0.71287 | 1094059.8036 | 67.4082 | 162.7378
0.68129 | 1018428.6227 | 67.5608 | 163.1062
0.68502 | 1027280.3935 | 71.4439 | 172.4808
0.69655 | 1055067.9198 | 70.3346 | 169.8027
0.70767 | 1078419.3918 | 56.2337 | 135.7602
0.69082 | 1042551.5458 | 77.1017 | 186.1401
0.68265 | 1019950.3018 | 74.6543 | 180.2314
0.66341 | 986287.0788 | 82.2958 | 198.6796
0.69624 | 1058569.315 | 74.2223 | 179.1885
0.67477 | 997186.4549 | 68.6216 | 165.6672
0.69658 | 1070222.3457 | 75.0025 | 181.072
0.67719 | 1019095.9092 | 66.0974 | 159.5734
0.68606 | 1028711.7591 | 74.9213 | 180.8761
0.70835 | 1088720.1775 | 80.2122 | 193.6494
0.66936 | 1001148.0079 | 63.732 | 153.8627
0.67279 | 1005470.6595 | 69.3945 | 167.5331
0.66372 | 962117.8995 | 78.7415 | 190.0988
0.68915 | 1029918.6444 | 72.9566 | 176.1329
0.6699 | 986074.9934 | 71.8575 | 173.4795
0.67704 | 1004533.1839 | 77.4329 | 186.9395
0.6842 | 1025458.5838 | 74.8877 | 180.7948
0.689 | 1041664.0284 | 60.2593 | 145.4789
0.67732 | 1008383.0287 | 79.2994 | 191.4457
0.67456 | 1012175.8016 | 84.5845 | 204.2051
0.69029 | 1059906.7097 | 61.2006 | 147.7514
0.68061 | 1013512.9749 | 63.4526 | 153.1881
0.67835 | 1021662.1018 | 76.9453 | 185.7623
0.681 | 1020407.5198 | 74.8693 | 180.7504
0.74259 | 1165609.0098 | 72.8852 | 175.9605
0.68307 | 1014404.1966 | 72.4971 | 175.0236
0.68016 | 1024723.7749 | 76.9714 | 185.8254
0.6963 | 1039778.3982 | 72.8987 | 175.993
0.69817 | 1068218.9918 | 65.1954 | 157.3956
0.67812 | 992202.6261 | 60.366 | 145.7365
0.69274 | 1033788.19 | 71.0952 | 171.6389
0.68892 | 1036773.1056 | 84.8332 | 204.8053
0.65899 | 959674.6671 | 72.7645 | 175.6692
0.67249 | 1006759.7765 | 69.5152 | 167.8245
0.69214 | 1047477.0495 | 83.7732 | 202.2465
0.69183 | 1038422.7073 | 71.5079 | 172.6352
0.70261 | 1068717.0774 | 74.2543 | 179.2657
0.70727 | 1079610.1804 | 65.0428 | 157.0272
0.69203 | 1047607.5443 | 79.0041 | 190.7328
0.68412 | 1024431.0507 | 76.5032 | 184.6951
0.69936 | 1068702.9571 | 81.3944 | 196.5034
0.68013 | 1011818.7143 | 69.0797 | 166.7731
0.70489 | 1074110.9473 | 72.5952 | 175.2602
0.68907 | 1049640.8868 | 66.905 | 161.523
0.68013 | 1016330.8231 | 87.1925 | 210.5013
0.6638 | 975829.5343 | 60.6514 | 146.4254
0.66696 | 994994.593 | 76.217 | 184.0042
0.67298 | 1001629.7945 | 76.0388 | 183.574
0.6881 | 1036267.5043 | 64.1705 | 154.9213
0.68015 | 1009056.8774 | 68.8887 | 166.3121
0.68875 | 1053863.0836 | 58.7691 | 141.8812
0.67846 | 1021129.65 | 73.4737 | 177.3813
0.70826 | 1082799.8284 | 94.9682 | 229.2734
0.68521 | 1028612.4815 | 71.5159 | 172.6548
0.7049 | 1086116.2115 | 92.8551 | 224.172
0.67564 | 1002609.6599 | 63.2312 | 152.6537
0.67315 | 1017902.8416 | 72.0402 | 173.9204
0.66942 | 987728.5961 | 81.6757 | 197.1827
0.70467 | 1080274.1845 | 83.2728 | 201.0384
0.65416 | 964528.2786 | 68.4647 | 165.2885
0.68546 | 1044530.9928 | 77.8515 | 187.9501
0.67883 | 995595.6877 | 81.1846 | 195.9971
0.67475 | 1004135.0302 | 77.6646 | 187.499
0.7017 | 1071335.9426 | 76.2874 | 184.174
0.68619 | 1023945.7651 | 63.5838 | 153.5048
0.69235 | 1043689.398 | 62.1935 | 150.1483
0.67956 | 1012129.6348 | 74.4193 | 179.6642
0.7152 | 1100333.9858 | 75.1049 | 181.3192
0.7068 | 1070577.7722 | 68.457 | 165.2698
0.67126 | 1001811.1404 | 67.4516 | 162.8426
0.70702 | 1071384.0579 | 72.1919 | 174.2867
0.69491 | 1040552.9197 | 73.0987 | 176.4758
0.69294 | 1049017.3161 | 72.1217 | 174.1173
0.68456 | 1030590.2566 | 71.722 | 173.1523
0.70054 | 1068193.3053 | 64.9734 | 156.8596
0.67784 | 991378.3031 | 79.4453 | 191.798
0.6885 | 1026035.1561 | 74.4619 | 179.767
0.66819 | 994341.0823 | 63.4092 | 153.0834
0.68122 | 1011106.8625 | 70.7106 | 170.7104
0.68503 | 1032088.9631 | 71.7774 | 173.2861
0.72324 | 1127094.2476 | 77.8667 | 187.9869
0.68516 | 1015573.8107 | 74.0622 | 178.8019
0.68111 | 1019730.2 | 76.4887 | 184.6599
0.67185 | 992106.5709 | 75.443 | 182.1354
0.68632 | 1034998.6781 | 62.2181 | 150.2077
0.73115 | 1142354.3472 | 72.549 | 175.1488
0.69517 | 1031350.9911 | 54.7468 | 132.1704
0.68015 | 1017221.7636 | 71.25 | 172.0128
0.68661 | 1039361.2185 | 72.1163 | 174.1041
0.69539 | 1054916.9787 | 79.171 | 191.1356
0.67761 | 1010637.805 | 65.5082 | 158.1507
0.71584 | 1107021.6558 | 77.2714 | 186.5496
0.68634 | 1040387.519 | 70.3875 | 169.9306
0.69732 | 1049893.7555 | 69.9116 | 168.7814
0.69938 | 1078800.3842 | 80.5007 | 194.3459
0.68185 | 1033775.0649 | 60.4304 | 145.8918
0.67064 | 986151.0984 | 68.2796 | 164.8415
0.70166 | 1069084.5719 | 72.5878 | 175.2424
0.68511 | 1015566.1381 | 75.732 | 182.8331
0.6823 | 1025551.1126 | 66.0184 | 159.3826
0.70878 | 1077581.2896 | 87.2193 | 210.5659
0.67783 | 1030108.5692 | 95.7811 | 231.2361
0.69295 | 1051547.3702 | 80.8291 | 195.1388
0.69758 | 1062330.2772 | 68.8276 | 166.1646
0.69927 | 1061770.2938 | 77.5737 | 187.2796
0.65609 | 958329.2003 | 60.2636 | 145.4893
0.71009 | 1098248.7662 | 84.0131 | 202.8256
0.6723 | 1003217.6275 | 72.8458 | 175.8652
0.69481 | 1044187.5266 | 66.0082 | 159.3579
0.70287 | 1071542.9542 | 70.7899 | 170.902
0.69647 | 1051806.7934 | 80.4405 | 194.2006
0.67046 | 1000922.3199 | 81.093 | 195.7758
0.67443 | 995012.3953 | 72.0219 | 173.8761
0.70251 | 1073254.665 | 74.5805 | 180.0533
0.68377 | 1026092.9277 | 78.3154 | 189.0701
0.68052 | 1032263.5779 | 80.7959 | 195.0586
0.69053 | 1050072.6449 | 79.3611 | 191.5947
0.68228 | 1022604.7832 | 71.4036 | 172.3834
0.68132 | 1012138.6593 | 75.896 | 183.2292
0.70218 | 1075934.2714 | 83.233 | 200.9421
0.71153 | 1100626.1894 | 65.9119 | 159.1255
0.70135 | 1084582.8636 | 73.673 | 177.8623
0.69121 | 1058044.8829 | 61.0117 | 147.2952
0.69341 | 1047083.5784 | 71.9396 | 173.6775
0.67695 | 1011197.1703 | 77.9567 | 188.204
0.69181 | 1041636.9653 | 66.1317 | 159.6561
0.68865 | 1060327.8113 | 75.1651 | 181.4645
0.71704 | 1097393.1132 | 58.7815 | 141.9111
0.67049 | 1002660.183 | 82.7779 | 199.8435
0.70778 | 1068520.1143 | 80.413 | 194.1342
0.67507 | 994204.8971 | 71.8956 | 173.5714
0.69637 | 1057905.3696 | 77.6137 | 187.3762
0.70225 | 1080986.8507 | 63.5083 | 153.3226
0.69193 | 1037357.4886 | 81.4374 | 196.6073
0.69724 | 1071879.706 | 78.5406 | 189.6139
0.69192 | 1051710.5182 | 64.7675 | 156.3625
0.69482 | 1036221.4328 | 88.458 | 213.5566
0.67643 | 1017196.8474 | 55.4909 | 133.9668
0.66964 | 973288.681 | 72.6237 | 175.3291
0.7334 | 1158729.8222 | 69.3946 | 167.5333
0.70224 | 1057803.9868 | 78.613 | 189.7886
0.6895 | 1035436.7458 | 59.8556 | 144.5043
0.70804 | 1085733.5454 | 63.2652 | 152.7357
0.68139 | 1021135.8521 | 71.3942 | 172.3608
0.68596 | 1035088.28 | 92.3451 | 222.9408
0.71393 | 1101085.1032 | 69.7046 | 168.2818
0.68746 | 1031015.5922 | 74.9369 | 180.9138
0.70356 | 1067807.4242 | 74.091 | 178.8714
0.69128 | 1053000.9323 | 82.7429 | 199.7591
0.66515 | 991453.6319 | 67.367 | 162.6383
0.71462 | 1093877.962 | 74.7255 | 180.4032
0.67703 | 1020825.7628 | 86.4998 | 208.8289
0.70286 | 1075195.402 | 71.0324 | 171.4874
0.70495 | 1081426.7493 | 73.0834 | 176.439
0.69827 | 1063565.1957 | 84.3127 | 203.5488
0.71159 | 1079280.0679 | 70.4273 | 170.0265
0.64762 | 937547.0556 | 67.2223 | 162.289
0.65652 | 971844.8318 | 68.736 | 165.9433
0.71193 | 1100133.292 | 82.3094 | 198.7124
0.70892 | 1085834.1902 | 86.3415 | 208.4467
0.72118 | 1104151.1565 | 76.4997 | 184.6866
0.70335 | 1069242.8613 | 75.9596 | 183.3827
0.69211 | 1036342.2862 | 71.3352 | 172.2184
0.6845 | 1038002.021 | 53.5764 | 129.3448
0.68612 | 1030622.959 | 65.9756 | 159.2792
0.689 | 1024841.5535 | 63.3805 | 153.0142
0.67988 | 1004602.0355 | 62.9623 | 152.0044
0.66632 | 989293.222 | 75.8535 | 183.1265
0.68035 | 1022540.7642 | 64.0416 | 154.6102
0.69198 | 1038397.0997 | 61.9114 | 149.4673
0.69804 | 1078421.527 | 71.7374 | 173.1895
0.69143 | 1033898.2474 | 67.5811 | 163.1551
0.68162 | 1023267.7208 | 63.9686 | 154.4338
0.6969 | 1049608.5996 | 64.1761 | 154.9348
0.67694 | 1017264.7186 | 67.9172 | 163.9667
0.68506 | 1029292.8283 | 73.3475 | 177.0764
0.6994 | 1069323.3208 | 79.792 | 192.635
0.70229 | 1059471.4068 | 81.9463 | 197.8359
0.67703 | 999164.4028 | 77.0199 | 185.9425
0.71295 | 1112420.6154 | 77.9621 | 188.2172
0.71568 | 1102138.2319 | 85.2454 | 205.8006
0.7073 | 1077468.2795 | 74.3662 | 179.5358
0.65964 | 965116.6605 | 73.1052 | 176.4917
0.68254 | 1029805.7182 | 71.8794 | 173.5322
0.69579 | 1066011.1921 | 71.1409 | 171.7494
0.68242 | 1029155.7258 | 70.0273 | 169.0608
0.69868 | 1067159.3637 | 65.0828 | 157.1238
0.68671 | 1043525.6242 | 80.9411 | 195.4091
0.6916 | 1064691.3907 | 61.3603 | 148.1369
0.68538 | 1028232.5458 | 76.6462 | 185.0404
0.69316 | 1048822.0237 | 82.1917 | 198.4284
0.69669 | 1064295.8124 | 74.5752 | 180.0405
0.68448 | 1045710.9357 | 60.8929 | 147.0084
0.69617 | 1055291.4234 | 72.2165 | 174.3462
0.69057 | 1060478.1898 | 71.3066 | 172.1494
0.70997 | 1092413.4257 | 92.7262 | 223.8608
0.68854 | 1026600.2947 | 76.1736 | 183.8993
0.65627 | 949679.7761 | 75.5206 | 182.323
0.68458 | 1043451.0339 | 74.4087 | 179.6386
0.6946 | 1053907.3672 | 77.5173 | 187.1434
0.69473 | 1052326.9906 | 79.0513 | 190.8467
0.67596 | 1002139.6401 | 59.0914 | 142.6593
0.72794 | 1137382.8479 | 68.6272 | 165.6807
0.68832 | 1038629.1098 | 55.6816 | 134.4272
0.69474 | 1044557.0685 | 73.3976 | 177.1975
0.68822 | 1053485.2465 | 83.1489 | 200.7393
0.6885 | 1043469.5612 | 75.7269 | 182.8208
0.70517 | 1076366.6014 | 74.5141 | 179.893
0.70017 | 1079347.6658 | 86.8977 | 209.7896
0.71075 | 1097212.3084 | 67.8045 | 163.6946
0.68423 | 1029172.5703 | 62.056 | 149.8165
0.72926 | 1120133.8907 | 85.0239 | 205.2659
0.6907 | 1033905.4413 | 82.8602 | 200.0422
0.69999 | 1061963.9104 | 77.5082 | 187.1215
0.70813 | 1074604.2324 | 68.5854 | 165.5798
0.69511 | 1035491.3822 | 86.9219 | 209.8481
0.69874 | 1056452.171 | 78.9056 | 190.4949
0.68341 | 1047261.3113 | 78.2892 | 189.0069
0.70244 | 1076330.4187 | 67.3008 | 162.4785
0.68511 | 1034051.1055 | 73.6812 | 177.8821
0.68868 | 1051523.8428 | 81.3276 | 196.3422
0.6783 | 1018375.2507 | 81.5532 | 196.8869
0.70951 | 1091264.6715 | 76.851 | 185.5347
0.69935 | 1056094.0759 | 73.7194 | 177.9744
0.69871 | 1062336.6696 | 77.6775 | 187.53
0.68067 | 1026936.8754 | 73.6509 | 177.809
0.67601 | 1008860.2671 | 71.0979 | 171.6454
0.68988 | 1037562.7783 | 62.0501 | 149.8023
0.69541 | 1047028.6825 | 64.1551 | 154.8841
0.69476 | 1056344.8348 | 79.3499 | 191.5677
0.66897 | 987449.1461 | 56.6934 | 136.8701
0.69836 | 1071515.9835 | 61.1554 | 147.6422
0.71525 | 1115647.5711 | 86.342 | 208.4481
0.70816 | 1093446.9374 | 78.8855 | 190.4466
0.69552 | 1057953.0356 | 51.2395 | 123.7031
0.66351 | 967955.7687 | 67.7402 | 163.5393
0.70364 | 1076139.3418 | 79.8163 | 192.6935
0.66828 | 988743.8132 | 70.4861 | 170.1685
0.68865 | 1039201.5484 | 71.8917 | 173.562
0.70341 | 1076554.6288 | 72.913 | 176.0274
0.69153 | 1057584.1244 | 67.1719 | 162.1674
0.68235 | 1009507.5301 | 77.2104 | 186.4023
0.69222 | 1051642.8041 | 62.679 | 151.3205
0.70669 | 1072408.7342 | 74.081 | 178.8475
0.69526 | 1057512.0675 | 66.8073 | 161.287
0.67477 | 1005240.8456 | 59.3786 | 143.3527
0.658 | 956067.161 | 61.9124 | 149.4697
0.68255 | 1011997.6873 | 80.5381 | 194.4362
0.66393 | 988751.9191 | 65.4107 | 157.9154
0.68469 | 1011425.9584 | 66.7626 | 161.1793
0.68227 | 1014014.9459 | 63.3289 | 152.8896
0.68749 | 1020939.693 | 77.5008 | 187.1036
0.68697 | 1030036.571 | 71.4318 | 172.4515
0.69902 | 1069468.9881 | 86.2449 | 208.2135
0.69104 | 1041605.1303 | 77.0072 | 185.9119
0.67429 | 1019596.4528 | 73.8171 | 178.2103
0.69096 | 1026185.095 | 75.8759 | 183.1806
0.71166 | 1096489.8257 | 59.1775 | 142.8671
0.67519 | 1005362.4244 | 85.4227 | 206.2285
0.6828 | 1022439.1029 | 85.9049 | 207.3927
0.67474 | 1008730.3919 | 67.5666 | 163.1201
0.70285 | 1077802.6369 | 76.6214 | 184.9804
0.71034 | 1096930.3269 | 67.826 | 163.7464
0.69325 | 1050675.5412 | 73.9864 | 178.619
0.68445 | 1029766.1582 | 78.5181 | 189.5595
0.67014 | 991984.3699 | 83.444 | 201.4516
0.68341 | 1037112.6918 | 70.4666 | 170.1215
0.6638 | 975997.9895 | 75.1175 | 181.3498
0.69825 | 1057205.5044 | 79.0176 | 190.7653
0.68404 | 1024170.9265 | 65.7614 | 158.7621
0.71136 | 1106405.4882 | 77.7141 | 187.6183
0.6753 | 1015035.5236 | 76.7924 | 185.3931
0.6707 | 994815.4506 | 77.4078 | 186.8791
0.68331 | 1033799.1529 | 103.2047 | 249.1581
0.70794 | 1082788.295 | 58.7449 | 141.8228
0.68799 | 1044763.81 | 66.6987 | 161.0249
0.69933 | 1057502.6006 | 64.434 | 155.5574
0.69447 | 1056733.4067 | 64.8383 | 156.5335
0.69478 | 1039027.1384 | 71.2449 | 172.0004
0.70709 | 1085490.8772 | 83.3894 | 201.3197
0.6873 | 1034473.7956 | 71.7052 | 173.1117
0.69413 | 1058970.7393 | 77.9428 | 188.1705
0.68368 | 1028344.1928 | 69.8431 | 168.6162
0.68016 | 1031477.5 | 70.4183 | 170.0049
0.70541 | 1075682.9235 | 81.2237 | 196.0914
0.69092 | 1055409.5219 | 78.0413 | 188.4083
0.69723 | 1051449.1711 | 70.0247 | 169.0545
0.69093 | 1044307.1794 | 83.6657 | 201.9868
0.67428 | 998376.1905 | 72.282 | 174.5043
0.70201 | 1062295.291 | 66.464 | 160.4584
0.68737 | 1059071.264 | 75.2154 | 181.5861
0.68276 | 1028542.384 | 76.9632 | 185.8056
0.68936 | 1040691.4361 | 70.0988 | 169.2335
0.6773 | 1014570.6669 | 81.3694 | 196.4432
0.68484 | 1028227.4677 | 73.2947 | 176.9491
0.69349 | 1053824.7696 | 72.3948 | 174.7764
0.68264 | 1042841.4015 | 82.9825 | 200.3374
0.67708 | 1002560.0581 | 70.5421 | 170.3037
0.70742 | 1081011.8574 | 72.7997 | 175.7541
0.70636 | 1077719.7178 | 62.948 | 151.9699
0.70436 | 1073953.0063 | 70.7994 | 170.9248
0.69375 | 1051793.9698 | 72.2283 | 174.3745
0.69519 | 1045283.7666 | 66.0413 | 159.4378
0.67188 | 999667.7248 | 74.3429 | 179.4796
0.69991 | 1057260.6081 | 64.3661 | 155.3934
0.71501 | 1088840.0431 | 66.8858 | 161.4767
0.67269 | 1006638.2567 | 74.7713 | 180.5138
0.69409 | 1064528.5846 | 69.3166 | 167.345
0.68062 | 1019481.9161 | 74.8428 | 180.6865
0.67611 | 1012138.4378 | 76.481 | 184.6415
0.68827 | 1034360.6797 | 67.9214 | 163.9767
0.69578 | 1039987.6959 | 61.6679 | 148.8794
0.69591 | 1073446.2442 | 75.3009 | 181.7923
0.7067 | 1064854.2401 | 62.1667 | 150.0837
0.67655 | 1009794.5297 | 72.8629 | 175.9066
0.68594 | 1029322.3814 | 78.314 | 189.0666
0.67294 | 1012734.0217 | 75.1957 | 181.5385
0.69528 | 1060377.4741 | 72.6574 | 175.4105
0.67588 | 1000509.5732 | 84.9226 | 205.0214
0.66595 | 991536.5292 | 61.4279 | 148.3001
0.70274 | 1088948.663 | 84.172 | 203.2093
0.70854 | 1080974.7438 | 81.2625 | 196.185
0.67739 | 1019279.6703 | 71.6857 | 173.0647
0.69013 | 1035287.4373 | 63.9212 | 154.3193
0.69446 | 1048500.2022 | 64.334 | 155.316
0.67921 | 1010307.8346 | 83.8493 | 202.4301
0.70312 | 1059551.64 | 68.1065 | 164.4235
0.69843 | 1061023.4598 | 82.5198 | 199.2204
0.71074 | 1076037.4388 | 73.9078 | 178.4292
0.67895 | 1004142.6859 | 70.459 | 170.1031
0.70378 | 1091348.8649 | 77.4829 | 187.0604
0.67517 | 1013595.9422 | 65.9376 | 159.1876
0.6791 | 1026093.2462 | 71.9835 | 173.7836
0.67847 | 1009027.3449 | 83.8355 | 202.3968
0.70847 | 1074888.732 | 68.5706 | 165.5442
0.6948 | 1042420.3101 | 84.5738 | 204.1793
0.70141 | 1068219.4451 | 93.2019 | 225.0092
0.67462 | 988892.6728 | 58.426 | 141.0527
0.69285 | 1036113.5786 | 73.0036 | 176.2462
0.69493 | 1062394.4892 | 67.1472 | 162.1076
0.70245 | 1069749.8138 | 76.0928 | 183.7043
0.68582 | 1037187.5376 | 91.1382 | 220.0271
0.69388 | 1062953.1757 | 84.7981 | 204.7207
0.69462 | 1057275.0001 | 85.8008 | 207.1416
0.69032 | 1048897.5107 | 66.7366 | 161.1164
0.66275 | 971099.2633 | 71.5302 | 172.6893
0.69118 | 1039791.8502 | 67.9907 | 164.1441
0.69997 | 1059220.8803 | 83.7652 | 202.2272
0.69219 | 1039006.3058 | 65.929 | 159.1666
0.70932 | 1058551.6929 | 63.5895 | 153.5186
0.6662 | 996990.1574 | 52.8725 | 127.6456
0.68317 | 1015704.3688 | 65.3016 | 157.6521
0.71214 | 1081341.2682 | 66.6933 | 161.0119
0.67738 | 1021365.7698 | 62.2083 | 150.184
0.68504 | 1036711.0036 | 82.2055 | 198.4617
0.68571 | 1022951.5172 | 77.8558 | 187.9606
0.6732 | 1007935.8177 | 67.6686 | 163.3663
0.70142 | 1077695.9107 | 81.8389 | 197.5767
0.69254 | 1044841.1769 | 71.8753 | 173.5224
0.68331 | 1027658.8687 | 78.4881 | 189.487
0.67979 | 1007602.3023 | 68.8263 | 166.1614
0.68511 | 1017836.0147 | 74.2322 | 179.2123
0.69256 | 1046317.2501 | 70.1523 | 169.3627
0.6763 | 1006962.7875 | 65.3965 | 157.8812
0.68941 | 1053202.0862 | 76.3332 | 184.2846
0.6971 | 1054233.151 | 78.9254 | 190.5427
0.71709 | 1107512.3541 | 80.6022 | 194.5908
0.70371 | 1073074.4742 | 70.8062 | 170.9414
0.68944 | 1029232.7111 | 72.8774 | 175.9416
0.70591 | 1065348.5906 | 78.7617 | 190.1475
0.69958 | 1055878.4413 | 60.1578 | 145.2337
0.65641 | 971523.0999 | 61.063 | 147.4191
0.6748 | 991281.5311 | 68.2121 | 164.6786
0.69562 | 1043478.0862 | 65.886 | 159.0629
0.66224 | 979804.7018 | 68.7009 | 165.8587
0.70355 | 1078792.9542 | 63.2623 | 152.7286
0.71116 | 1093151.5685 | 76.8047 | 185.423
0.69012 | 1048388.1491 | 77.6113 | 187.3703
0.67603 | 1006387.6385 | 79.2938 | 191.4323
0.70205 | 1039931.0092 | 75.6772 | 182.7009
0.67006 | 985500.7092 | 70.0686 | 169.1605
0.68203 | 1042511.9445 | 86.165 | 208.0206
0.70657 | 1083014.4137 | 79.9935 | 193.1214
0.67982 | 1025831.3924 | 73.6646 | 177.842
0.68644 | 1043286.1566 | 66.7549 | 161.1605
0.69404 | 1056379.8747 | 77.7616 | 187.7332
0.68424 | 1025719.8684 | 74.9025 | 180.8305
0.68768 | 1026874.0661 | 76.2887 | 184.1771
0.69757 | 1047528.4124 | 78.3704 | 189.2029
0.69273 | 1058841.0464 | 84.6099 | 204.2665
0.68304 | 1026499.1696 | 86.2154 | 208.1423
0.71152 | 1098663.1006 | 77.1534 | 186.2647
0.68912 | 1043902.2017 | 88.3902 | 213.3929
0.67181 | 992051.7162 | 81.7415 | 197.3415
0.69158 | 1053797.6656 | 48.6946 | 117.5593
0.69798 | 1061886.0529 | 76.0173 | 183.5221
0.71911 | 1111746.2814 | 67.904 | 163.9348
0.67084 | 996477.7434 | 67.1342 | 162.0764
0.67467 | 1011699.7578 | 67.773 | 163.6185
0.6721 | 993335.4378 | 72.1496 | 174.1846
0.67055 | 1013440.2824 | 53.3961 | 128.9097
0.69383 | 1037894.7412 | 67.6082 | 163.2206
0.69678 | 1062574.0254 | 58.5718 | 141.4048
0.71556 | 1098074.8814 | 78.3995 | 189.2732
0.69406 | 1052216.6293 | 83.9197 | 202.6
0.66325 | 980987.904 | 93.2194 | 225.0516
0.69448 | 1090313.4048 | 74.7946 | 180.5702
0.69471 | 1036637.2944 | 67.4407 | 162.8162
0.70737 | 1074584.2688 | 82.519 | 199.2185
0.70736 | 1089946.181 | 65.6573 | 158.5107
0.72512 | 1140690.3606 | 91.7898 | 221.6001
0.67972 | 1011094.0662 | 75.9958 | 183.4702
0.68009 | 1006087.2466 | 72.2179 | 174.3494
0.70005 | 1071663.6286 | 63.1586 | 152.4783
0.69444 | 1046086.8521 | 55.5871 | 134.1991
0.68764 | 1047687.34 | 74.3782 | 179.5647
0.67093 | 997538.157 | 70.8317 | 171.0028
0.69326 | 1054185.9539 | 82.6 | 199.414
0.69857 | 1074016.1056 | 78.5472 | 189.6296
0.67131 | 1012379.3932 | 71.8456 | 173.4506
0.70369 | 1066358.9352 | 66.559 | 160.6876
0.68285 | 1020535.8915 | 72.9589 | 176.1383
0.67232 | 1009198.6378 | 65.9303 | 159.1699
0.66621 | 982569.4437 | 75.3501 | 181.9112
0.67898 | 1008703.7822 | 74.2303 | 179.2079
0.68687 | 1032624.3117 | 80.0772 | 193.3234
0.68613 | 1048744.9654 | 72.818 | 175.7982
0.70876 | 1088349.3662 | 69.4005 | 167.5477
0.6872 | 1040088.1195 | 80.6174 | 194.6277
0.7061 | 1078645.6403 | 78.0446 | 188.4163
0.70454 | 1055552 | 76.5431 | 184.7915
0.67773 | 996851.7105 | 71.8032 | 173.3484
0.70977 | 1100512.4899 | 84.224 | 203.3348
0.69529 | 1051340.2544 | 68.34 | 164.9873
0.69066 | 1052982.3107 | 72.6328 | 175.351
0.69581 | 1061362.9133 | 81.2478 | 196.1494
0.70493 | 1081907.4422 | 65.3979 | 157.8845
0.67339 | 1012700.084 | 80.1033 | 193.3865
0.71259 | 1095172.0218 | 71.3646 | 172.2894
0.66528 | 988112.7198 | 63.2876 | 152.7899
0.69252 | 1047327.3593 | 62.3951 | 150.6352
0.68665 | 1029051.4997 | 81.4066 | 196.5328
0.68079 | 1022590.1599 | 78.3906 | 189.2517
0.7069 | 1078780.7379 | 68.9717 | 166.5124
0.69684 | 1055680.082 | 64.5528 | 155.8443
0.68548 | 1031387.4057 | 68.0018 | 164.171
0.69607 | 1069322.7391 | 67.4742 | 162.8972
0.68558 | 1043887.2122 | 79.4632 | 191.8412
0.67281 | 988578.1914 | 64.5359 | 155.8034
0.69781 | 1056549.5886 | 66.5019 | 160.5499
0.68336 | 1031258.9476 | 57.9616 | 139.9318
0.68924 | 1033808.7017 | 74.9458 | 180.9352
0.66825 | 983066.7074 | 75.1582 | 181.4479
0.67956 | 1014605.7898 | 76.2777 | 184.1506
0.68877 | 1044334.5792 | 75.7606 | 182.9022
0.6677 | 988460.6239 | 86.3806 | 208.5413
0.69243 | 1061413.5701 | 74.8311 | 180.6583
0.68241 | 1023622.637 | 81.4802 | 196.7107
0.67466 | 1005783.1881 | 76.1646 | 183.8775
0.65626 | 957829.8247 | 60.216 | 145.3744
0.67266 | 1001959.3585 | 79.3399 | 191.5435
0.68465 | 1024187.7606 | 74.7457 | 180.452
0.69335 | 1054181.6853 | 77.6771 | 187.529
0.70363 | 1069649.3646 | 66.7702 | 161.1975
0.68606 | 1042957.6889 | 72.4234 | 174.8456
0.68844 | 1037681.911 | 81.8757 | 197.6655
0.69985 | 1082519.7654 | 67.9379 | 164.0166
0.67112 | 1005305.4655 | 42.2175 | 101.922
0.70932 | 1093357.7751 | 69.6648 | 168.1858
0.70078 | 1075711.1888 | 74.9359 | 180.9114
0.68382 | 1030444.652 | 80.2639 | 193.7741
0.70279 | 1062999.2865 | 77.2745 | 186.5572
0.69788 | 1063792.9474 | 85.9268 | 207.4457
0.69242 | 1049490.9711 | 72.2077 | 174.3248
0.70422 | 1076763.9277 | 72.8436 | 175.86
0.66963 | 993455.9247 | 67.052 | 161.8779
0.68313 | 1027618.9376 | 66.0709 | 159.5092
0.69233 | 1062632.7063 | 80.1024 | 193.3844
0.68428 | 1028360.7655 | 62.0026 | 149.6876
0.67762 | 1018914.916 | 67.959 | 164.0676
0.7072 | 1084986.7648 | 84.9247 | 205.0263
0.6548 | 968913.4064 | 73.9354 | 178.4959
0.68919 | 1045922.7044 | 60.4892 | 146.0338
0.69333 | 1052732.2953 | 77.5731 | 187.278
0.69752 | 1055262.4932 | 73.8542 | 178.2997
0.69331 | 1065624.7436 | 83.4905 | 201.5639
0.69326 | 1060236.077 | 78.1329 | 188.6295
0.69735 | 1054134.1983 | 50.9207 | 122.9335
0.68795 | 1023537.4524 | 65.3476 | 157.763
0.6809 | 1016995.7583 | 67.5043 | 162.9697
0.69966 | 1057796.7702 | 77.2411 | 186.4766
0.66284 | 973514.9881 | 71.6975 | 173.0931
0.68837 | 1045241.4042 | 79.2446 | 191.3133
0.69838 | 1073173.9994 | 75.0161 | 181.1049
0.69414 | 1038061.0446 | 69.4408 | 167.6449
0.70031 | 1056223.4874 | 62.2256 | 150.2258
0.69084 | 1035482.7705 | 65.8339 | 158.937
0.70165 | 1073201.8511 | 72.8681 | 175.9192
0.68575 | 1041874.1583 | 74.1687 | 179.059
0.6735 | 1000514.919 | 83.7414 | 202.1696
0.68672 | 1019478.1747 | 67.0957 | 161.9833
0.67896 | 1017283.8255 | 76.5473 | 184.8016
0.67739 | 1006725.991 | 63.3271 | 152.8851
0.69474 | 1050043.9729 | 77.2179 | 186.4206
0.66923 | 992017.7319 | 52.3712 | 126.4354
0.7091 | 1095736.3578 | 78.9456 | 190.5916
0.69319 | 1059476.4183 | 67.1312 | 162.0691
0.67883 | 1024758.226 | 66.3045 | 160.0732
0.67282 | 1006919.7746 | 74.778 | 180.53
0.67185 | 1003386.7174 | 69.3479 | 167.4207
0.70276 | 1061915.9865 | 73.7655 | 178.0858
0.69936 | 1071703.7499 | 71.156 | 171.7857
0.72156 | 1121061.4954 | 89.1318 | 215.1832
0.69003 | 1052823.7458 | 68.958 | 166.4793
0.69242 | 1043498.8091 | 73.569 | 177.6113
0.67186 | 995068.0754 | 80.4513 | 194.2267
0.69467 | 1047980.6063 | 77.073 | 186.0706
0.68669 | 1029549.9078 | 69.6183 | 168.0733
0.69026 | 1055606.2097 | 74.6556 | 180.2347
0.69032 | 1037483.4578 | 62.0057 | 149.6951
0.70322 | 1076259.1048 | 74.7136 | 180.3746
0.69865 | 1059145.617 | 72.2222 | 174.3599
0.70118 | 1071070.6319 | 79.7189 | 192.4583
0.70701 | 1088040.2149 | 83.5944 | 201.8147
0.69539 | 1061487.8066 | 67.0451 | 161.8613
0.68271 | 1023751.822 | 61.8669 | 149.3598
0.67375 | 1015005.0078 | 71.7706 | 173.2695
0.67584 | 993690.636 | 71.7758 | 173.2821
0.69963 | 1060474.2277 | 70.3837 | 169.9213
0.68233 | 1016925.919 | 74.4683 | 179.7824
0.69972 | 1058259.4448 | 73.8316 | 178.2453
0.68382 | 1031527.6164 | 85.0342 | 205.2906
0.70836 | 1082234.5195 | 72.7857 | 175.7202
0.72373 | 1102319.7832 | 62.8393 | 151.7074
0.68213 | 1024302.2681 | 69.7262 | 168.334
0.68292 | 1038342.0951 | 75.9706 | 183.4092
0.71961 | 1095054.8503 | 68.7754 | 166.0384
0.71172 | 1074795.4307 | 68.2514 | 164.7733
0.69671 | 1063529.862 | 70.8815 | 171.123
0.68276 | 1032600.6996 | 78.9294 | 190.5525
0.68571 | 1026424.3616 | 84.3797 | 203.7107
0.69589 | 1060031.2917 | 70.7316 | 170.7613
0.7207 | 1086783.4876 | 66.2323 | 159.8989
0.68367 | 1027356.7684 | 71.5905 | 172.8346
0.66628 | 988533.4681 | 68.8924 | 166.3209
0.68631 | 1028783.9746 | 90.3353 | 218.0887
0.70197 | 1068367.9635 | 72.5071 | 175.0477
0.67526 | 1008588.9816 | 76.3088 | 184.2257
0.67653 | 1008416.8905 | 82.5106 | 199.1983
0.69912 | 1075334.7846 | 76.8543 | 185.5428
0.67445 | 1004013.6912 | 67.2799 | 162.428
0.68064 | 1022313.1372 | 75.0491 | 181.1845
0.69706 | 1056176.4703 | 72.1849 | 174.2697
0.70332 | 1071957.6767 | 72.6106 | 175.2974
0.67319 | 1006953.9337 | 88.7217 | 214.1932
0.67934 | 1023374.3985 | 68.0842 | 164.3697
0.69819 | 1050171.4237 | 71.658 | 172.9976
0.69216 | 1045635.7121 | 64.3711 | 155.4057
0.67896 | 1018229.7228 | 65.8696 | 159.0233
0.68133 | 1000621.5416 | 73.3818 | 177.1594
0.70124 | 1068391.3194 | 79.5124 | 191.96
0.67729 | 1028955.5748 | 75.9096 | 183.2621
0.70557 | 1073226.5311 | 70.6304 | 170.5169
0.6978 | 1053324.4696 | 60.1661 | 145.2538
0.70213 | 1071997.4694 | 72.7363 | 175.6009
0.67581 | 1006773.6092 | 72.724 | 175.5713
0.70242 | 1070335.184 | 79.5237 | 191.9873
0.7016 | 1079905.7615 | 76.9085 | 185.6735
0.67385 | 986254.6989 | 79.5024 | 191.9359
0.68873 | 1030384.8332 | 79.6875 | 192.3827
0.68911 | 1037916.4822 | 69.938 | 168.8452
0.70526 | 1091549.9653 | 77.334 | 186.7007
0.69319 | 1044450.6638 | 78.6337 | 189.8385
0.69056 | 1034064.3503 | 85.905 | 207.3931
0.69751 | 1042080.3153 | 71.1756 | 171.833
0.72633 | 1123486.1526 | 79.0253 | 190.7839
0.70222 | 1054581.0268 | 71.5459 | 172.727
0.69824 | 1056502.171 | 78.9655 | 190.6396
0.68888 | 1044055.7189 | 83.8941 | 202.5382
0.66508 | 983152.0912 | 83.3293 | 201.1748
0.70173 | 1077094.3042 | 65.4038 | 157.8986
0.67352 | 989747.1017 | 64.5444 | 155.8239
0.70886 | 1074510.6198 | 79.1086 | 190.985
0.69662 | 1061239.0847 | 76.1193 | 183.7682
0.66691 | 991873.6972 | 82.3266 | 198.7539
0.68425 | 1027045.0336 | 69.6963 | 168.2617
0.68305 | 1027855.6842 | 81.7408 | 197.3398
0.67184 | 995670.2195 | 75.7149 | 182.792
0.68724 | 1030183.3274 | 65.8399 | 158.9515
0.68888 | 1035721.4833 | 64.0642 | 154.6647
0.70406 | 1085111.5152 | 78.7399 | 190.095
0.66937 | 989452.7102 | 62.9862 | 152.0622
0.68273 | 1030950.288 | 80.4321 | 194.1803
0.69122 | 1045843.2512 | 73.8627 | 178.3203
0.69433 | 1046542.3812 | 73.5805 | 177.6391
0.68108 | 1023524.872 | 67.9199 | 163.9732
0.67699 | 1008541.5217 | 88.3631 | 213.3275
0.70354 | 1086108.6177 | 80.5676 | 194.5075
0.70223 | 1064390.4661 | 79.104 | 190.9739
0.70217 | 1054600.4429 | 78.382 | 189.2308
0.69105 | 1049563.0514 | 75.8884 | 183.2107
0.69206 | 1059077.1073 | 86.266 | 208.2645
0.66972 | 998796.9513 | 74.5175 | 179.9013
0.71314 | 1114183.869 | 76.2701 | 184.1323
0.68835 | 1052979.3433 | 69.1359 | 166.9087
0.69863 | 1061318.8928 | 68.6583 | 165.7558
0.69111 | 1034836.1738 | 78.1445 | 188.6576
0.71974 | 1092172.1725 | 70.4621 | 170.1106
0.65413 | 938145.8864 | 74.8478 | 180.6987
0.7116 | 1101411.4003 | 88.7953 | 214.3709
0.71458 | 1087068.2233 | 72.4738 | 174.9673
0.73022 | 1146739.5013 | 64.7908 | 156.4189
0.70301 | 1080150.9532 | 78.4965 | 189.5074
0.69057 | 1043205.3117 | 53.3717 | 128.8507
0.69689 | 1055604.2774 | 80.682 | 194.7836
0.70465 | 1070016.6309 | 74.1298 | 178.9652
0.70089 | 1083363.0985 | 75.4572 | 182.1697
0.70173 | 1048455.5602 | 74.6545 | 180.2318
0.68839 | 1029290.6549 | 72.9927 | 176.22
0.69112 | 1041201.6874 | 69.9546 | 168.8854
0.67029 | 988550.8008 | 66.6553 | 160.9201
0.67648 | 1013929.252 | 71.4048 | 172.3864
0.70856 | 1089759.5356 | 85.135 | 205.5342
0.70387 | 1064051.933 | 64.4151 | 155.5118
0.70442 | 1085526.4011 | 71.5556 | 172.7504
0.65594 | 949976.0574 | 63.6107 | 153.5697
0.68587 | 1033595.5486 | 71.7934 | 173.3245
0.69321 | 1062695.1424 | 65.3221 | 157.7014
0.66134 | 964825.1913 | 76.1378 | 183.8128
0.66742 | 980073.6057 | 60.5842 | 146.2632
0.70383 | 1074471.1676 | 77.9932 | 188.2923
0.678 | 1010109.3487 | 67.3339 | 162.5585
0.67889 | 1015398.6973 | 73.0146 | 176.2728
0.66966 | 995351.8199 | 65.2837 | 157.6089
0.67103 | 986064.9055 | 76.6897 | 185.1454
0.6926 | 1047470.2004 | 86.763 | 209.4644
0.70466 | 1068003.5253 | 66.57 | 160.7142
0.68258 | 1028721.6245 | 86.4605 | 208.7341
0.68522 | 1039087.8183 | 70.4143 | 169.9952
0.70645 | 1078625.4537 | 76.2894 | 184.1789
0.70602 | 1082810.3397 | 78.3544 | 189.1641
0.67925 | 1026768.1198 | 66.3935 | 160.288
0.70532 | 1083333.4072 | 73.1735 | 176.6565
0.67271 | 988852.8331 | 58.5899 | 141.4486
0.71148 | 1080589.178 | 77.5937 | 187.3278
0.68309 | 1016903.3884 | 70.9218 | 171.2203
0.6837 | 1023197.9936 | 82.8414 | 199.9969
0.66292 | 981083.1882 | 66.8871 | 161.4796
0.70193 | 1068047.9282 | 85.7732 | 207.0748
0.68178 | 1010481.8386 | 63.1537 | 152.4666
0.67411 | 999170.1874 | 74.0598 | 178.7961
0.67908 | 1011925.6745 | 69.1935 | 167.0479
0.69905 | 1057041.443 | 66.4148 | 160.3395
0.66677 | 980743.825 | 60.6852 | 146.507
0.64763 | 938293.9566 | 74.0188 | 178.6972
0.70627 | 1066382.4647 | 68.24 | 164.7458
0.67327 | 997608.6758 | 74.0368 | 178.7408
0.67198 | 1017710.616 | 89.8315 | 216.8724
0.67953 | 1018032.0805 | 69.5645 | 167.9436
0.67063 | 1000265.8968 | 57.2869 | 138.3028
0.70379 | 1089008.7316 | 65.9334 | 159.1774
0.68608 | 1045878.6212 | 64.9569 | 156.8199
0.6794 | 1025728.7087 | 71.9256 | 173.6438
0.69718 | 1044334.2891 | 78.1754 | 188.7322
0.67921 | 1023237.9746 | 68.0872 | 164.3771
0.69815 | 1059473.1335 | 64.6637 | 156.112
0.72184 | 1120336.8544 | 79.2941 | 191.4329
0.70595 | 1086388.3786 | 69.5516 | 167.9124
0.68504 | 1028717.4524 | 75.0336 | 181.1472
0.67086 | 995656.9163 | 83.6626 | 201.9793
0.68615 | 1034398.8953 | 68.134 | 164.49
0.69513 | 1063500.7167 | 74.6872 | 180.3109
0.70146 | 1074002.073 | 81.5534 | 196.8874
0.67574 | 996049.7372 | 75.3059 | 181.8044
0.71177 | 1085596.0666 | 79.5983 | 192.1672
0.66647 | 977203.072 | 76.5199 | 184.7355
0.70567 | 1084162.9352 | 75.0011 | 181.0687
0.73169 | 1137470.1315 | 75.8898 | 183.2143
0.67288 | 989698.1838 | 76.7605 | 185.3161
0.69257 | 1050251.5883 | 77.4748 | 187.0406
0.68713 | 1028730.9276 | 76.4883 | 184.659
0.67816 | 1025011.3249 | 70.1836 | 169.4382
0.70612 | 1076868.7789 | 81.2308 | 196.1086
0.68925 | 1045806.6038 | 68.5755 | 165.5558
0.68964 | 1038850.0803 | 64.187 | 154.9612
0.71502 | 1092941.3493 | 70.9772 | 171.3541
0.6885 | 1032312.9199 | 67.765 | 163.5993
0.69472 | 1052210.0837 | 77.3795 | 186.8107
0.65913 | 960375.7503 | 65.8513 | 158.9792
0.66707 | 996118.2972 | 76.2904 | 184.1812
0.68095 | 1010488.8282 | 72.4418 | 174.8899
0.68892 | 1034553.6847 | 63.9879 | 154.4805
0.66107 | 973164.8107 | 66.5326 | 160.624
0.68772 | 1052404.2498 | 74.7038 | 180.3509
0.69313 | 1054160.0714 | 79.3715 | 191.6197
0.67904 | 1016908.2449 | 71.3446 | 172.241
0.68888 | 1046203.2421 | 68.6381 | 165.7071
0.69545 | 1065506.8976 | 69.245 | 167.1723
0.68993 | 1053668.3004 | 63.0089 | 152.1171
0.67378 | 1005712.7196 | 57.3298 | 138.4065
0.68792 | 1046311.9821 | 88.7488 | 214.2585
0.69079 | 1012538.5708 | 60.474 | 145.9971
0.70078 | 1069779.4794 | 65.0471 | 157.0376
0.68388 | 1032375.7183 | 82.0001 | 197.9659
0.66686 | 988947.8845 | 75.7657 | 182.9145
0.71891 | 1116413.1363 | 73.9827 | 178.6102
0.67267 | 998924.6856 | 75.8689 | 183.1638
0.68829 | 1031988.5079 | 84.5192 | 204.0474
0.67759 | 1006774.7545 | 61.3012 | 147.9942
0.68033 | 1020864.6814 | 67.8181 | 163.7273
0.68763 | 1039667.1967 | 82.1285 | 198.2758
0.68338 | 1036534.6896 | 79.1584 | 191.1052
0.67749 | 1019513.416 | 92.0934 | 222.3331
0.70286 | 1065247.5068 | 70.3755 | 169.9015
0.68678 | 1040045.5639 | 86.8784 | 209.743
0.68445 | 1020024.747 | 69.2002 | 167.0641
0.68066 | 1011821.1664 | 75.9455 | 183.3486
0.67961 | 1017733.9922 | 82.8956 | 200.1277
0.69441 | 1046863.4468 | 66.2152 | 159.8576
0.68766 | 1035892.7514 | 66.1647 | 159.7357
0.7028 | 1079188.6989 | 76.7261 | 185.2332
0.67142 | 1002749.2461 | 81.1246 | 195.8522
0.68661 | 1028234.5023 | 61.7846 | 149.1613
0.68831 | 1035496.6324 | 81.7183 | 197.2853
0.6854 | 1038368.68 | 70.3471 | 169.833
0.69644 | 1039432.8858 | 71.1036 | 171.6592
0.67757 | 1013715.3497 | 69.4701 | 167.7156
0.69095 | 1032128.6659 | 82.6884 | 199.6273
0.70351 | 1067991.5477 | 64.6863 | 156.1666
0.69934 | 1059683.2917 | 75.712 | 182.785
0.7003 | 1088224.4435 | 66.2142 | 159.8552
0.66833 | 989754.1066 | 83.4816 | 201.5425
0.68511 | 1046176.0219 | 81.3361 | 196.3627
0.66944 | 983261.7095 | 66.382 | 160.2604
0.6868 | 1031241.4342 | 67.8615 | 163.8321
0.69399 | 1049372.9884 | 70.213 | 169.5092
0.69035 | 1043871.2439 | 72.0367 | 173.912
0.68273 | 1036429.3855 | 80.8597 | 195.2125
0.69836 | 1071481.3033 | 76.6624 | 185.0794
0.70425 | 1079256.4892 | 73.009 | 176.2593
0.67376 | 1005860.5325 | 77.9117 | 188.0956
0.69338 | 1052192.5778 | 72.2943 | 174.5338
0.66865 | 982675.3422 | 62.6217 | 151.1822
0.67921 | 1023110.7802 | 62.1471 | 150.0364
0.6903 | 1058482.3475 | 71.2005 | 171.8932
0.65845 | 972867.7832 | 73.8255 | 178.2305
0.67936 | 998445.1441 | 62.624 | 151.1878
0.70869 | 1087363.783 | 74.9019 | 180.8291
0.69722 | 1065231.1686 | 69.0511 | 166.7041
0.72885 | 1146613.4054 | 80.3729 | 194.0374
0.69373 | 1045111.4886 | 57.8623 | 139.692
0.66445 | 974983.4529 | 69.9053 | 168.7663
0.71138 | 1099692.4835 | 61.6151 | 148.7521
0.68491 | 1026608.7417 | 69.5998 | 168.0287
0.65711 | 961729.4981 | 54.0077 | 130.3862
0.67483 | 1006221.162 | 71.0358 | 171.4955
0.67313 | 996007.4604 | 77.2214 | 186.4289
0.70257 | 1080233.1915 | 63.4349 | 153.1454
0.68337 | 1031919.7687 | 78.8073 | 190.2576
0.69705 | 1063382.6311 | 76.2256 | 184.0248
0.71377 | 1096344.6642 | 74.983 | 181.0251
0.69825 | 1066573.6 | 76.3232 | 184.2604
0.70044 | 1053864.5174 | 69.8637 | 168.6659
0.69603 | 1061958.1146 | 77.9485 | 188.1843
0.69 | 1033868.2146 | 78.1965 | 188.7831
0.67382 | 1015262.5751 | 78.6334 | 189.8378
0.68355 | 1003332.7965 | 66.3394 | 160.1575
0.68684 | 1031708.8363 | 76.7422 | 185.2719
0.72256 | 1110537.8846 | 65.4906 | 158.1083
0.67816 | 1013207.0557 | 72.9353 | 176.0815
0.69044 | 1035864.5598 | 47.945 | 115.7496
0.67817 | 1000131.419 | 64.7737 | 156.3774
0.70317 | 1071032.5969 | 76.1062 | 183.7366
0.71394 | 1092745.1566 | 83.1404 | 200.7187
0.6774 | 1001773.1503 | 58.9446 | 142.305
0.67717 | 1014942.4759 | 74.8889 | 180.7978
0.69611 | 1042614.0995 | 74.8385 | 180.6761
0.6948 | 1032765.839 | 83.791 | 202.2893
0.69648 | 1059632.3643 | 77.7274 | 187.6506
0.67812 | 1009974.0284 | 61.4472 | 148.3466
0.70025 | 1060034.8429 | 64.5039 | 155.7263
0.69455 | 1058461.1761 | 76.1457 | 183.832
0.68493 | 1028099.1432 | 65.9278 | 159.1638
0.681 | 1018521.5124 | 84.3094 | 203.5408
0.71185 | 1085144.8715 | 85.9313 | 207.4566
0.68521 | 1016069.4011 | 85.0697 | 205.3763
0.69541 | 1053451.9866 | 68.4651 | 165.2894
0.70912 | 1072733.5829 | 78.9486 | 190.5987
0.71157 | 1077472.858 | 77.3608 | 186.7656
0.68016 | 1020142.7157 | 57.6072 | 139.0761
0.70434 | 1063685.2679 | 68.2536 | 164.7787
0.67244 | 1001394.7714 | 65.4084 | 157.9099
0.69511 | 1065379.098 | 71.1683 | 171.8154
0.68694 | 1036334.4407 | 86.5479 | 208.9451
0.70881 | 1073442.3646 | 82.2634 | 198.6014
0.70603 | 1102345.2981 | 70.5655 | 170.3603
0.70122 | 1069863.0832 | 74.4816 | 179.8145
0.68341 | 1024322.3255 | 65.8318 | 158.932
0.6924 | 1045290.4251 | 68.23 | 164.7218
0.68597 | 1021072.4498 | 59.3296 | 143.2342
0.70247 | 1049922.6887 | 76.3586 | 184.3459
0.72006 | 1104086.08 | 78.3371 | 189.1226
0.67867 | 1010211.5706 | 80.8552 | 195.2017
0.67901 | 1017172.8047 | 81.3722 | 196.4499
0.68876 | 1028890.9457 | 63.6529 | 153.6716
0.68799 | 1038359.3832 | 78.7663 | 190.1587
0.66061 | 965298.4967 | 62.2635 | 150.3173
0.67994 | 1014718.3384 | 69.2971 | 167.2981
0.69716 | 1067462.8518 | 74.7473 | 180.456
0.66312 | 976837.7373 | 58.7151 | 141.7509
0.69803 | 1065724.0111 | 65.0565 | 157.0602
0.6797 | 1036206.9687 | 68.9006 | 166.3409
0.68204 | 1012370.5977 | 63.248 | 152.6941
0.67005 | 998569.1976 | 84.0081 | 202.8135
0.67558 | 1009212.3306 | 67.5152 | 162.996
0.6746 | 1006730.4394 | 72.8523 | 175.881
0.69293 | 1049614.2338 | 94.094 | 227.163
0.69041 | 1045269.1107 | 75.6482 | 182.6308
0.68952 | 1036757.111 | 71.6323 | 172.9356
0.6923 | 1046695.5323 | 85.2276 | 205.7577
0.71211 | 1107323.1433 | 68.2986 | 164.8874
0.68661 | 1026691.6176 | 65.6771 | 158.5586
0.68295 | 1018584.0176 | 72.8088 | 175.776
0.69582 | 1059233.2946 | 72.5224 | 175.0846
0.66929 | 992142.4784 | 78.9588 | 190.6234
0.70167 | 1066915.2208 | 58.401 | 140.9925
0.67687 | 1012046.7172 | 74.7083 | 180.3618
0.67539 | 1008268.5891 | 76.2695 | 184.1308
0.67353 | 1018715.4487 | 88.2538 | 213.0636
0.70199 | 1095170.6631 | 66.476 | 160.4873
0.66961 | 1014574.2238 | 68.4795 | 165.3241
0.70252 | 1076314.5997 | 86.3635 | 208.4999
0.68506 | 1037010.2534 | 75.6366 | 182.6028
0.68222 | 1024922.1819 | 57.7017 | 139.3043
0.67925 | 1013168.1606 | 61.5397 | 148.57
0.69571 | 1059711.1132 | 81.1337 | 195.874
0.69737 | 1058281.5533 | 75.8251 | 183.058
0.71001 | 1092142.8651 | 76.8584 | 185.5525
0.70293 | 1073247.9879 | 82.936 | 200.2251
0.70514 | 1078979.6854 | 69.1567 | 166.9591
0.67893 | 1007853.9535 | 63.2538 | 152.7081
0.70068 | 1075557.9594 | 76.6995 | 185.1691
0.67594 | 1009911.5143 | 81.1582 | 195.9331
0.69806 | 1055117.0066 | 68.3128 | 164.9217
0.69456 | 1051841.6764 | 77.3699 | 186.7874
0.71713 | 1099835.229 | 65.3499 | 157.7686
0.67193 | 997981.3886 | 65.2354 | 157.4922
0.67005 | 1000761.2461 | 82.676 | 199.5976
0.69496 | 1050556.045 | 64.4131 | 155.5069
0.67429 | 1008929.5457 | 55.2145 | 133.2997
0.67613 | 1006318.6252 | 77.6748 | 187.5237
0.68847 | 1034707.4959 | 69.7861 | 168.4786
0.69242 | 1040456.332 | 71.5866 | 172.8253
0.69124 | 1042542.8588 | 80.3349 | 193.9455
0.69395 | 1058553.5311 | 66.999 | 161.7499
0.70876 | 1091761.4897 | 78.6977 | 189.993
0.70517 | 1070859.1818 | 70.5145 | 170.237
0.69913 | 1068571.826 | 63.0585 | 152.2367
0.72152 | 1122858.9015 | 79.5076 | 191.9483
0.69183 | 1036278.7991 | 69.5525 | 167.9145
0.6906 | 1030910.1191 | 81.0868 | 195.7608
0.69388 | 1038678.6647 | 84.3823 | 203.717
0.65938 | 973062.8477 | 80.1609 | 193.5254
0.68898 | 1029334.7265 | 69.3724 | 167.4799
0.69427 | 1049414.4898 | 68.3152 | 164.9274
0.68526 | 1032909.3916 | 83.4369 | 201.4345
0.68566 | 1033358.705 | 71.7995 | 173.3393
0.69116 | 1047168.3055 | 77.4049 | 186.872
0.70174 | 1060128.9866 | 80.4676 | 194.266
0.66329 | 978643.3285 | 66.6834 | 160.9879
0.66572 | 978364.4303 | 62.6221 | 151.1832
0.65667 | 959502.9126 | 73.009 | 176.2593
0.67383 | 1001080.8823 | 74.2324 | 179.2128
0.68777 | 1034285.5565 | 67.9664 | 164.0854
0.69389 | 1060257.4228 | 76.3475 | 184.3192
0.68754 | 1042758.6988 | 80.7834 | 195.0284
0.67335 | 1012685.4432 | 69.8343 | 168.5949
0.6881 | 1021913.6087 | 67.1209 | 162.0442
0.6943 | 1051530.1359 | 79.0398 | 190.8188
0.68589 | 1043258.0258 | 81.0174 | 195.5934
0.67445 | 1000973.0094 | 85.1421 | 205.5513
0.65756 | 957872.867 | 62.0715 | 149.8539
0.69658 | 1046936.5484 | 83.1961 | 200.8531
0.71529 | 1093175.68 | 79.7588 | 192.5548
0.6737 | 1007294.3782 | 76.3872 | 184.4151
0.70191 | 1071970.1091 | 70.9148 | 171.2035
0.6842 | 1007468.3107 | 54.336 | 131.1787
0.66259 | 982666.6659 | 72.3994 | 174.7876
0.67521 | 1015339.6764 | 67.9624 | 164.0756
0.7195 | 1113672.0473 | 69.9557 | 168.888
0.67241 | 997206.189 | 83.5994 | 201.8269
0.67546 | 1018271.8895 | 64.2361 | 155.0798
0.6807 | 1015362.0373 | 84.0849 | 202.9988
0.6864 | 1016218.2833 | 72.2745 | 174.486
0.68085 | 1025195.3501 | 81.792 | 197.4634
0.6844 | 1038010.1727 | 68.5467 | 165.4863
0.68546 | 1030999.3011 | 75.1546 | 181.4393
0.66721 | 975868.4589 | 71.4003 | 172.3756
0.667 | 995942.5975 | 84.1704 | 203.2052
0.67915 | 1010172.8058 | 65.3295 | 157.7194
0.68722 | 1026387.8889 | 65.5993 | 158.3708
0.69727 | 1060526.2892 | 87.2559 | 210.6543
0.70287 | 1071098.7358 | 62.8721 | 151.7867
0.70344 | 1081337.2 | 76.8085 | 185.432
0.66135 | 989985.216 | 59.6642 | 144.0422
0.70679 | 1074867.9168 | 62.7427 | 151.4742
0.68814 | 1044180.7346 | 75.8185 | 183.0421
0.68801 | 1046310.8652 | 74.24 | 179.2312
0.68537 | 1013857.4674 | 73.9141 | 178.4445
0.69053 | 1031575.8304 | 72.9272 | 176.0619
0.69692 | 1055472.9155 | 79.0379 | 190.8144
0.67768 | 1011059.6167 | 91.3042 | 220.4278
0.68848 | 1058457.9216 | 69.4826 | 167.7458
0.6996 | 1069059.7248 | 70.9561 | 171.3033
0.67817 | 1014105.7524 | 69.4356 | 167.6324
0.70321 | 1066033.712 | 79.1149 | 191.0002
0.67641 | 1017914.5534 | 69.8776 | 168.6995
0.69231 | 1035217.7902 | 76.3941 | 184.4317
0.68352 | 1040231.3855 | 72.5026 | 175.0368
0.70295 | 1071586.4505 | 76.1686 | 183.8874
0.68405 | 1037482.5824 | 80.7853 | 195.0329
0.69394 | 1056625.616 | 74.3314 | 179.4518
0.67282 | 985780.5418 | 71.5241 | 172.6744
0.70667 | 1083233.4433 | 69.279 | 167.2542
0.67346 | 1003940.986 | 72.7731 | 175.6897
0.67782 | 995956.0578 | 66.4077 | 160.3224
0.69104 | 1017645.6212 | 79.9925 | 193.1191
0.66643 | 994390.3486 | 71.0857 | 171.6159
0.69902 | 1042357.7256 | 63.5399 | 153.3988
0.68768 | 1046917.3235 | 68.805 | 166.1099
0.67705 | 1027060.3544 | 74.0544 | 178.783
0.68635 | 1052359.3232 | 84.8733 | 204.9024
0.71887 | 1107670.4056 | 69.6391 | 168.1237
0.70408 | 1084896.9108 | 69.6402 | 168.1263
0.70212 | 1082563.067 | 73.4548 | 177.3355
0.67792 | 1013581.6969 | 87.2399 | 210.6159
0.69246 | 1038082.2331 | 81.4318 | 196.5937
0.68103 | 1014819.6185 | 63.4953 | 153.2912
0.6733 | 1014281.6872 | 75.2708 | 181.7198
0.68865 | 1051249.4333 | 81.3106 | 196.3012
0.68374 | 1031932.2742 | 78.7655 | 190.1568
0.70868 | 1082360.489 | 74.5576 | 179.9981
0.69402 | 1065153.0277 | 80.0963 | 193.3696
0.69118 | 1050962.221 | 75.9752 | 183.4204
0.6901 | 1038782.5226 | 66.1141 | 159.6135
0.71438 | 1084590.8379 | 60.5134 | 146.0924
0.69357 | 1048200.6944 | 70.1766 | 169.4213
0.69998 | 1068523.8542 | 60.8435 | 146.8891
0.65722 | 974319.5687 | 70.6267 | 170.5079
0.68823 | 1037671.0659 | 80.3095 | 193.8844
0.67916 | 1016829.4501 | 81.3264 | 196.3392
0.67906 | 1016840.6295 | 57.5133 | 138.8494
0.68925 | 1033457.0836 | 59.7628 | 144.2802
0.69024 | 1020557.0337 | 75.5387 | 182.3667
0.69105 | 1024400.2185 | 72.3654 | 174.7056
0.68799 | 1027104.908 | 63.6133 | 153.5761
0.67111 | 995322.9437 | 78.9966 | 190.7148
0.6878 | 1039775.6015 | 66.3992 | 160.3017
0.69793 | 1072370.6318 | 63.3923 | 153.0426
0.68087 | 1013482.465 | 77.6833 | 187.5441
0.70509 | 1075334.0327 | 76.2041 | 183.973
0.69431 | 1041336.8752 | 71.4526 | 172.5018
0.69783 | 1054910.622 | 77.2316 | 186.4535
0.69312 | 1046933.5478 | 83.6108 | 201.8544
0.68287 | 1023471.915 | 57.5651 | 138.9744
0.68346 | 1026984.244 | 75.0804 | 181.2601
0.69055 | 1048926.9818 | 66.7162 | 161.0672
0.73022 | 1139903.9863 | 76.2054 | 183.9762
0.69742 | 1060799.753 | 72.4625 | 174.94
0.69711 | 1064192.0463 | 72.7274 | 175.5796
0.68352 | 1011827.936 | 65.0184 | 156.9684
0.69022 | 1029309.2757 | 79.433 | 191.7682
0.68437 | 1022515.823 | 69.3767 | 167.4902
0.70596 | 1064387.3245 | 83.4589 | 201.4875
0.69429 | 1040608.1647 | 69.5955 | 168.0184
0.71025 | 1094050.9079 | 83.6732 | 202.005
0.67526 | 1007939.7789 | 72.8685 | 175.9201
0.71042 | 1091256.2796 | 73.1793 | 176.6704
0.68068 | 1016451.5272 | 76.2909 | 184.1826
0.68643 | 1027903.4779 | 79.7616 | 192.5615
0.69701 | 1052178.2179 | 73.6929 | 177.9105
0.67938 | 1020822.9497 | 71.5835 | 172.8179
0.69541 | 1061897.7807 | 76.2992 | 184.2025
0.71046 | 1091017.3752 | 76.7303 | 185.2433
0.71218 | 1085775.0301 | 72.6315 | 175.3478
0.71098 | 1087479.4932 | 79.8247 | 192.7138
0.68136 | 1018677.1428 | 71.2258 | 171.9542
0.69844 | 1060512.0578 | 72.3373 | 174.6377
0.71741 | 1116382.3183 | 77.7544 | 187.7157
0.72394 | 1120077.4816 | 93.2389 | 225.0986
0.6896 | 1044232.8057 | 68.4379 | 165.2238
0.68655 | 1031878.484 | 67.9281 | 163.9931
0.66422 | 995464.5765 | 65.8255 | 158.9169
0.68993 | 1039829.8525 | 76.9997 | 185.8937
0.6937 | 1044372.2555 | 71.7233 | 173.1555
0.68099 | 1030465.905 | 72.0895 | 174.0395
0.72592 | 1130737.9985 | 68.4156 | 165.1699
0.66819 | 990484.6807 | 82.0538 | 198.0955
0.68846 | 1021942.1703 | 70.7177 | 170.7276
0.67722 | 1018566.3662 | 75.1507 | 181.4299
0.68879 | 1031703.0218 | 63.8239 | 154.0846
0.68648 | 1022353.0286 | 61.4434 | 148.3375
0.70399 | 1084215.8017 | 61.3859 | 148.1987
0.67665 | 1000869.1717 | 68.6835 | 165.8168
0.70874 | 1079590.3505 | 74.3609 | 179.5232
0.69025 | 1030498.6955 | 67.4742 | 162.897
0.67332 | 995236.6336 | 68.635 | 165.6995
0.68672 | 1025727.08 | 52.7714 | 127.4014
0.67979 | 1021761.342 | 69.7927 | 168.4945
0.71556 | 1098891.2767 | 65.4375 | 157.98
0.69342 | 1044482.1916 | 64.1836 | 154.953
0.67933 | 1008804.3738 | 50.0239 | 120.7684
0.67849 | 1016824.5479 | 76.565 | 184.8441
0.68913 | 1046049.5094 | 85.5661 | 206.5747
0.71174 | 1095389.3847 | 76.8991 | 185.6508
0.70095 | 1067495.9564 | 62.8319 | 151.6896
0.69616 | 1057040.0279 | 91.5027 | 220.9071
0.68922 | 1040352.7916 | 88.0759 | 212.634
0.68566 | 1035561.1225 | 71.9523 | 173.7083
0.70609 | 1086669.9925 | 63.4948 | 153.29
0.71933 | 1087162.5147 | 78.8756 | 190.4226
0.66527 | 971325.3787 | 83.7607 | 202.2161
0.70133 | 1051319.1669 | 78.2328 | 188.8707
0.69554 | 1056888.3532 | 75.925 | 183.2991
0.68974 | 1028828.6701 | 74.5308 | 179.9332
0.69493 | 1056923.2782 | 90.2115 | 217.7899
0.68606 | 1032342.6691 | 62.0812 | 149.8773
0.67951 | 1013793.1509 | 63.573 | 153.4787
0.67137 | 992031.4702 | 78.7051 | 190.011
0.6781 | 1018131.5481 | 63.9777 | 154.4558
0.71175 | 1100477.5829 | 71.7946 | 173.3276
0.70162 | 1060349.4293 | 68.3038 | 164.9
0.69524 | 1055705.0264 | 68.6281 | 165.6829
0.71188 | 1096621.0721 | 59.1665 | 142.8406
0.68987 | 1050637.8939 | 76.0956 | 183.711
0.69528 | 1065834.2377 | 73.0308 | 176.312
0.68446 | 1026801.9114 | 73.2342 | 176.8029
0.66695 | 1001411.5533 | 69.7199 | 168.3187
0.69478 | 1045827.9468 | 84.8527 | 204.8525
0.68775 | 1021865.0341 | 74.2407 | 179.2329
0.70441 | 1090343.0474 | 88.7999 | 214.3819
0.68419 | 1018373.671 | 69.1147 | 166.8577
0.65096 | 933067.9606 | 45.8063 | 110.5862
0.68335 | 1016889.2259 | 77.6609 | 187.4901
0.66675 | 978574.1152 | 78.1756 | 188.7326
0.69885 | 1045697.522 | 59.6402 | 143.9841
0.66292 | 968493.0952 | 72.9978 | 176.2323
0.68306 | 1023677.7832 | 76.5661 | 184.8468
0.70384 | 1054042.678 | 63.8122 | 154.0562
0.69371 | 1047652.9311 | 73.7249 | 177.9877
0.67936 | 1034656.9473 | 68.8325 | 166.1765
0.67843 | 1020169.7738 | 55.4895 | 133.9636
0.69774 | 1047690.5132 | 81.5641 | 196.9133
0.69573 | 1048261.9882 | 73.5532 | 177.5731
0.69183 | 1040825.1034 | 78.5145 | 189.5507
0.70992 | 1104135.0255 | 81.1686 | 195.9583
0.69264 | 1044424.8231 | 92.2794 | 222.7823
0.70672 | 1078455.8854 | 72.7001 | 175.5136
0.70533 | 1067853.624 | 72.1273 | 174.1307
0.68879 | 1033467.2754 | 67.8121 | 163.7128
0.68036 | 1013572.6174 | 74.4372 | 179.7073
0.6879 | 1036541.9 | 73.5585 | 177.586
0.68674 | 1024842.6694 | 77.8828 | 188.0256
0.68056 | 1006433.9633 | 68.6687 | 165.7809
0.69254 | 1035788.6604 | 63.7302 | 153.8582
0.68894 | 1046843.8716 | 82.5393 | 199.2676
0.6791 | 1023726.0526 | 80.2949 | 193.8491
0.69144 | 1053335.2507 | 70.738 | 170.7767
0.70079 | 1063640.3517 | 77.5423 | 187.2038
0.68556 | 1037948.8792 | 73.9102 | 178.435
0.69547 | 1061462.9877 | 79.6531 | 192.2996
0.68558 | 1033125.659 | 71.7909 | 173.3185
0.70289 | 1069658.5484 | 74.2906 | 179.3534
0.69011 | 1044474.9082 | 60.3252 | 145.638
0.70481 | 1065592.5561 | 66.4053 | 160.3167
0.71801 | 1119734.0097 | 81.3876 | 196.4871
0.68223 | 1002874.3952 | 68.2576 | 164.7884
0.68975 | 1036729.8922 | 57.7299 | 139.3722
0.67353 | 998096.9303 | 65.0877 | 157.1356
0.68335 | 1020357.0747 | 71.1357 | 171.7368
0.69723 | 1066651.7849 | 79.8396 | 192.7499
0.67834 | 1000167.6513 | 68.4645 | 165.288
0.67665 | 998733.224 | 77.0302 | 185.9673
0.72307 | 1125238.0604 | 84.5321 | 204.0785
0.68692 | 1018950.4201 | 68.879 | 166.2886
0.71945 | 1110721.8225 | 82.2326 | 198.527
0.6947 | 1066124.826 | 65.3921 | 157.8705
0.68752 | 1041779.6202 | 72.1883 | 174.2781
0.68246 | 1036994.5051 | 62.6905 | 151.3482
0.70257 | 1071771.0649 | 60.8041 | 146.7942
0.70967 | 1098882.897 | 61.7181 | 149.0007
0.67197 | 1004218.9064 | 61.6565 | 148.8519
0.69569 | 1040901.3873 | 79.6956 | 192.4021
0.69286 | 1041629.5039 | 73.3908 | 177.181
0.67815 | 1007493.4203 | 85.2292 | 205.7616
0.68991 | 1061760.7515 | 75.129 | 181.3776
0.67571 | 1015408.4659 | 75.0857 | 181.273
0.70006 | 1055557.098 | 66.2261 | 159.8839
0.67092 | 1007340.9056 | 73.2955 | 176.9509
0.70944 | 1084954.8901 | 72.8571 | 175.8926
0.70126 | 1074080.2284 | 86.3394 | 208.4419
0.67575 | 1002105.5523 | 83.5952 | 201.8166
0.7005 | 1048367.4155 | 65.7337 | 158.6953
0.69891 | 1078229.3955 | 77.0845 | 186.0983
0.68033 | 1025559.0058 | 75.321 | 181.841
0.69564 | 1052482.4837 | 75.5977 | 182.5091
0.68858 | 1040505.6101 | 64.3885 | 155.4475
0.67263 | 991650.5502 | 75.3208 | 181.8404
0.69391 | 1043146.7124 | 80.0721 | 193.3111
0.71511 | 1114027.7856 | 72.3368 | 174.6364
0.6637 | 982794.4706 | 72.2881 | 174.519
0.68138 | 1036998.8105 | 64.849 | 156.5592
0.68114 | 1026414.4905 | 84.9453 | 205.076
0.68696 | 1036155.4102 | 62.2293 | 150.2349
0.68927 | 1047124.9912 | 66.7256 | 161.0898
0.70396 | 1080329.434 | 69.3594 | 167.4484
0.70031 | 1071946.0024 | 81.5306 | 196.8323
0.65524 | 961496.8904 | 74.9688 | 180.9908
0.67108 | 1002270.8251 | 52.7667 | 127.39
0.7234 | 1124607.3895 | 68.2821 | 164.8477
0.68935 | 1019553.2166 | 77.1874 | 186.3468
0.69395 | 1050106.9337 | 72.0917 | 174.0448
0.70996 | 1081755.5728 | 58.5222 | 141.2851
0.68116 | 985899.5105 | 84.4543 | 203.8907
0.68273 | 1033816.6624 | 64.5227 | 155.7716
0.68619 | 1044877.2822 | 74.3357 | 179.4623
0.69831 | 1053056.8076 | 68.1557 | 164.5424
0.69095 | 1039690.198 | 78.5293 | 189.5866
0.69829 | 1060063.1089 | 79.2261 | 191.2688
0.70329 | 1066224.4307 | 72.2258 | 174.3685
0.7044 | 1081438.3564 | 78.4351 | 189.3592
0.7074 | 1072610.9483 | 83.9322 | 202.6304
0.70432 | 1079126.8659 | 81.6099 | 197.0236
0.67838 | 1019044.8729 | 71.564 | 172.7707
0.68527 | 1034215.238 | 71.9001 | 173.5821
0.68923 | 1043412.8896 | 83.7999 | 202.3109
0.68565 | 1023214.0783 | 57.3172 | 138.3761
0.70107 | 1082229.7195 | 56.9638 | 137.5227
0.68506 | 1038625.0842 | 66.5904 | 160.7634
0.70645 | 1078466.4915 | 81.3143 | 196.31
0.70738 | 1091352.9085 | 65.7395 | 158.7092
0.69996 | 1060614.3909 | 68.4828 | 165.3321
0.66117 | 975695.2264 | 73.7359 | 178.0142
0.684 | 1014706.8874 | 65.1576 | 157.3043
0.71337 | 1110028.2973 | 72.2535 | 174.4355
0.72323 | 1120368.0669 | 73.8006 | 178.1703
0.70382 | 1066701.3873 | 68.4981 | 165.3689
0.69893 | 1053149.0622 | 92.214 | 222.6242
0.70621 | 1084434.7426 | 67.4454 | 162.8277
0.67979 | 999428.9709 | 79.8608 | 192.801
0.68551 | 1034882.4344 | 71.9209 | 173.6323
0.68415 | 1024271.1989 | 63.7939 | 154.0121
0.69599 | 1063404.7965 | 65.105 | 157.1773
0.69037 | 1023908.1198 | 72.402 | 174.7938
0.69268 | 1048236.8984 | 70.2687 | 169.6436
0.70589 | 1078667.1886 | 69.063 | 166.7328
0.68521 | 1021205.3725 | 85.4989 | 206.4127
0.68927 | 1029714.8321 | 71.1718 | 171.8238
0.70547 | 1086146.8079 | 86.7885 | 209.526
0.67679 | 997163.8255 | 64.7136 | 156.2325
0.6766 | 1007484.3991 | 65.4019 | 157.8943
0.72559 | 1116799.6412 | 76.9064 | 185.6685
0.68581 | 1015223.0685 | 66.9039 | 161.5202
0.69867 | 1061988.1157 | 66.762 | 161.1777
0.683 | 1025237.3116 | 61.9302 | 149.5128
0.6704 | 984906.03 | 60.6125 | 146.3314
0.6871 | 1038036.5444 | 78.843 | 190.3439
0.67697 | 1018318.4147 | 67.9848 | 164.1298
0.694 | 1057314.1033 | 90.3662 | 218.1633
0.70302 | 1073029.7685 | 67.0387 | 161.8456
0.68008 | 1025386.1346 | 70.011 | 169.0215
0.68158 | 1020797.2064 | 69.5445 | 167.8954
0.68932 | 1044518.0417 | 73.9724 | 178.5853
0.65755 | 968512.4758 | 80.0085 | 193.1576
0.68826 | 1023631.9103 | 79.2197 | 191.2533
0.70257 | 1069452.1743 | 53.7798 | 129.8359
0.67973 | 1006664.0749 | 70.5542 | 170.3329
0.6805 | 1018393.294 | 79.7127 | 192.4435
0.68541 | 1027444.0248 | 42.2477 | 101.9949
0.7009 | 1058889.8924 | 73.7512 | 178.0511
0.68147 | 1018960.7142 | 67.5784 | 163.1486
0.70432 | 1079301.577 | 73.0047 | 176.249
0.64408 | 926740.5005 | 68.5486 | 165.4909
0.67407 | 977832.6125 | 69.1094 | 166.8448
0.69085 | 1052293.6795 | 84.3157 | 203.5562
0.69133 | 1038153.1815 | 73.3075 | 176.9799
0.68736 | 1043769.9402 | 84.0575 | 202.9327
0.65759 | 970505.3492 | 77.022 | 185.9475
0.69749 | 1057538.0665 | 70.4036 | 169.9692
0.69535 | 1046026.1091 | 73.4367 | 177.2919
0.69269 | 1045070.2528 | 71.8765 | 173.5251
0.70432 | 1085419.0553 | 82.0859 | 198.1728
0.68568 | 1024100.4459 | 65.6004 | 158.3734
0.69039 | 1032732.2824 | 84.3235 | 203.5749
0.69556 | 1061618.8528 | 76.2628 | 184.1148
0.68462 | 1022695.4077 | 74.6398 | 180.1965
0.68562 | 1029303.6757 | 73.2885 | 176.9341
0.7043 | 1075374.1277 | 74.538 | 179.9507
0.68289 | 1026093.7068 | 76.7659 | 185.3294
0.71332 | 1101716.8808 | 78.9075 | 190.4995
0.69473 | 1034624.1579 | 65.4107 | 157.9154
0.69973 | 1070277.0271 | 76.4266 | 184.5102
0.67516 | 989013.1146 | 68.0196 | 164.2139
0.68139 | 1002004.4991 | 73.4769 | 177.3889
0.67115 | 991987.961 | 72.154 | 174.1952
0.69879 | 1066565.4778 | 67.038 | 161.844
0.70185 | 1075914.9918 | 80.0033 | 193.1451
0.70227 | 1034981.4771 | 70.211 | 169.5043
0.70062 | 1069586.2245 | 74.5259 | 179.9215
0.6741 | 1004494.5918 | 76.1296 | 183.7932
0.67769 | 1002024.0991 | 75.5406 | 182.3713
0.6776 | 1017033.4756 | 78.3309 | 189.1076
0.70973 | 1101949.8059 | 68.1949 | 164.6371
0.68597 | 1028231.8161 | 61.8422 | 149.3004
0.69686 | 1048167.9965 | 76.8446 | 185.5192
0.68016 | 1019246.9064 | 67.7278 | 163.5095
0.68143 | 1020332.3881 | 63.8022 | 154.0321
0.6877 | 1032408.4327 | 63.9686 | 154.4338
0.67321 | 1004821.6368 | 69.3072 | 167.3224
0.68771 | 1033960.8364 | 92.0101 | 222.1319
0.72068 | 1098712.8953 | 64.962 | 156.8322
0.69134 | 1056572.9854 | 69.0727 | 166.7563
0.70156 | 1063553.277 | 85.4572 | 206.312
0.70716 | 1071523.9469 | 88.4139 | 213.4499
0.70184 | 1073537.6979 | 69.1317 | 166.8987
0.69817 | 1056292.8195 | 82.1824 | 198.4058
0.69788 | 1063091.5642 | 64.9096 | 156.7055
0.68953 | 1022079.2049 | 77.3575 | 186.7576
0.7007 | 1067039.9781 | 56.1939 | 135.6641
0.68329 | 1025408.4214 | 71.4942 | 172.6022
0.68225 | 1008446.1106 | 66.8043 | 161.2798
0.67829 | 1031069.836 | 77.6608 | 187.4898
0.68883 | 1038888.0504 | 55.6541 | 134.3609
0.68405 | 1020958.1276 | 71.105 | 171.6628
0.70293 | 1074144.5878 | 78.0691 | 188.4754
0.71209 | 1100906.4155 | 61.6214 | 148.7671
0.68637 | 1016767.5934 | 82.0013 | 197.9686
0.68808 | 1034095.9204 | 69.7848 | 168.4755
0.66585 | 993067.2028 | 65.5062 | 158.1458
0.67476 | 1014299.4782 | 66.0235 | 159.3949
0.69413 | 1053858.5555 | 70.4982 | 170.1976
0.6961 | 1046644.1254 | 66.0408 | 159.4365
0.68883 | 1037231.8392 | 73.509 | 177.4663
0.69274 | 1047282.3118 | 75.5383 | 182.3656
0.68642 | 1040464.9239 | 72.2134 | 174.3385
0.7016 | 1057665.1168 | 71.8187 | 173.3856
0.67798 | 1019012.619 | 78.4403 | 189.3716
0.68794 | 1025021.8541 | 67.3407 | 162.5749
0.69078 | 1045671.2591 | 75.6361 | 182.6018
0.70702 | 1076373.4198 | 84.2719 | 203.4503
0.68848 | 1048503.5332 | 66.8902 | 161.4872
0.68821 | 1031037.6107 | 74.8569 | 180.7204
0.69445 | 1045704.4535 | 86.3862 | 208.5546
0.69349 | 1056713.0306 | 74.3863 | 179.5843
0.71327 | 1113299.8675 | 92.5799 | 223.5076
0.6594 | 966874.5381 | 68.0863 | 164.3748
0.69814 | 1072868.5677 | 76.7853 | 185.3761
0.68734 | 1029326.3221 | 69.5373 | 167.8779
0.70462 | 1073829.7181 | 51.337 | 123.9385
0.68456 | 1035883.2093 | 95.3704 | 230.2445
0.68452 | 1031581.5087 | 68.1101 | 164.4323
0.69201 | 1054277.5041 | 88.8949 | 214.6112
0.68748 | 1052312.3054 | 78.201 | 188.7939
0.6978 | 1064451.1863 | 84.4397 | 203.8556
0.65704 | 968951.0146 | 86.0487 | 207.7401
0.67859 | 1015313.0308 | 74.7596 | 180.4857
0.6895 | 1038042.7007 | 98.8697 | 238.6927
0.70201 | 1056196.8493 | 78.5195 | 189.5628
0.68512 | 1014332.3571 | 72.2743 | 174.4856
0.68156 | 1017019.0812 | 76.2348 | 184.0472
0.69922 | 1078005.8767 | 67.7162 | 163.4814
0.71221 | 1090507.3497 | 54.096 | 130.5993
0.69737 | 1060949.3153 | 68.7589 | 165.9987
0.69422 | 1065795.7206 | 73.1398 | 176.5752
0.69749 | 1100380.3519 | 71.3018 | 172.1378
0.69628 | 1085869.4495 | 76.1321 | 183.7991
0.67495 | 1002512.7888 | 75.8839 | 183.2
0.68765 | 1031475.8002 | 76.6346 | 185.0124
0.68697 | 1026406.351 | 82.7058 | 199.6693
0.67941 | 1016271.8573 | 87.28 | 210.7125
0.70249 | 1076658.9766 | 84.8388 | 204.819
0.68666 | 1025822.5393 | 74.6462 | 180.2118
0.6943 | 1056728.3367 | 66.8498 | 161.3896
0.67507 | 1002512.0216 | 83.7134 | 202.102
0.67366 | 1011158.8948 | 65.9262 | 159.1599
0.67064 | 996735.5035 | 63.3295 | 152.8908
0.69992 | 1074706.0547 | 71.2986 | 172.13
0.691 | 1048348.3068 | 75.806 | 183.012
0.70029 | 1061013.1592 | 77.4053 | 186.8729
0.70339 | 1061714.2271 | 55.6997 | 134.4709
0.70093 | 1058384.7164 | 78.8033 | 190.248
0.68484 | 1018050.0925 | 65.2925 | 157.63
0.69915 | 1064026.7462 | 68.7824 | 166.0554
0.66676 | 977157.2141 | 68.0505 | 164.2885
0.71312 | 1086339.3452 | 70.7362 | 170.7724
0.70375 | 1069162.4961 | 79.2866 | 191.4147
0.65571 | 965026.17 | 56.5262 | 136.4662
0.68895 | 1056190.3264 | 67.3342 | 162.5592
0.69952 | 1060917.7057 | 55.9656 | 135.1128
0.69346 | 1059609.5969 | 79.7441 | 192.5193
0.68651 | 1046636.5289 | 90.4296 | 218.3165
0.68119 | 1027540.601 | 77.0658 | 186.0534
0.72054 | 1101263.7258 | 76.3672 | 184.3667
0.71466 | 1099140.0956 | 85.135 | 205.534
0.708 | 1098919.56 | 70.15 | 169.3572
0.6908 | 1048902.3569 | 60.5199 | 146.108
0.69432 | 1059835.2042 | 74.5159 | 179.8974
0.71789 | 1111273.8028 | 69.9614 | 168.9018
0.68117 | 1024425.0601 | 71.6481 | 172.9739
0.69904 | 1047111.554 | 71.362 | 172.2831
0.67684 | 1004579.1475 | 68.9183 | 166.3835
0.68779 | 1056795.3512 | 75.0079 | 181.085
0.68985 | 1035640.9742 | 69.7094 | 168.2934
0.69605 | 1062457.2275 | 65.041 | 157.0229
0.69985 | 1070791.6805 | 83.1107 | 200.6471
0.67477 | 1002286.9028 | 75.8864 | 183.2059
0.67627 | 1006168.4512 | 57.1442 | 137.9583
0.70416 | 1060334.9808 | 64.1632 | 154.9038
0.68509 | 1031497.6997 | 62.1187 | 149.9678
0.67375 | 1008740.8352 | 59.6762 | 144.0711
0.66303 | 969078.4821 | 58.2565 | 140.6437
0.68671 | 1026224.9463 | 63.6776 | 153.7313
0.67943 | 1006009.9962 | 80.4293 | 194.1735
0.67913 | 1007032.6424 | 78.9823 | 190.6801
0.6758 | 1016854.2359 | 80.1457 | 193.4888
0.71571 | 1092542.4174 | 76.2745 | 184.143
0.68611 | 1032469.5801 | 78.5334 | 189.5965
0.69518 | 1049379.8458 | 78.7784 | 190.1878
0.70085 | 1071554.6151 | 84.6335 | 204.3233
0.7023 | 1075590.2089 | 79.9274 | 192.9618
0.69902 | 1059525.9456 | 79.104 | 190.9739
0.67037 | 1000988.4029 | 79.5347 | 192.0139
0.70631 | 1073505.2993 | 74.1959 | 179.1248
0.68293 | 1009919.804 | 64.9517 | 156.8072
0.66936 | 991011.7938 | 61.0825 | 147.4663
0.69784 | 1060786.6812 | 72.4185 | 174.8337
0.67683 | 988614.5284 | 60.2626 | 145.4867
0.67229 | 1007979.2669 | 65.8537 | 158.9848
0.69444 | 1051435.389 | 55.871 | 134.8846
0.68579 | 1037429.253 | 68.3691 | 165.0577
0.70006 | 1066140.549 | 88.8626 | 214.5332
0.68924 | 1045567.6384 | 70.4402 | 170.0577
0.65532 | 949107.611 | 65.4571 | 158.0275
0.69078 | 1035646.5847 | 71.425 | 172.4352
0.69966 | 1054794.299 | 68.1359 | 164.4946
0.70805 | 1098455.8805 | 79.2595 | 191.3495
0.68453 | 1028181.2843 | 75.0504 | 181.1876
0.70228 | 1071165.7309 | 58.5832 | 141.4324
0.70617 | 1081513.768 | 74.765 | 180.4986
0.68661 | 1027355.4275 | 57.469 | 138.7423
0.7018 | 1065864.7068 | 77.5923 | 187.3243
0.6779 | 1012709.9612 | 84.2114 | 203.3043
0.71047 | 1100045.6131 | 85.7077 | 206.9168
0.68168 | 1021964.9259 | 65.8473 | 158.9693
0.66707 | 984377.7701 | 77.0125 | 185.9247
0.68245 | 1026548.3136 | 80.1211 | 193.4293
0.67821 | 1009530.9029 | 80.8558 | 195.2033
0.69605 | 1065673.8595 | 69.1954 | 167.0525
0.6735 | 1011742.5477 | 71.5668 | 172.7776
0.68856 | 1021761.5547 | 65.176 | 157.3487
0.68736 | 1034691.7889 | 55.5477 | 134.1041
0.6699 | 1000782.8577 | 76.3242 | 184.263
0.687 | 1013952.9561 | 70.3131 | 169.7509
0.7024 | 1068640.1119 | 68.7608 | 166.0033
0.69406 | 1052773.5958 | 82.394 | 198.9167
0.66726 | 991700.8433 | 72.5536 | 175.1598
0.70507 | 1070812.5436 | 94.341 | 227.7594
0.70304 | 1072679.4987 | 81.2702 | 196.2037
0.67593 | 1014917.7149 | 81.5616 | 196.9071
0.69511 | 1053499.9801 | 79.0965 | 190.9558
0.67079 | 979807.9493 | 60.2637 | 145.4895
0.67943 | 1019570.2689 | 59.8126 | 144.4003
0.69395 | 1050714.6399 | 83.8124 | 202.341
0.68472 | 1024993.3397 | 63.7859 | 153.9927
0.70074 | 1061027.2934 | 72.1354 | 174.1504
0.66541 | 976125.3392 | 73.2982 | 176.9576
0.68475 | 1035298.3235 | 66.4073 | 160.3215
0.69261 | 1047849.1918 | 66.4501 | 160.4246
0.68512 | 1020769.091 | 70.9808 | 171.3627
0.71366 | 1093543.2938 | 71.1512 | 171.7741
0.66758 | 980200.6935 | 75.0141 | 181.1001
0.69164 | 1032414.8797 | 61.585 | 148.6793
0.67207 | 993685.2 | 72.1781 | 174.2535
0.69651 | 1056026.0154 | 66.5934 | 160.7708
0.65435 | 957054.2101 | 73.0811 | 176.4335
0.70009 | 1038478.6045 | 63.2059 | 152.5925
0.673 | 992365.7108 | 64.7647 | 156.3557
0.70394 | 1068596.9551 | 80.8571 | 195.2062
0.6697 | 990214.5515 | 66.8261 | 161.3326
0.71193 | 1094423.4435 | 80.2275 | 193.6863
0.70043 | 1070776.8624 | 69.0919 | 166.8026
0.69961 | 1049899.0313 | 67.8243 | 163.7425
0.69106 | 1043076.4424 | 68.2177 | 164.6921
0.71314 | 1079931.8088 | 82.2584 | 198.5894
0.68659 | 1029782.8249 | 71.4536 | 172.5042
0.69066 | 1049801.8197 | 68.697 | 165.8493
0.71719 | 1110584.1197 | 75.5948 | 182.5019
0.68106 | 1016951.4027 | 71.1366 | 171.7391
0.71172 | 1089313.9279 | 81.5233 | 196.8147
0.70056 | 1052173.3283 | 90.5096 | 218.5095
0.68457 | 1027426.9613 | 73.2668 | 176.8817
0.69599 | 1065637.8716 | 72.9429 | 176.0997
0.6929 | 1062597.133 | 71.2096 | 171.9153
0.69621 | 1047333.429 | 87.7774 | 211.9134
0.70252 | 1073418.9955 | 74.5907 | 180.0779
0.70445 | 1078848.0953 | 67.2657 | 162.3938
0.69656 | 1053346.9567 | 88.2367 | 213.0223
0.68751 | 1044353.7971 | 87.1445 | 210.3854
0.69685 | 1059981.261 | 72.4663 | 174.949
0.68694 | 1021871.9169 | 69.0744 | 166.7604
0.68022 | 1024243.2914 | 76.9275 | 185.7194
0.67944 | 1004343.0018 | 78.7358 | 190.0851
0.70457 | 1077177.3297 | 74.6344 | 180.1833
0.69125 | 1044280.2035 | 70.149 | 169.3546
0.69422 | 1053362.2412 | 68.5766 | 165.5586
0.72442 | 1135468.3926 | 72.7817 | 175.7107
0.671 | 988748.9657 | 62.7889 | 151.5857
0.68225 | 1037534.6243 | 79.3658 | 191.606
0.68842 | 1039938.343 | 62.3237 | 150.4626
0.65902 | 982736.4868 | 71.0424 | 171.5115
0.7059 | 1077029.1527 | 81.2685 | 196.1994
0.67601 | 1010761.4648 | 71.2317 | 171.9686
0.68724 | 1009422.5287 | 66.3837 | 160.2645
0.7124 | 1115849.9752 | 89.8578 | 216.936
0.67805 | 1014482.8117 | 63.7533 | 153.9141
0.70851 | 1083615.0337 | 84.6607 | 204.3889
0.68568 | 1017141.716 | 59.1261 | 142.7429
0.6986 | 1066280.346 | 72.7143 | 175.5478
0.68301 | 1018596.5143 | 74.5166 | 179.899
0.67356 | 1008868.6221 | 69.5503 | 167.9092
0.7105 | 1109422.6392 | 76.2264 | 184.0269
0.67308 | 1000817.8523 | 65.3088 | 157.6694
0.68945 | 1047167.2891 | 75.8908 | 183.2167
0.68308 | 1021145.8317 | 63.6986 | 153.7821
0.67341 | 980559.5463 | 77.5597 | 187.2457
0.69782 | 1065022.6253 | 76.2586 | 184.1046
0.68851 | 1043941.9045 | 74.018 | 178.6953
0.71024 | 1089609.3949 | 61.1484 | 147.6252
0.67797 | 997199.7403 | 71.8038 | 173.3496
0.69461 | 1047417.1234 | 75.3451 | 181.8992
0.6836 | 1028684.9529 | 68.0109 | 164.1929
0.6892 | 1036204.3772 | 76.812 | 185.4406
0.68932 | 1037701.9454 | 62.1084 | 149.9429
0.68612 | 1045188.1934 | 66.9074 | 161.5288
0.70576 | 1084163.019 | 80.1694 | 193.546
0.69747 | 1078971.2275 | 65.9178 | 159.1396
0.69206 | 1034005.8708 | 67.5266 | 163.0237
0.69711 | 1044739.9995 | 76.5658 | 184.8461
0.6969 | 1064326.6726 | 79.7035 | 192.4212
0.66805 | 996411.3655 | 59.4494 | 143.5234
0.68961 | 1045464.8101 | 64.5983 | 155.954
0.67237 | 1018015.1973 | 72.1314 | 174.1407
0.69396 | 1065608.1023 | 76.7251 | 185.2307
0.7018 | 1082370.1349 | 73.4755 | 177.3855
0.70041 | 1060498.363 | 77.3546 | 186.7504
0.69592 | 1048345.3325 | 78.6169 | 189.7979
0.68042 | 1028614.472 | 86.7261 | 209.3753
0.72039 | 1116417.8993 | 71.9608 | 173.7288
0.7025 | 1081443.044 | 67.6143 | 163.2354
0.67066 | 991380.1207 | 66.8084 | 161.2897
0.70425 | 1061028.5202 | 64.1975 | 154.9865
0.65951 | 977485.5324 | 70.61 | 170.4676
0.71603 | 1113693.7273 | 84.4871 | 203.9699
0.70692 | 1078635.244 | 70.8508 | 171.0489
0.71318 | 1079642.0055 | 64.5053 | 155.7296
0.68508 | 1046772.2575 | 74.5901 | 180.0765
0.70026 | 1075584.2609 | 93.9987 | 226.933
0.67328 | 1006069.2938 | 65.6319 | 158.4495
0.65985 | 975051.4137 | 77.1182 | 186.1798
0.6681 | 976324.8589 | 77.838 | 187.9176
0.67525 | 1018336.3895 | 71.8459 | 173.4513
0.71061 | 1058624.0316 | 74.4365 | 179.7055
0.70976 | 1083706.2126 | 72.7402 | 175.6103
0.68352 | 1019467.1052 | 67.4086 | 162.7389
0.6914 | 1038684.7863 | 71.5703 | 172.7859
0.69173 | 1044669.1057 | 62.3079 | 150.4246
0.69631 | 1064831.5333 | 78.1286 | 188.6192
0.67387 | 990353.3543 | 78.2803 | 188.9853
0.69072 | 1029148.8302 | 69.1854 | 167.0284
0.71166 | 1106300.5166 | 77.6739 | 187.5214
0.66315 | 971635.2586 | 66.2194 | 159.8678
0.66636 | 976576.4918 | 62.4473 | 150.761
0.68835 | 1052856.4748 | 78.1434 | 188.6548
0.7 | 1075167.4193 | 65.4478 | 158.005
0.69667 | 1050733.2104 | 78.7121 | 190.0279
0.70306 | 1052684.8409 | 67.671 | 163.3723
0.67174 | 997684.546 | 79.1105 | 190.9897
0.69377 | 1043750.7023 | 69.8587 | 168.6539
0.68295 | 1031708.636 | 70.1467 | 169.3491
0.69088 | 1052460.5621 | 79.5017 | 191.9342
0.68617 | 1030515.3103 | 73.3151 | 176.9983
0.67852 | 1007641.8632 | 65.227 | 157.4719
0.7076 | 1084912.4275 | 83.2841 | 201.0657
0.68988 | 1038532.58 | 69.9015 | 168.7572
0.67196 | 1000747.5533 | 61.4448 | 148.3408
0.69137 | 1044350.8562 | 71.9636 | 173.7355
0.68715 | 1038850.892 | 75.5402 | 182.3703
0.68945 | 1038102.4747 | 75.2345 | 181.6322
0.71068 | 1086185.811 | 81.408 | 196.5364
0.68536 | 1031159.2914 | 73.4477 | 177.3184
0.71243 | 1092115.6205 | 67.3297 | 162.5483
0.69897 | 1076965.6089 | 69.637 | 168.1187
0.68505 | 1032669.0891 | 62.193 | 150.1471
0.69244 | 1041071.2705 | 78.4644 | 189.4298
0.6946 | 1057307.6211 | 62.8008 | 151.6146
0.69197 | 1050764.3562 | 75.8741 | 183.1762
0.68693 | 1036090.3692 | 75.364 | 181.9449
0.68639 | 1026953.4477 | 79.693 | 192.396
0.69154 | 1053607.6856 | 62.858 | 151.7526
0.66223 | 980555.8327 | 65.0353 | 157.009
0.73207 | 1140512.5727 | 86.238 | 208.1969
0.69122 | 1033543.7637 | 74.7114 | 180.3692
0.70713 | 1080068.637 | 72.7662 | 175.6731
0.68322 | 1029874.8402 | 78.6676 | 189.9205
0.69615 | 1052119.2374 | 83.6595 | 201.9719
0.70019 | 1071601.2497 | 69.1695 | 166.9901
0.68111 | 1027332.3503 | 83.4922 | 201.568
0.69161 | 1032482.915 | 61.5916 | 148.6952
0.69222 | 1054819.8382 | 90.3841 | 218.2065
0.71466 | 1102628.728 | 67.7657 | 163.6008
0.68021 | 1035467.9935 | 68.8 | 166.0978
0.69854 | 1061260.5602 | 62.6805 | 151.3241
0.66817 | 981517.1339 | 78.2442 | 188.8982
0.67382 | 1004366.753 | 88.4111 | 213.4434
0.69157 | 1049251.897 | 68.5709 | 165.5448
0.68549 | 1051298.6344 | 71.2136 | 171.9248
0.72004 | 1114397.7007 | 69.9491 | 168.872
0.69367 | 1048882.3712 | 64.4305 | 155.549
0.67605 | 1002112.0001 | 78.955 | 190.6143
0.68226 | 1023837.8594 | 59.8131 | 144.4015
0.67851 | 1011065.7977 | 86.948 | 209.911
0.6856 | 1028808.2995 | 84.589 | 204.2159
0.6991 | 1079625.8936 | 65.8038 | 158.8645
0.70184 | 1066093.4619 | 74.1535 | 179.0224
0.70067 | 1071799.2494 | 93.613 | 226.0017
0.69218 | 1061123.9871 | 83.2729 | 201.0386
0.6831 | 1021124.6424 | 76.3708 | 184.3755
0.70289 | 1072362.2238 | 56.2986 | 135.9168
0.68265 | 1022180.4356 | 59.6828 | 144.087
0.69458 | 1042646.2566 | 79.9561 | 193.0311
0.65965 | 964262.7189 | 67.8762 | 163.8676
0.71232 | 1086687.5568 | 77.9399 | 188.1635
0.6891 | 1057461.8174 | 80.0858 | 193.3443
0.69319 | 1045351.0369 | 71.8439 | 173.4466
0.67552 | 1013467.6259 | 63.9273 | 154.3342
0.68005 | 1018023.428 | 71.3174 | 172.1754
0.7075 | 1077712.5755 | 82.7434 | 199.7601
0.71298 | 1096907.8473 | 71.4383 | 172.4674
0.6972 | 1053931.6082 | 60.2523 | 145.4619
0.68769 | 1028950.1875 | 85.8372 | 207.2292
0.69589 | 1051059.9211 | 77.1642 | 186.2909
0.68917 | 1042376.1718 | 73.8367 | 178.2576
0.70748 | 1089660.4445 | 82.0383 | 198.0579
0.6849 | 1032183.9508 | 71.7974 | 173.3341
0.68688 | 1036597.3523 | 90.9463 | 219.5639
0.69432 | 1056121.8704 | 82.6246 | 199.4734
0.6753 | 999834.3151 | 66.9155 | 161.5483
0.69441 | 1058150.0354 | 70.7573 | 170.8232
0.69957 | 1058032.086 | 68.3539 | 165.021
0.68507 | 1024585.1704 | 76.4087 | 184.4668
0.67417 | 997313.7075 | 73.9165 | 178.4502
0.7013 | 1069375.7026 | 74.4393 | 179.7123
0.6828 | 1026012.0188 | 87.679 | 211.6757
0.68379 | 1020264.0515 | 72.8384 | 175.8474
0.67862 | 997304.0948 | 78.2785 | 188.981
0.70063 | 1068113.4502 | 77.8859 | 188.0332
0.70043 | 1064202.7165 | 74.1705 | 179.0633
0.70419 | 1074702.3638 | 58.6168 | 141.5135
0.6848 | 1022913.5301 | 76.3146 | 184.2398
0.70345 | 1070334.8123 | 77.7351 | 187.6691
0.6715 | 1005109.8897 | 63.7799 | 153.9782
0.69502 | 1058510.2598 | 81.1218 | 195.8454
0.68641 | 1044249.6606 | 72.4766 | 174.9741
0.68764 | 1033140.8254 | 69.7649 | 168.4275
0.67027 | 983609.1619 | 78.1651 | 188.7073
0.68151 | 1032670.0092 | 70.3605 | 169.8653
0.68262 | 1015386.0109 | 74.3243 | 179.4348
0.70527 | 1065620.469 | 75.4955 | 182.2623
0.67243 | 986980.4002 | 58.5032 | 141.2393
0.68381 | 1012128.7343 | 70.5148 | 170.2378
0.65855 | 951449.5551 | 63.4572 | 153.1991
0.68577 | 1041931.0513 | 66.7338 | 161.1097
0.69804 | 1066693.8033 | 78.0716 | 188.4814
0.6926 | 1051237.0468 | 53.6246 | 129.4613
0.67599 | 1002146.1464 | 86.5024 | 208.8353
0.66819 | 990108.987 | 70.5925 | 170.4254
0.66531 | 972363.9482 | 66.0547 | 159.4701
0.69943 | 1062435.9419 | 85.6267 | 206.7212
0.70424 | 1086093.9726 | 75.2872 | 181.7593
0.69403 | 1040066.5954 | 74.8787 | 180.7732
0.70382 | 1067245.6483 | 77.47 | 187.0292
0.67571 | 1009641.1155 | 81.1645 | 195.9484
0.69263 | 1044045.9783 | 64.2298 | 155.0644
0.68809 | 1053299.0854 | 63.155 | 152.4697
0.69061 | 1050750.2259 | 93.1118 | 224.7917
0.68963 | 1056257.633 | 54.1129 | 130.64
0.70467 | 1086016.8844 | 78.7372 | 190.0884
0.69416 | 1073039.4659 | 77.238 | 186.4691
0.6848 | 1011988.3589 | 72.1415 | 174.1651
0.70093 | 1075238.0205 | 81.7094 | 197.264
0.67901 | 1019023.1825 | 91.8432 | 221.7292
0.68845 | 1034059.9106 | 59.8672 | 144.5321
0.67711 | 1000876.6303 | 73.8585 | 178.3102
0.69074 | 1036888.6294 | 78.3415 | 189.1331
0.7009 | 1063914.141 | 70.4838 | 170.1629
0.69853 | 1067812.4642 | 84.6182 | 204.2864
0.70943 | 1094433.1411 | 73.4947 | 177.4319
0.69858 | 1062839.9341 | 66.8247 | 161.329
0.69349 | 1044907.7198 | 81.8461 | 197.5938
0.67402 | 997835.7511 | 83.4521 | 201.4712
0.67801 | 1017789.7735 | 61.933 | 149.5194
0.69712 | 1046684.8126 | 78.3579 | 189.1727
0.69854 | 1075016.8174 | 67.5939 | 163.1862
0.68251 | 1023588.2422 | 52.9858 | 127.919
0.69282 | 1064864.368 | 69.4658 | 167.7053
0.70158 | 1071517.9734 | 82.9621 | 200.2881
0.68641 | 1019658.2382 | 63.2981 | 152.8152
0.70173 | 1066015.3074 | 64.7311 | 156.2746
0.67902 | 1009526.163 | 77.5907 | 187.3205
0.71454 | 1087725.8341 | 72.5223 | 175.0844
0.7048 | 1063629.3985 | 64.3046 | 155.245
0.67687 | 1015789.8373 | 67.8261 | 163.7466
0.69039 | 1039415.0281 | 80.1931 | 193.6032
0.67753 | 1016446.3621 | 74.9601 | 180.9697
0.7061 | 1073200.5082 | 79.3213 | 191.4986
0.6792 | 1016575.8013 | 61.7509 | 149.0798
0.68378 | 1017039.6285 | 70.8594 | 171.0698
0.69901 | 1068177.0701 | 81.3957 | 196.5067
0.70367 | 1060514.3749 | 68.115 | 164.4441
0.69971 | 1063390.696 | 66.7513 | 161.1519
0.66866 | 980463.0591 | 68.0868 | 164.3762
0.6849 | 1047717.3907 | 71.0324 | 171.4874
0.7033 | 1081044.7093 | 67.5129 | 162.9905
0.69861 | 1066554.1967 | 87.5838 | 211.4461
0.65578 | 963638.5394 | 71.7346 | 173.1827
0.69029 | 1032625.245 | 74.7458 | 180.4523
0.69379 | 1045039.9722 | 64.183 | 154.9514
0.71648 | 1086663.4905 | 75.7822 | 182.9543
0.70153 | 1069713.9288 | 74.4376 | 179.7083
0.67631 | 1000089.1646 | 83.3573 | 201.2423
0.69184 | 1029354.1272 | 70.8819 | 171.1241
0.70802 | 1062333.1807 | 78.3043 | 189.0434
0.6587 | 965498.0014 | 68.061 | 164.3139
0.68593 | 1035881.4655 | 75.4581 | 182.172
0.65643 | 954382.3178 | 81.9576 | 197.8633
0.69789 | 1080494.5342 | 77.3597 | 186.7629
0.70343 | 1087750.7166 | 79.0017 | 190.7271
0.68956 | 1044024.0501 | 82.2422 | 198.5502
0.70266 | 1098784.6366 | 88.1492 | 212.8111
0.69709 | 1051983.5161 | 66.2831 | 160.0215
0.68392 | 1019710.4214 | 65.487 | 158.0995
0.68736 | 1022425.9737 | 72.6145 | 175.3069
0.6894 | 1033722.2982 | 93.1434 | 224.8681
0.70346 | 1065643.5794 | 73.9512 | 178.5341
0.67845 | 1008929.1378 | 69.0529 | 166.7084
0.69112 | 1047879.6566 | 74.0171 | 178.6932
0.67654 | 998935.9564 | 58.8524 | 142.0823
0.68921 | 1030343.8622 | 65.6615 | 158.5208
0.68967 | 1050386.8965 | 64.8968 | 156.6747
0.72602 | 1123198.0105 | 59.4195 | 143.4514
0.69191 | 1049433.5315 | 79.1547 | 191.0964
0.67758 | 994379.9902 | 68.1448 | 164.5161
0.66884 | 993233.3754 | 65.2249 | 157.4667
0.69727 | 1045020.3617 | 73.1644 | 176.6344
0.69489 | 1063217.71 | 71.047 | 171.5227
0.68988 | 1038947.4493 | 71.7456 | 173.2091
0.70135 | 1076816.3811 | 71.2924 | 172.115
0.69251 | 1044977.1581 | 50.8314 | 122.7177
0.69289 | 1042021.821 | 76.8319 | 185.4887
0.70557 | 1065281.7996 | 62.2908 | 150.3832
0.70358 | 1057562.453 | 66.0197 | 159.3857
0.68485 | 1021805.1588 | 65.5322 | 158.2088
0.69535 | 1054052.0117 | 70.3874 | 169.9302
0.7044 | 1065710.5657 | 64.467 | 155.6372
0.6814 | 1037384.3059 | 70.6926 | 170.667
0.67364 | 1012136.2931 | 67.9062 | 163.9401
0.70867 | 1086257.2131 | 61.0948 | 147.4958
0.68851 | 1050949.837 | 82.9145 | 200.1733
0.67751 | 1008338.6561 | 69.1081 | 166.8416
0.72587 | 1120373.7907 | 63.68 | 153.7372
0.68185 | 1017621.2928 | 72.857 | 175.8923
0.68956 | 1045763.8899 | 64.6212 | 156.0094
0.67625 | 999569.7702 | 77.1395 | 186.2313
0.67708 | 1023865.904 | 60.2726 | 145.511
0.67153 | 990264.1131 | 70.2524 | 169.6043
0.68355 | 1031959.966 | 86.6097 | 209.0943
0.69016 | 1038992.7634 | 64.1472 | 154.865
0.71686 | 1100339.3057 | 67.4177 | 162.7607
0.70312 | 1061662.923 | 75.4453 | 182.1411
0.6858 | 1038708.0559 | 89.8371 | 216.886
0.70189 | 1049427.3455 | 71.2301 | 171.9647
0.69063 | 1045181.8131 | 77.3171 | 186.6601
0.68889 | 1031314.348 | 63.7912 | 154.0056
0.70424 | 1085709.3462 | 78.7629 | 190.1505
0.68675 | 1026385.2246 | 67.5473 | 163.0736
0.71039 | 1094798.1414 | 77.0037 | 185.9034
0.70131 | 1073217.3634 | 63.1366 | 152.4253
0.68177 | 1040097.3934 | 74.4754 | 179.7994
0.68524 | 1048180.7452 | 69.1336 | 166.9034
0.69047 | 1035777.3833 | 76.6392 | 185.0234
0.70716 | 1075611.5577 | 79.3796 | 191.6393
0.68342 | 1025751.1521 | 81.529 | 196.8284
0.71295 | 1109457.1791 | 95.2334 | 229.9137
0.67447 | 985451.9127 | 69.6065 | 168.0449
0.67115 | 1009670.0172 | 76.8586 | 185.5531
0.71665 | 1103843.8777 | 56.8381 | 137.2192
0.69894 | 1073037.8376 | 84.9821 | 205.1649
0.69685 | 1053768.7682 | 75.5585 | 182.4143
0.67957 | 1029619.1538 | 72.181 | 174.2604
0.69193 | 1041290.0468 | 71.6947 | 173.0863
0.69188 | 1052021.0185 | 79.3542 | 191.5779
0.70864 | 1087927.3869 | 81.7734 | 197.4183
0.70641 | 1070506.299 | 80.7177 | 194.8697
0.71895 | 1109245.1881 | 75.3181 | 181.8341
0.6803 | 998407.9234 | 68.1464 | 164.52
0.68401 | 1019894.6153 | 85.5398 | 206.5113
0.70378 | 1079117.5775 | 76.6145 | 184.9637
0.6676 | 999130.7259 | 73.0596 | 176.3815
0.68599 | 1030275.5526 | 75.0433 | 181.1705
0.67744 | 1032965.1167 | 76.1884 | 183.9352
0.69065 | 1053544.3794 | 77.9334 | 188.1478
0.70043 | 1059109.2964 | 73.8157 | 178.207
0.66547 | 980254.0322 | 80.0146 | 193.1723
0.68509 | 1033887.1506 | 81.9751 | 197.9054
0.70358 | 1070886.7555 | 72.5116 | 175.0585
0.69783 | 1063903.6076 | 66.9631 | 161.6632
0.68632 | 1026073.4967 | 74.7017 | 180.3458
0.66207 | 971685.7051 | 73.3329 | 177.0413
0.68456 | 1040156.169 | 74.1751 | 179.0746
0.70356 | 1070032.3965 | 84.1875 | 203.2466
0.68989 | 1026675.5547 | 61.3976 | 148.2269
0.68732 | 1053949.4579 | 69.304 | 167.3147
0.66225 | 980419.401 | 80.4174 | 194.1447
0.67207 | 1000970.5418 | 83.9572 | 202.6906
0.69117 | 1059793.5103 | 73.0492 | 176.3564
0.70036 | 1076612.1648 | 67.5177 | 163.0021
0.69817 | 1057137.5608 | 72.1355 | 174.1505
0.67665 | 1026753.5072 | 72.2907 | 174.5251
0.66915 | 992350.9619 | 72.3957 | 174.7786
0.68668 | 1040121.4776 | 68.4578 | 165.2718
0.67576 | 989375.8278 | 67.9843 | 164.1287
0.69449 | 1042096.9042 | 76.661 | 185.0761
0.69636 | 1044789.0269 | 74.9674 | 180.9873
0.71929 | 1101915.2473 | 78.2666 | 188.9524
0.69586 | 1061400.9938 | 70.8264 | 170.9902
0.70368 | 1066937.6416 | 70.4047 | 169.9719
0.68344 | 1027298.9735 | 69.7962 | 168.503
0.66385 | 984699.5097 | 85.096 | 205.4399
0.70615 | 1075085.1054 | 72.5312 | 175.1058
0.68185 | 1012340.3714 | 74.1847 | 179.0977
0.69913 | 1047347.5406 | 73.1313 | 176.5546
0.70462 | 1090421.4424 | 84.0609 | 202.941
0.68917 | 1032906.8533 | 69.5937 | 168.0141
0.69176 | 1062012.9322 | 73.1391 | 176.5734
0.69764 | 1064800.021 | 82.9078 | 200.1571
0.67371 | 1013863.2662 | 64.1242 | 154.8096
0.70574 | 1066877.3784 | 75.5347 | 182.3569
0.70324 | 1081604.7459 | 73.2724 | 176.8953
0.66042 | 974748.7072 | 88.9464 | 214.7355
0.69565 | 1066759.7856 | 80.0291 | 193.2072
0.69869 | 1054724.6785 | 53.3338 | 128.7591
0.69603 | 1050496.8925 | 81.9827 | 197.9239
0.69615 | 1068370.7367 | 78.4056 | 189.2878
0.7015 | 1054429.9623 | 75.1047 | 181.3188
0.68426 | 1017653.6597 | 67.2163 | 162.2744
0.70075 | 1074952.4119 | 65.3646 | 157.804
0.70124 | 1064843.2611 | 75.8741 | 183.1763
0.68996 | 1038129.0366 | 71.6812 | 173.0537
0.70242 | 1049422.3749 | 81.1536 | 195.9222
0.68764 | 1050537.2573 | 69.0472 | 166.6946
0.68205 | 1017597.9688 | 74.5484 | 179.9758
0.68147 | 1025584.0855 | 68.0175 | 164.2087
0.69803 | 1073511.7832 | 76.8436 | 185.5168
0.68835 | 1033850.9172 | 59.8645 | 144.5258
0.67352 | 981169.1775 | 69.4015 | 167.5501
0.70453 | 1083994.5811 | 67.579 | 163.1502
0.67586 | 1021119.8215 | 75.4918 | 182.2532
0.71947 | 1124694.2083 | 72.5868 | 175.2401
0.70616 | 1085940.5857 | 80.256 | 193.755
0.68827 | 1044779.639 | 64.6816 | 156.1552
0.6902 | 1051197.0396 | 74.3503 | 179.4976
0.70164 | 1069497.622 | 76.1624 | 183.8723
0.68239 | 1037858.5577 | 56.1237 | 135.4945
0.67765 | 1012099.4519 | 57.3647 | 138.4906
0.67787 | 1020895.8535 | 69.9318 | 168.8302
0.68739 | 1024735.0845 | 74.4875 | 179.8287
0.6811 | 1018506.4861 | 87.2546 | 210.6512
0.69691 | 1050733.7515 | 62.2034 | 150.1722
0.68028 | 1021537.7415 | 76.7176 | 185.2126
0.70639 | 1071402.4306 | 77.4885 | 187.0739
0.66151 | 978268.674 | 75.5536 | 182.4026
0.68684 | 1038841.8757 | 73.8357 | 178.2553
0.6968 | 1056379.5738 | 88.4606 | 213.5628
0.68675 | 1042036.9867 | 74.0727 | 178.8274
0.71272 | 1088252.7131 | 82.9406 | 200.2363
0.70462 | 1086342.1288 | 69.6799 | 168.2222
0.69374 | 1041529.4887 | 79.9681 | 193.0601
0.69463 | 1046491.2309 | 93.5846 | 225.9333
0.70645 | 1095124.8625 | 85.8098 | 207.1631
0.68824 | 1048859.1407 | 70.7649 | 170.8416
0.70883 | 1093623.9923 | 83.8078 | 202.33
0.70317 | 1098600.136 | 64.4587 | 155.617
0.70914 | 1107491.8676 | 79.7888 | 192.6271
0.69054 | 1040559.7181 | 69.9711 | 168.9252
0.68417 | 1017576.2555 | 75.9864 | 183.4474
0.6765 | 1015353.0187 | 71.5032 | 172.624
0.68888 | 1040551.3371 | 53.3662 | 128.8375
0.66156 | 976517.8565 | 66.7125 | 161.0583
0.69862 | 1070025.3557 | 58.8267 | 142.0202
0.69938 | 1057428.427 | 72.0306 | 173.8972
0.68543 | 1020026.1117 | 80.8363 | 195.1562
0.68721 | 1040423.7018 | 79.0118 | 190.7514
0.66186 | 977142.035 | 78.6377 | 189.8481
0.71439 | 1089544.7035 | 81.2359 | 196.1208
0.66777 | 980128.9044 | 68.1546 | 164.5398
0.71417 | 1092075.4211 | 74.6494 | 180.2195
0.67843 | 1013194.623 | 72.9056 | 176.0096
0.68085 | 1017845.5878 | 63.7007 | 153.7872
0.71047 | 1074929.1818 | 75.5985 | 182.5108
0.70116 | 1070263.7193 | 79.6618 | 192.3206
0.70366 | 1085182.9674 | 77.0577 | 186.0338
0.7086 | 1071574.5857 | 78.9215 | 190.5334
0.68761 | 1042246.9504 | 64.5873 | 155.9275
0.67361 | 985504.7123 | 76.3528 | 184.332
0.66964 | 992144.2574 | 65.0754 | 157.1059
0.69654 | 1064058.5251 | 82.9801 | 200.3317
0.67567 | 989898.6593 | 66.2147 | 159.8563
0.69733 | 1054112.0677 | 73.7938 | 178.1539
0.67703 | 1022305.4206 | 73.7136 | 177.9604
0.68954 | 1037024.6457 | 67.9713 | 164.0972
0.6842 | 1024277.6616 | 69.5135 | 167.8205
0.71921 | 1120492.3685 | 68.428 | 165.1999
0.67317 | 1009058.9916 | 74.8977 | 180.8189
0.70434 | 1083181.1416 | 85.1277 | 205.5165
0.66182 | 994201.7931 | 69.7543 | 168.4017
0.68036 | 1026177.2275 | 73.6301 | 177.7589
0.70697 | 1072693.1733 | 82.4197 | 198.9787
0.68171 | 1001352.7947 | 69.9786 | 168.9433
0.68359 | 1038540.3254 | 80.9217 | 195.3624
0.67792 | 1017453.1924 | 71.5017 | 172.6203
0.69159 | 1041018.8171 | 71.7114 | 173.1266
0.65699 | 970534.8661 | 62.9459 | 151.9649
0.67795 | 1011209.622 | 63.5566 | 153.4392
0.6949 | 1063777.2712 | 76.4903 | 184.664
0.69382 | 1051326.4339 | 80.7101 | 194.8514
0.68062 | 1019534.1568 | 63.8284 | 154.0955
0.67203 | 984305.8178 | 81.0769 | 195.737
0.68285 | 1043423.9941 | 70.9538 | 171.2977
0.67787 | 1008928.1464 | 70.9072 | 171.1851
0.67881 | 1012636.7581 | 82.5772 | 199.3589
0.66733 | 992264.302 | 82.3074 | 198.7075
0.70128 | 1071719.3779 | 74.6045 | 180.1113
0.68262 | 1026963.4376 | 67.9939 | 164.1517
0.68374 | 1030662.4508 | 66.208 | 159.8402
0.68593 | 1023964.7563 | 67.469 | 162.8845
0.69687 | 1055768.3684 | 79.066 | 190.8823
0.68912 | 1036357.8288 | 69.8322 | 168.5898
0.67715 | 1000098.7973 | 75.4552 | 182.165
0.70872 | 1108147.7092 | 68.6763 | 165.7993
0.69813 | 1072213.9878 | 74.9753 | 181.0064
0.68347 | 1017305.4853 | 69.1149 | 166.8581
0.6924 | 1044273.1935 | 73.626 | 177.7489
0.68693 | 1026008.2092 | 74.6308 | 180.1746
0.69232 | 1046430.1413 | 92.4998 | 223.3142
0.67565 | 1001479.2201 | 82.0487 | 198.0831
0.70507 | 1073600.5698 | 77.7932 | 187.8095
0.68959 | 1023123.2263 | 63.2203 | 152.6273
0.67317 | 999321.1042 | 83.6823 | 202.0269
0.68311 | 1025834.4788 | 83.1265 | 200.685
0.70221 | 1061130.1224 | 83.6477 | 201.9434
0.7039 | 1061211.2287 | 73.5982 | 177.6817
0.68535 | 1027808.628 | 65.869 | 159.0219
0.7133 | 1093770.3311 | 80.0255 | 193.1985
0.67713 | 1020989.4325 | 71.8336 | 173.4216
0.70172 | 1091995.9467 | 72.2592 | 174.4492
0.70817 | 1069972.2091 | 71.9803 | 173.7758
0.66845 | 978130.1026 | 73.1145 | 176.514
0.6706 | 986752.9673 | 68.3259 | 164.9533
0.69754 | 1047286.3441 | 76.7053 | 185.1829
0.66902 | 999368.6672 | 76.3158 | 184.2427
0.69925 | 1063503.6118 | 74.2476 | 179.2495
0.68186 | 1016249.5094 | 59.4046 | 143.4154
0.67124 | 993422.2742 | 83.4274 | 201.4116
0.70154 | 1054572.8806 | 71.6147 | 172.8931
0.70636 | 1088160.0369 | 73.3908 | 177.1812
0.68419 | 1034188.9271 | 72.011 | 173.8499
0.70232 | 1055960.9769 | 69.8151 | 168.5485
0.69097 | 1063131.0983 | 78.6069 | 189.7738
0.70346 | 1074545.5908 | 56.3674 | 136.0828
0.72583 | 1126862.1743 | 84.3324 | 203.5964
0.69461 | 1055772.1889 | 62.7086 | 151.392
0.6985 | 1047614.2073 | 98.4952 | 237.7883
0.69966 | 1063289.8872 | 75.9347 | 183.3225
0.70624 | 1073536.829 | 60.5874 | 146.271
0.67757 | 1004213.8116 | 74.0544 | 178.7832
0.68412 | 1024608.116 | 69.5446 | 167.8955
0.7081 | 1083383.4285 | 87.7732 | 211.9033
0.68087 | 1034588.7433 | 79.3007 | 191.4489
0.70086 | 1067481.1124 | 74.3536 | 179.5056
0.6737 | 1012124.2485 | 71.5719 | 172.7898
0.69931 | 1062879.1688 | 94.4135 | 227.9344
0.67298 | 1003376.0318 | 76.1718 | 183.8949
0.67225 | 1007680.2498 | 73.1752 | 176.6606
0.68166 | 1021496.3645 | 71.3918 | 172.3551
0.69988 | 1063437.8911 | 70.5551 | 170.3351
0.68887 | 1049648.5758 | 77.8462 | 187.9374
0.68677 | 1034650.3838 | 75.2766 | 181.7338
0.70033 | 1078873.424 | 83.6802 | 202.0218
0.70136 | 1066783.8902 | 74.2522 | 179.2608
0.66928 | 997675.7046 | 74.4831 | 179.8182
0.7048 | 1067668.7373 | 73.9508 | 178.533
0.69422 | 1046007.1001 | 78.6339 | 189.8391
0.71273 | 1090946.4616 | 72.9216 | 176.0483
0.72654 | 1143587.7863 | 85.4997 | 206.4146
0.67422 | 1014813.4936 | 69.8908 | 168.7314
0.68608 | 1026918.3447 | 76.4703 | 184.6156
0.67753 | 1003410.55 | 81.9797 | 197.9165
0.70895 | 1087401.5399 | 80.0475 | 193.2518
0.68629 | 1024011.9721 | 69.2844 | 167.2674
0.70573 | 1076534.9343 | 77.9329 | 188.1467
0.68442 | 1025654.155 | 65.8204 | 158.9045
0.65425 | 951742.3332 | 57.9824 | 139.9818
0.68335 | 1010806.7805 | 78.7885 | 190.2123
0.68288 | 1018845.486 | 79.4702 | 191.858
0.68516 | 1044959.1721 | 56.2793 | 135.8702
0.69867 | 1050070.6939 | 86.2225 | 208.1595
0.70553 | 1081888.8802 | 73.0539 | 176.3677
0.68787 | 1036383.542 | 83.3998 | 201.345
0.70213 | 1091695.3798 | 82.8176 | 199.9393
0.66924 | 990649.7637 | 81.9399 | 197.8204
0.70676 | 1099190.31 | 82.8403 | 199.9943
0.67377 | 1001560.0199 | 86.7371 | 209.4018
0.67456 | 998016.0215 | 73.9257 | 178.4724
0.70284 | 1074890.3643 | 69.1199 | 166.8703
0.715 | 1093126.6575 | 72.8355 | 175.8404
0.68153 | 1022075.937 | 78.2663 | 188.9515
0.7105 | 1088722.0184 | 65.2937 | 157.6329
0.7101 | 1086915.3565 | 86.2837 | 208.3072
0.69534 | 1064041.9717 | 67.2118 | 162.2636
0.69262 | 1045209.2804 | 81.9716 | 197.8969
0.70013 | 1049722.628 | 73.1931 | 176.7038
0.68977 | 1045022.4964 | 82.296 | 198.6801
0.68459 | 1034011.3747 | 86.6502 | 209.1922
0.67343 | 1020190.0624 | 72.172 | 174.2387
0.70502 | 1081202.6371 | 78.3494 | 189.1522
0.68581 | 1027464.3708 | 65.803 | 158.8625
0.69722 | 1063995.6625 | 81.2845 | 196.2382
0.68708 | 1023691.9402 | 79.3603 | 191.5927
0.71675 | 1103835.0951 | 69.6039 | 168.0386
0.69751 | 1065732.6582 | 56.3822 | 136.1187
0.70328 | 1072699.8656 | 72.6652 | 175.4294
0.71852 | 1097622.6925 | 67.0959 | 161.9838
0.70129 | 1066950.4074 | 81.0373 | 195.6413
0.70089 | 1069787.0941 | 68.9836 | 166.5411
0.71455 | 1101009.7571 | 88.3963 | 213.4075
0.69668 | 1043769.0576 | 87.4193 | 211.0488
0.69264 | 1026905.4089 | 82.0683 | 198.1304
0.69715 | 1066636.8604 | 76.4483 | 184.5625
0.71957 | 1123009.1743 | 70.5351 | 170.2868
0.70792 | 1080372.2964 | 81.2881 | 196.2469
0.67209 | 990744.4256 | 75.3125 | 181.8204
0.6684 | 984639.2682 | 87.3777 | 210.9485
0.68857 | 1045448.8983 | 70.5008 | 170.204
0.68176 | 1021854.4764 | 59.7413 | 144.2283
0.70065 | 1052665.4915 | 96.034 | 231.8465
0.70727 | 1097761.0911 | 84.3107 | 203.544
0.66606 | 982851.9001 | 64.8122 | 156.4705
0.71722 | 1102190.595 | 80.2008 | 193.6218
0.70291 | 1068034.3598 | 45.7801 | 110.523
0.70073 | 1073115.928 | 76.5195 | 184.7344
0.68191 | 1035709.2634 | 70.5264 | 170.2657
0.69464 | 1054552.7615 | 70.4933 | 170.1859
0.67666 | 1017262.483 | 67.9456 | 164.0352
0.70605 | 1085521.8878 | 69.4863 | 167.7547
0.68685 | 1035980.3544 | 71.8562 | 173.4762
0.68413 | 1042608.6746 | 68.8773 | 166.2845
0.70502 | 1069481.2485 | 72.2686 | 174.4718
0.70214 | 1068695.7133 | 76.0517 | 183.605
0.68693 | 1028970.1638 | 76.5277 | 184.7542
0.67584 | 1005703.9782 | 76.0249 | 183.5403
0.69484 | 1040316.2653 | 67.6669 | 163.3624
0.6649 | 978415.0242 | 56.5185 | 136.4478
0.66894 | 983103.7338 | 78.2798 | 188.984
0.67979 | 1010280.3557 | 80.7277 | 194.8938
0.66178 | 981284.0957 | 87.9511 | 212.3326
0.67998 | 1009237.4411 | 85.1473 | 205.5637
0.65815 | 980595.6574 | 74.4464 | 179.7295
0.69461 | 1054116.1365 | 84.004 | 202.8036
0.68237 | 1003549.4176 | 70.0643 | 169.1503
0.67113 | 1003514.4905 | 79.645 | 192.2802
0.7039 | 1054548.3719 | 85.9466 | 207.4935
0.68827 | 1025888.9893 | 79.3926 | 191.6706
0.71122 | 1098404.8839 | 73.5761 | 177.6285
0.67218 | 994693.8264 | 68.7136 | 165.8894
0.67157 | 1001308.5703 | 67.3861 | 162.6844
0.71079 | 1088025.3395 | 71.0776 | 171.5964
0.67638 | 1001868.394 | 78.897 | 190.4742
0.67953 | 1011048.9519 | 76.014 | 183.514
0.67856 | 1016467.3415 | 67.7019 | 163.4469
0.69127 | 1050835.7382 | 77.6633 | 187.4958
0.72377 | 1115742.607 | 71.5814 | 172.8129
0.69294 | 1059716.7085 | 76.4134 | 184.4783
0.71034 | 1092086.6735 | 71.3881 | 172.346
0.68464 | 1029908.2218 | 73.4364 | 177.2912
0.69287 | 1042202.2353 | 78.5008 | 189.5177
0.69046 | 1051503.173 | 79.4779 | 191.8766
0.70559 | 1059082.012 | 78.3337 | 189.1144
0.67912 | 1027904.8489 | 72.1077 | 174.0834
0.7007 | 1081363.6446 | 78.8442 | 190.3468
0.68573 | 1019384.2118 | 67.1866 | 162.2028
0.70177 | 1063299.3407 | 73.9658 | 178.5693
0.6879 | 1033975.4345 | 75.1042 | 181.3176
0.67215 | 1004446.9709 | 76.3468 | 184.3176
0.71579 | 1092026.9987 | 84.4506 | 203.8819
0.68265 | 1017775.3594 | 69.2298 | 167.1356
0.69066 | 1047868.6016 | 82.4142 | 198.9654
0.68126 | 1007678.928 | 68.6822 | 165.8134
0.72182 | 1120725.9544 | 75.7958 | 182.9873
0.7016 | 1070847.4906 | 81.2974 | 196.2693
0.69373 | 1040948.2462 | 78.3092 | 189.055
0.66116 | 969861.7606 | 79.6823 | 192.3702
0.69752 | 1060863.6639 | 81.0108 | 195.5773
0.68414 | 1020375.8166 | 67.4079 | 162.737
0.66217 | 958659.8149 | 88.5996 | 213.8984
0.71498 | 1088164.5279 | 85.7937 | 207.1243
0.68837 | 1041400.0375 | 72.0734 | 174.0006
0.69305 | 1053168.3493 | 60.5689 | 146.2263
0.6844 | 1023575.6332 | 69.446 | 167.6575
0.68435 | 1024991.8815 | 79.778 | 192.601
0.69689 | 1053454.7231 | 86.7208 | 209.3626
0.68311 | 1035290.8165 | 80.7257 | 194.8891
0.70817 | 1082085.1058 | 72.7954 | 175.7436
0.68259 | 1039114.1837 | 60.677 | 146.4872
0.686 | 1023299.4294 | 72.8198 | 175.8024
0.69996 | 1075753.8058 | 71.3633 | 172.2862
0.69762 | 1048199.9274 | 64.0237 | 154.5668
0.67488 | 1006217.9532 | 61.5126 | 148.5046
0.71049 | 1079414.6376 | 64.7363 | 156.2871
0.72252 | 1117300.2045 | 67.9205 | 163.9745
0.68283 | 1032153.1248 | 78.8866 | 190.449
0.68747 | 1022299.6467 | 69.0496 | 166.7004
0.6897 | 1053985.7265 | 72.8033 | 175.7627
0.69476 | 1060397.2368 | 76.2622 | 184.1132
0.67869 | 1019890.7046 | 86.2096 | 208.1284
0.68165 | 1027254.8233 | 78.6486 | 189.8744
0.70639 | 1097542.6063 | 79.3793 | 191.6387
0.71454 | 1106341.847 | 80.8045 | 195.0793
0.69218 | 1040433.6555 | 78.4456 | 189.3844
0.70193 | 1076462.4208 | 78.3491 | 189.1516
0.67893 | 1012413.9095 | 72.7961 | 175.7454
0.69019 | 1053800.5859 | 76.2893 | 184.1788
0.69578 | 1055081.9308 | 72.2422 | 174.4082
0.68568 | 1053832.9707 | 73.2202 | 176.7691
0.69921 | 1072318.6172 | 69.3124 | 167.3348
0.71668 | 1098938.967 | 71.2009 | 171.8943
0.69445 | 1050869.3364 | 72.0919 | 174.0452
0.69387 | 1049721.6455 | 75.5914 | 182.4939
0.69579 | 1068080.4792 | 74.9378 | 180.916
0.70174 | 1077890.8793 | 63.3727 | 152.9952
0.68297 | 1014535.2509 | 68.9767 | 166.5245
0.6949 | 1053244.6007 | 70.3794 | 169.911
0.66774 | 982446.2509 | 75.1681 | 181.4719
0.63895 | 922704.2779 | 67.0558 | 161.887
0.68251 | 1030610.4363 | 53.3499 | 128.798
0.71078 | 1105396.9641 | 64.1629 | 154.9029
0.70597 | 1089975.267 | 75.3706 | 181.9607
0.70034 | 1059196.1728 | 66.4275 | 160.3701
0.6657 | 979889.6474 | 71.8579 | 173.4804
0.70865 | 1097917.9029 | 77.4127 | 186.8909
0.70311 | 1087543.095 | 87.192 | 210.5
0.69355 | 1037658.495 | 76.4372 | 184.5357
0.67693 | 1013280.3071 | 57.4921 | 138.7982
0.72024 | 1117165.0017 | 72.0243 | 173.8821
0.68491 | 1030534.173 | 80.143 | 193.4824
0.69024 | 1048725.2505 | 68.6687 | 165.7809
0.68176 | 1020692.8424 | 84.3951 | 203.7479
0.69264 | 1049014.1044 | 73.9347 | 178.4942
0.70114 | 1059015.3524 | 62.316 | 150.4442
0.70591 | 1080137.9535 | 69.1553 | 166.9557
0.70228 | 1070361.9087 | 84.3668 | 203.6795
0.67198 | 994316.7142 | 90.5536 | 218.6158
0.68846 | 1034929.034 | 81.6564 | 197.1361
0.66722 | 999658.0584 | 69.5692 | 167.9549
0.69428 | 1058773.5514 | 79.5871 | 192.1403
0.66427 | 966931.084 | 63.934 | 154.3503
0.68099 | 1002312.6154 | 70.1198 | 169.2842
0.68972 | 1038939.1842 | 76.9564 | 185.7892
0.71085 | 1104743.5689 | 90.6816 | 218.9247
0.69028 | 1032832.2672 | 71.282 | 172.0899
0.68266 | 1031177.1578 | 73.74 | 178.024
0.72411 | 1095618.6168 | 66.4568 | 160.4409
0.69249 | 1039844.384 | 71.5371 | 172.7058
0.67231 | 993638.786 | 58.9082 | 142.2169
0.68596 | 1029090.1753 | 81.4917 | 196.7383
0.71159 | 1095932.401 | 56.8321 | 137.2048
0.71249 | 1084126.4073 | 77.7374 | 187.6748
0.70865 | 1076322.9618 | 68.6442 | 165.7218
0.6938 | 1046716.231 | 73.6496 | 177.8059
0.69019 | 1026475.0571 | 67.2164 | 162.2748
0.69463 | 1047939.5226 | 71.8722 | 173.515
0.68624 | 1033939.5547 | 73.5514 | 177.5688
0.69034 | 1032372.774 | 76.4019 | 184.4506
0.69084 | 1038163.0992 | 84.7111 | 204.5107
0.67133 | 995621.4642 | 68.8646 | 166.2538
0.69223 | 1034310.2351 | 74.6578 | 180.2399
0.69731 | 1061866.6429 | 84.2945 | 203.5049
0.6911 | 1042861.0562 | 68.1997 | 164.6487
0.71705 | 1101337.2489 | 83.4311 | 201.4205
0.69022 | 1053087.3732 | 87.5291 | 211.314
0.71972 | 1105814.4787 | 83.4595 | 201.4891
0.71583 | 1109793.4831 | 77.4657 | 187.0187
0.67239 | 1000530.5327 | 77.6575 | 187.4818
0.6846 | 1017756.9707 | 72.5744 | 175.21
0.68676 | 1025545.3227 | 74.6153 | 180.1373
0.68751 | 1039520.5759 | 88.2254 | 212.9948
0.69153 | 1052706.412 | 68.8009 | 166.1001
0.69127 | 1046771.5595 | 80.6578 | 194.7252
0.66249 | 977421.0179 | 72.0236 | 173.8805
0.70166 | 1065629.5149 | 66.7054 | 161.041
0.67479 | 1012809.7342 | 76.6812 | 185.1247
0.69455 | 1037721.9477 | 79.582 | 192.1281
0.70263 | 1074590.1857 | 76.4173 | 184.4878
0.67891 | 1022734.5272 | 80.2401 | 193.7166
0.71618 | 1116296.1681 | 72.3758 | 174.7307
0.67665 | 985791.7841 | 80.6446 | 194.6934
0.6877 | 1042448.7803 | 77.4439 | 186.966
0.71119 | 1089571.1543 | 81.6039 | 197.0093
0.68342 | 1019609.5649 | 67.4287 | 162.7873
0.69422 | 1053701.155 | 72.3098 | 174.5714
0.70935 | 1079253.1476 | 84.2203 | 203.3258
0.6806 | 1026453.1268 | 60.1122 | 145.1237
0.68043 | 1029906.8788 | 72.1098 | 174.0884
0.67879 | 1030398.7258 | 79.2212 | 191.2568
0.68812 | 1038674.343 | 78.775 | 190.1798
0.68589 | 1024743.3858 | 81.1561 | 195.9281
0.70324 | 1065491.3847 | 79.0715 | 190.8956
0.69171 | 1065343.2574 | 88.3572 | 213.3133
0.66565 | 977730.8064 | 71.7042 | 173.1092
0.70174 | 1049619.5119 | 74.7362 | 180.4292
0.70609 | 1073577.0933 | 82.5902 | 199.3904
0.69753 | 1054357.8581 | 77.2248 | 186.4372
0.68459 | 1034805.5109 | 62.363 | 150.5577
0.69591 | 1069864.9858 | 71.386 | 172.3411
0.68247 | 1030545.5729 | 73.7156 | 177.9653
0.69437 | 1057237.1351 | 79.4613 | 191.8366
0.67938 | 1028271.5949 | 73.8876 | 178.3804
0.67833 | 1019873.1261 | 78.4654 | 189.4322
0.69379 | 1062109.8126 | 81.5414 | 196.8584
0.70945 | 1084430.2754 | 69.0833 | 166.7819
0.68905 | 1027632.6255 | 74.5186 | 179.9038
0.7054 | 1070716.6936 | 74.0988 | 178.8902
0.68492 | 1026767.5174 | 67.753 | 163.5702
0.67126 | 1011800.1222 | 73.5825 | 177.6439
0.68741 | 1029622.7657 | 73.1203 | 176.5281
0.68712 | 1025273.9012 | 79.4783 | 191.8776
0.68376 | 1048862.9772 | 73.0791 | 176.4287
0.69174 | 1048237.6366 | 62.5201 | 150.9369
0.69316 | 1044950.0661 | 73.593 | 177.6691
0.6866 | 1020462.26 | 80.7331 | 194.9069
0.6959 | 1042692.1509 | 74.8662 | 180.743
0.66583 | 990063.4536 | 69.0459 | 166.6915
0.70586 | 1062719.6675 | 84.8692 | 204.8925
0.68305 | 1023206.0423 | 63.8302 | 154.0998
0.65536 | 968032.6062 | 68.5878 | 165.5856
0.68239 | 1017029.3352 | 71.0028 | 171.4159
0.69256 | 1050300.5099 | 79.1463 | 191.0761
0.68678 | 1026204.7726 | 79.5898 | 192.1469
0.70353 | 1094319.7214 | 61.9983 | 149.6772
0.68809 | 1040226.8554 | 80.5239 | 194.402
0.69794 | 1074353.6127 | 80.3345 | 193.9447
0.6752 | 1004719.8329 | 82.3691 | 198.8565
0.68602 | 1033726.843 | 83.4107 | 201.3712
0.68682 | 1038318.7089 | 78.8976 | 190.4755
0.685 | 1026796.691 | 65.8384 | 158.9481
0.70634 | 1081963.4532 | 71.1277 | 171.7175
0.67749 | 1016080.5848 | 84.5428 | 204.1044
0.66779 | 998873.6512 | 74.7384 | 180.4346
0.6795 | 1016537.2125 | 73.0313 | 176.3132
0.68657 | 1026937.7198 | 82.7973 | 199.8904
0.67503 | 1016025.8023 | 81.7597 | 197.3854
0.67382 | 1005887.5994 | 77.9071 | 188.0844
0.69412 | 1047871.9968 | 82.0034 | 197.9737
0.69404 | 1044069.3549 | 64.1007 | 154.7528
0.69376 | 1023734.7094 | 77.0107 | 185.9204
0.69712 | 1059530.5583 | 88.6836 | 214.1011
0.65566 | 956396.8823 | 58.1402 | 140.3629
0.69151 | 1036771.0524 | 65.8519 | 158.9805
0.66835 | 979923.6394 | 78.0962 | 188.5409
0.66392 | 969828.4297 | 71.3102 | 172.158
0.67687 | 1007762.83 | 61.4259 | 148.2953
0.69479 | 1062817.5071 | 69.1356 | 166.908
0.67194 | 994422.4771 | 68.7189 | 165.9021
0.69124 | 1040127.0132 | 69.8705 | 168.6824
0.68969 | 1029759.2097 | 81.104 | 195.8024
0.73397 | 1133763.1802 | 78.9177 | 190.5242
0.69586 | 1045353.662 | 78.3998 | 189.2739
0.69283 | 1043144.4357 | 55.5595 | 134.1325
0.70153 | 1063650.6071 | 72.232 | 174.3834
0.67572 | 1007616.2912 | 69.2414 | 167.1635
0.66795 | 987243.7653 | 80.2858 | 193.827
0.71999 | 1106517.3805 | 88.1663 | 212.8524
0.68746 | 1036925.1446 | 71.858 | 173.4807
0.70024 | 1058972.3946 | 88.2418 | 213.0345
0.6987 | 1061446.2287 | 56.0594 | 135.3394
0.69352 | 1039251.3817 | 69.5825 | 167.9871
0.68565 | 1032349.992 | 78.577 | 189.7018
0.69323 | 1042420.8799 | 69.8241 | 168.5702
0.67678 | 1020398.6689 | 84.9917 | 205.188
0.69782 | 1054509.4225 | 66.3731 | 160.2389
0.68653 | 1021227.3498 | 79.2326 | 191.2845
0.70968 | 1084244.0504 | 67.1039 | 162.0032
0.68789 | 1027343.4518 | 65.5959 | 158.3626
0.69567 | 1035880.9011 | 70.9392 | 171.2624
0.67721 | 1022572.6402 | 86.6256 | 209.1328
0.70131 | 1073823.6745 | 76.5072 | 184.7046
0.69236 | 1047908.4711 | 75.6255 | 182.5761
0.69404 | 1048191.2219 | 53.358 | 128.8175
0.7046 | 1067034.6371 | 60.3602 | 145.7225
0.67918 | 998457.8564 | 58.5954 | 141.4619
0.69019 | 1040821.9821 | 80.3247 | 193.921
0.68854 | 1043671.4099 | 68.5064 | 165.3892
0.6931 | 1055068.5187 | 70.684 | 170.6464
0.66297 | 972758.2789 | 73.3144 | 176.9966
0.71914 | 1113521.2407 | 77.3683 | 186.7837
0.69199 | 1049821.2677 | 72.2763 | 174.4904
0.71722 | 1113374.4359 | 52.292 | 126.2442
0.69339 | 1057255.9453 | 72.6405 | 175.3697
0.69457 | 1060880.9999 | 79.712 | 192.4417
0.71112 | 1094600.735 | 69.5676 | 167.9511
0.70485 | 1073611.2932 | 77.8179 | 187.8691
0.6553 | 947182.8402 | 53.3376 | 128.7683
0.69821 | 1050504.653 | 83.2851 | 201.0679
0.69447 | 1050742.1841 | 86.7993 | 209.5521
0.69561 | 1070474.0767 | 49.1787 | 118.728
0.6764 | 998374.2079 | 70.3186 | 169.764
0.67976 | 1010957.4536 | 70.8527 | 171.0536
0.7092 | 1082339.5699 | 82.8997 | 200.1376
0.70481 | 1087717.9102 | 75.3383 | 181.8828
0.69843 | 1074833.4479 | 76.8944 | 185.6395
0.68891 | 1034392.1212 | 65.9486 | 159.2139
0.69185 | 1062595.7674 | 63.3662 | 152.9796
0.69102 | 1038687.9614 | 76.7939 | 185.397
0.67632 | 1012346.1851 | 74.7922 | 180.5642
0.70367 | 1075496.4732 | 81.4101 | 196.5414
0.71558 | 1105485.9505 | 83.918 | 202.596
0.68412 | 1030525.1663 | 75.2662 | 181.7087
0.70596 | 1073951.2597 | 74.2634 | 179.2878
0.68282 | 1034874.1252 | 83.8949 | 202.5403
0.69002 | 1035699.623 | 79.9482 | 193.012
0.67597 | 1007252.1336 | 80.9408 | 195.4084
0.68503 | 1036117.6927 | 93.9554 | 226.8283
0.69372 | 1041324.5504 | 71.5121 | 172.6456
0.68879 | 1048604.3098 | 79.4511 | 191.8118
0.66711 | 973268.349 | 74.5364 | 179.9469
0.66825 | 983713.5508 | 64.6568 | 156.0953
0.68471 | 1043950.1084 | 76.1812 | 183.9176
0.69627 | 1054056.817 | 72.1215 | 174.1167
0.69868 | 1063164.0598 | 72.494 | 175.016
0.68732 | 1027751.0408 | 65.6766 | 158.5573
0.67182 | 988994.6252 | 76.8272 | 185.4773
0.68275 | 1023105.7501 | 76.5583 | 184.828
0.67479 | 995753.506 | 72.037 | 173.9127
0.67512 | 1026556.6894 | 64.7917 | 156.4211
0.71278 | 1089607.6051 | 74.6253 | 180.1613
0.6616 | 969597.2545 | 60.4637 | 145.9723
0.68909 | 1035522.2486 | 69.7783 | 168.4598
0.67141 | 980158.2109 | 67.7868 | 163.6518
0.67673 | 994336.9445 | 81.3336 | 196.3567
0.71514 | 1093711.0482 | 72.8603 | 175.9003
0.69354 | 1042988.637 | 75.1418 | 181.4084
0.70328 | 1065332.0475 | 82.2834 | 198.6496
0.70525 | 1070226.0854 | 72.2955 | 174.5368
0.68825 | 1030541.5046 | 82.8846 | 200.101
0.72811 | 1137781.734 | 66.5856 | 160.7519
0.68013 | 1013747.1228 | 69.2113 | 167.0909
0.70544 | 1087372.2504 | 63.5941 | 153.5297
0.71011 | 1088622.165 | 76.5996 | 184.9277
0.68703 | 1043849.8745 | 87.1637 | 210.4318
0.68106 | 1020908.8434 | 76.5827 | 184.8869
0.68006 | 1005423.2058 | 73.8715 | 178.3415
0.6764 | 979123.1791 | 65.4228 | 157.9447
0.67238 | 989518.802 | 82.9583 | 200.279
0.68044 | 1030762.3872 | 73.9509 | 178.5332
0.68092 | 1021820.5069 | 67.8222 | 163.7374
0.68184 | 1006805.0449 | 88.9851 | 214.8291
0.70779 | 1088641.3734 | 73.2749 | 176.9013
0.69339 | 1046466.6611 | 70.0784 | 169.1841
0.66952 | 1000716.2721 | 57.4075 | 138.5939
0.70146 | 1074211.6964 | 67.2615 | 162.3836
0.71656 | 1092273.1704 | 81.1932 | 196.0178
0.70248 | 1071904.2529 | 74.4894 | 179.8332
0.69765 | 1064607.7736 | 87.5442 | 211.3504
0.68684 | 1043709.9991 | 62.6936 | 151.3558
0.68505 | 1018822.2799 | 74.3107 | 179.402
0.67966 | 1007166.6307 | 77.3366 | 186.7072
0.68539 | 1024798.9383 | 74.7095 | 180.3647
0.6695 | 996804.2518 | 72.6825 | 175.4712
0.70059 | 1063843.0358 | 75.8738 | 183.1757
0.64888 | 929139.7624 | 73.156 | 176.6142
0.70179 | 1064077.4305 | 82.3618 | 198.8389
0.70916 | 1078893.2338 | 70.6438 | 170.5493
0.68993 | 1045358.7287 | 72.1837 | 174.2669
0.71382 | 1096244.9065 | 76.7358 | 185.2567
0.70441 | 1065786.7086 | 70.2557 | 169.6124
0.67808 | 1004186.6855 | 65.0452 | 157.0331
0.70662 | 1091730.3405 | 90.1507 | 217.6429
0.67184 | 998042.8022 | 65.2479 | 157.5224
0.70381 | 1073045.8042 | 68.9065 | 166.355
0.6991 | 1063374.9364 | 76.0021 | 183.4853
0.67989 | 1015520.2247 | 61.6243 | 148.7743
0.69124 | 1039149.0765 | 57.7399 | 139.3965
0.6795 | 1018245.2603 | 59.7281 | 144.1964
0.72111 | 1101858.9848 | 65.1092 | 157.1875
0.69166 | 1054501.8466 | 72.6327 | 175.3509
0.67156 | 1002173.8006 | 79.4878 | 191.9006
0.68775 | 1049371.7887 | 70.8498 | 171.0466
0.70088 | 1065029.7025 | 72.3925 | 174.771
0.71088 | 1090330.7504 | 76.6371 | 185.0183
0.69751 | 1074960.0512 | 71.5611 | 172.7637
0.67109 | 994609.9913 | 70.6068 | 170.46
0.68493 | 1022608.6169 | 71.1276 | 171.7171
0.72836 | 1130864.6702 | 72.0948 | 174.0522
0.67817 | 1015577.2701 | 71.3433 | 172.2379
0.69316 | 1056049.2811 | 53.8268 | 129.9493
0.69959 | 1053986.3325 | 83.3958 | 201.3352
0.68118 | 1013255.9966 | 88.2294 | 213.0047
0.68106 | 1000124.4012 | 79.7673 | 192.5754
0.69606 | 1062001.1784 | 51.396 | 124.0809
0.6835 | 1022641.5687 | 78.0838 | 188.511
0.68535 | 1030680.7679 | 75.1426 | 181.4102
0.71191 | 1092781.9471 | 63.3297 | 152.8915
0.68517 | 1033685.899 | 85.0444 | 205.3153
0.69221 | 1041204.7911 | 78.5007 | 189.5175
0.70648 | 1090091.6669 | 65.748 | 158.7298
0.6781 | 1020544.3229 | 87.8165 | 212.0079
0.70462 | 1079075.1056 | 65.2545 | 157.5383
0.69935 | 1064956.965 | 72.5461 | 175.1417
0.64793 | 945537.1589 | 62.1821 | 150.1208
0.7011 | 1066448.1332 | 64.815 | 156.4773
0.69447 | 1053396.422 | 88.5063 | 213.673
0.71043 | 1095907.6948 | 77.0779 | 186.0825
0.72488 | 1112558.3478 | 69.367 | 167.4666
0.66332 | 971745.498 | 76.4548 | 184.5782
0.67232 | 989644.813 | 79.9579 | 193.0354
0.72281 | 1113650.0162 | 86.851 | 209.677
0.65696 | 947135.9956 | 81.2693 | 196.2014
0.66444 | 980825.9577 | 81.7129 | 197.2723
0.69017 | 1034250.4895 | 74.8765 | 180.7679
0.70084 | 1073826.3859 | 71.1463 | 171.7624
0.67103 | 995661.0181 | 83.642 | 201.9297
0.68067 | 1013718.7375 | 63.4597 | 153.2054
0.65941 | 967578.1417 | 64.4485 | 155.5925
0.66452 | 969280.7804 | 69.489 | 167.7612
0.68737 | 1044643.332 | 68.6863 | 165.8235
0.68632 | 1036687.3914 | 75.4741 | 182.2107
0.7047 | 1074043.2384 | 84.3668 | 203.6794
0.70233 | 1071187.3359 | 74.4561 | 179.753
0.69387 | 1045589.1677 | 73.5622 | 177.5948
0.68392 | 1015789.4354 | 70.7576 | 170.824
0.7221 | 1116746.1277 | 79.0105 | 190.7483
0.65896 | 968894.5197 | 60.6623 | 146.4517
0.7004 | 1071530.7264 | 60.9785 | 147.215
0.69053 | 1046027.3395 | 77.3907 | 186.8377
0.67702 | 1010447.4597 | 72.8594 | 175.8982
0.68392 | 1039898.7839 | 72.4375 | 174.8796
0.68171 | 1022107.6214 | 69.6202 | 168.0781
0.69074 | 1037417.6561 | 59.8625 | 144.5209
0.70544 | 1086616.08 | 78.6949 | 189.9863
0.67098 | 997656.6574 | 70.8357 | 171.0125
0.68435 | 1018504.5432 | 70.9027 | 171.1742
0.70177 | 1058076.3852 | 83.4595 | 201.4891
0.69048 | 1051625.6267 | 79.4853 | 191.8945
0.69977 | 1064826.1124 | 66.8354 | 161.355
0.67667 | 1019051.392 | 76.9393 | 185.7478
0.71452 | 1098299.9735 | 65.497 | 158.1238
0.70263 | 1064199.0299 | 66.5236 | 160.6022
0.70064 | 1048805.7253 | 76.477 | 184.6317
0.6948 | 1053362.7385 | 64.6003 | 155.9589
0.70585 | 1076205.7452 | 68.9092 | 166.3616
0.68914 | 1039952.2924 | 83.5337 | 201.6683
0.71093 | 1101001.1184 | 80.8228 | 195.1234
0.68176 | 1013453.0194 | 83.7967 | 202.3031
0.66547 | 984348.174 | 80.3488 | 193.9791
0.68351 | 1031220.2114 | 70.0557 | 169.1294
0.70534 | 1076176.6463 | 76.2363 | 184.0508
0.68339 | 1016958.6152 | 69.0996 | 166.8212
0.69299 | 1054285.345 | 77.725 | 187.6448
0.68537 | 1029599.6712 | 69.7556 | 168.4049
0.68936 | 1037700.1885 | 48.1053 | 116.1364
0.6722 | 1016103.5596 | 62.3648 | 150.5619
0.68211 | 1033379.2799 | 79.0645 | 190.8785
0.6978 | 1053963.6311 | 62.3151 | 150.4421
0.70346 | 1065878.2102 | 75.708 | 182.7753
0.69136 | 1052933.2498 | 82.7279 | 199.7227
0.67728 | 1004572.9701 | 77.4082 | 186.88
0.69429 | 1051721.2293 | 72.1666 | 174.2257
0.68485 | 1033656.4924 | 66.2925 | 160.0442
0.69591 | 1053873.4073 | 85.3666 | 206.0931
0.67184 | 993320.8421 | 73.8754 | 178.3509
0.69806 | 1060408.9165 | 72.3697 | 174.716
0.69298 | 1033695.0843 | 84.0859 | 203.0014
0.69478 | 1057056.7855 | 66.8238 | 161.3269
0.66373 | 971253.4747 | 69.7135 | 168.3032
0.67865 | 1014356.579 | 61.666 | 148.8749
0.67543 | 999013.162 | 63.0244 | 152.1544
0.67067 | 1005887.1718 | 71.4521 | 172.5006
0.66598 | 994464.9977 | 76.2879 | 184.1753
0.68545 | 1035155.8022 | 70.1238 | 169.2939
0.70195 | 1049390.4132 | 76.3773 | 184.3912
0.68663 | 1031011.1722 | 79.9797 | 193.0881
0.67406 | 1009303.5018 | 81.3358 | 196.362
0.69807 | 1084355.4384 | 61.9139 | 149.4733
0.68793 | 1023617.6897 | 77.6549 | 187.4756
0.71219 | 1075393.9176 | 64.3411 | 155.3332
0.68062 | 1012647.1901 | 76.0126 | 183.5107
0.70342 | 1068468.159 | 70.5314 | 170.2778
0.69234 | 1052383.9093 | 77.6583 | 187.4837
0.71621 | 1096611.9976 | 67.2504 | 162.3569
0.68493 | 1022699.7161 | 76.2835 | 184.1648
0.68304 | 1018915.8346 | 63.5637 | 153.4563
0.66753 | 987170.9879 | 64.9538 | 156.8124
0.67141 | 999566.4338 | 74.3872 | 179.5865
0.6838 | 1018417.3549 | 69.1564 | 166.9583
0.67304 | 1000807.8604 | 75.9696 | 183.4069
0.70925 | 1068709.7001 | 76.9818 | 185.8505
0.68359 | 1039275.1841 | 75.9645 | 183.3944
0.70523 | 1070092.0672 | 79.1893 | 191.18
0.69108 | 1054293.2143 | 81.2595 | 196.1778
0.71778 | 1104862.0166 | 85.2062 | 205.7058
0.70233 | 1077603.3745 | 63.302 | 152.8246
0.69272 | 1044450.2639 | 75.337 | 181.8796
0.69006 | 1048663.3852 | 79.309 | 191.4687
0.67145 | 1004336.1883 | 65.697 | 158.6066
0.67793 | 1023855.1676 | 68.2572 | 164.7876
0.70178 | 1074460.8283 | 69.1971 | 167.0567
0.66264 | 957010.6889 | 72.1635 | 174.2181
0.70019 | 1056509.6797 | 70.064 | 169.1496
0.68173 | 1033910.8468 | 77.4822 | 187.0585
0.68288 | 1021151.2851 | 74.7173 | 180.3836
0.68297 | 1020689.8027 | 76.3525 | 184.3311
0.68531 | 1035120.1513 | 68.2651 | 164.8065
0.66448 | 981857.2623 | 77.1156 | 186.1735
0.68105 | 1037900.2281 | 77.8582 | 187.9664
0.71788 | 1110815.6979 | 65.9333 | 159.177
0.70086 | 1065970.463 | 66.8034 | 161.2776
0.67865 | 1021862.9451 | 68.0522 | 164.2926
0.69516 | 1060434.2041 | 85.9907 | 207.5999
0.70606 | 1068861.176 | 70.2935 | 169.7034
0.68937 | 1037300.2965 | 55.5258 | 134.0511
0.66347 | 994635.4158 | 63.879 | 154.2174
0.70346 | 1057144.6983 | 57.7191 | 139.3463
0.71152 | 1087055.8208 | 72.7848 | 175.7181
0.68759 | 1032443.4015 | 79.9795 | 193.0876
0.70243 | 1090062.9275 | 73.9306 | 178.4843
0.6757 | 1021786.0002 | 82.1416 | 198.3074
0.71168 | 1108742.1156 | 62.0963 | 149.9136
0.69293 | 1044556.3104 | 71.8154 | 173.3778
0.69137 | 1037702.4513 | 79.947 | 193.0092
0.68909 | 1047770.6067 | 77.682 | 187.541
0.71336 | 1115409.24 | 84.9339 | 205.0486
0.68459 | 1026416.0623 | 67.7621 | 163.5921
0.6758 | 1009004.7139 | 72.8867 | 175.9641
0.69867 | 1055260.3204 | 77.1642 | 186.2909
0.67315 | 985593.3137 | 64.3086 | 155.2547
0.69648 | 1043634.6241 | 79.8137 | 192.6873
0.68324 | 1023161.313 | 69.5363 | 167.8755
0.70992 | 1082674.8495 | 62.9197 | 151.9016
0.68833 | 1039503.8785 | 78.8142 | 190.2743
0.7194 | 1108448.374 | 78.7181 | 190.0422
0.68414 | 1037801.1442 | 83.9703 | 202.7223
0.67938 | 1010253.1006 | 82.3139 | 198.7234
0.6768 | 1005181.3541 | 67.1246 | 162.0532
0.73048 | 1159598.7134 | 81.102 | 195.7975
0.68286 | 1024913.6001 | 97.2618 | 234.8108
0.68878 | 1040814.5272 | 68.2947 | 164.8779
0.69045 | 1039340.18 | 73.4845 | 177.4074
0.69871 | 1063634.653 | 84.2654 | 203.4347
0.70543 | 1068633.7843 | 62.499 | 150.886
0.68326 | 1033410.2603 | 85.2588 | 205.8329
0.70273 | 1066938.1202 | 74.1178 | 178.9363
0.71767 | 1097846.7858 | 84.6776 | 204.4298
0.70504 | 1088625.8433 | 77.1512 | 186.2595
0.7131 | 1104130.8645 | 79.1045 | 190.9752
0.69176 | 1051207.7907 | 79.3066 | 191.463
0.67373 | 1000246.4685 | 59.1752 | 142.8615
0.66986 | 1005217.8992 | 74.9816 | 181.0215
0.67303 | 992642.8086 | 83.1402 | 200.7182
0.71818 | 1111229.8626 | 88.7662 | 214.3004
0.69274 | 1029176.5532 | 86.6879 | 209.2832
0.69613 | 1031679.7325 | 75.715 | 182.7923
0.66719 | 988784.7201 | 74.0504 | 178.7734
0.70091 | 1069678.5482 | 74.5019 | 179.8635
0.68491 | 1032834.9454 | 75.3479 | 181.9059
0.68614 | 1040309.548 | 75.758 | 182.896
0.70806 | 1085305.9635 | 58.872 | 142.1295
0.66619 | 978819.3852 | 60.6188 | 146.3466
0.68916 | 1029409.6288 | 74.6349 | 180.1845
0.6595 | 968216.809 | 68.1702 | 164.5773
0.70442 | 1088215.686 | 69.8194 | 168.5589
0.70294 | 1080372.7298 | 67.5049 | 162.9712
0.66997 | 1009909.5821 | 80.2914 | 193.8405
0.68649 | 1032652.6264 | 87.7722 | 211.9009
0.71998 | 1112475.677 | 77.205 | 186.3893
0.70031 | 1065152.5651 | 87.2561 | 210.6549
0.69482 | 1060726.8879 | 84.505 | 204.0131
0.69626 | 1061328.7184 | 66.9517 | 161.6357
0.69311 | 1042790.7365 | 64.1076 | 154.7694
0.69912 | 1066207.6597 | 66.9839 | 161.7135
0.67286 | 991686.7174 | 60.8067 | 146.8005
0.68773 | 1039900.1894 | 75.5529 | 182.4009
0.69498 | 1036549.7118 | 67.4086 | 162.7388
0.68728 | 1019938.5696 | 82.1472 | 198.321
0.71469 | 1090397.8586 | 76.2329 | 184.0424
0.69473 | 1043543.5689 | 69.7486 | 168.3879
0.70899 | 1092801.107 | 91.4233 | 220.7154
0.688 | 1043207.5911 | 86.9873 | 210.006
0.70019 | 1066004.5699 | 79.4539 | 191.8188
0.68229 | 1009453.7572 | 78.8056 | 190.2536
0.69623 | 1066325.7385 | 90.8434 | 219.3153
0.68989 | 1039264.2071 | 75.2696 | 181.7169
0.68686 | 1044805.1582 | 82.6272 | 199.4798
0.68995 | 1041290.0138 | 60.1547 | 145.2262
0.69453 | 1036808.7455 | 72.8758 | 175.9377
0.68504 | 1020840.4896 | 72.7467 | 175.626
0.71264 | 1080813.314 | 58.2508 | 140.6299
0.71817 | 1089517.5845 | 62.5903 | 151.1063
0.70546 | 1075676.0559 | 56.2667 | 135.8398
0.67491 | 1004005.1367 | 67.2332 | 162.3152
0.68265 | 1034734.9567 | 60.4151 | 145.8551
0.72958 | 1146124.8829 | 74.8403 | 180.6804
0.68307 | 1031449.244 | 80.4307 | 194.1768
0.71697 | 1103564.0963 | 76.9083 | 185.6732
0.70656 | 1070040.9024 | 77.3714 | 186.7911
0.6597 | 961152.6487 | 55.9587 | 135.0963
0.67498 | 1010134.052 | 81.292 | 196.2563
0.71168 | 1090557.4943 | 65.2954 | 157.6371
0.6878 | 1039230.1815 | 70.1598 | 169.3808
0.68028 | 1033249.1311 | 74.1468 | 179.0063
0.67534 | 1010585.2651 | 84.3535 | 203.6473
0.67264 | 1010339.0287 | 94.6672 | 228.5468
0.703 | 1058259.9584 | 80.1817 | 193.5757
0.70218 | 1058210.2725 | 78.6496 | 189.8769
0.70333 | 1075560.3575 | 74.6535 | 180.2295
0.69568 | 1061558.6995 | 77.9582 | 188.2078
0.6845 | 1030682.843 | 64.1604 | 154.8968
0.71017 | 1083214.3361 | 68.9361 | 166.4265
0.71529 | 1110892.0016 | 92.1183 | 222.3933
0.66293 | 975635.3299 | 87.2927 | 210.7432
0.6875 | 1029667.0371 | 69.5443 | 167.8948
0.6892 | 1051210.3727 | 68.9353 | 166.4245
0.68669 | 1038785.2159 | 68.369 | 165.0573
0.66882 | 998486.5087 | 72.8802 | 175.9484
0.689 | 1039469.9976 | 66.2638 | 159.9751
0.69375 | 1055223.5888 | 77.7086 | 187.6051
0.69455 | 1055797.6483 | 62.7156 | 151.4088
0.68004 | 1000494.2652 | 60.699 | 146.5405
0.71632 | 1107270.8284 | 69.8627 | 168.6635
0.67967 | 1017906.1567 | 78.1594 | 188.6934
0.70429 | 1085960.8102 | 71.598 | 172.8529
0.66595 | 973591.0852 | 69.649 | 168.1475
0.69987 | 1066076.4342 | 77.8213 | 187.8772
0.68813 | 1047821.3459 | 81.1064 | 195.8082
0.68661 | 1029721.4136 | 69.6382 | 168.1216
0.70238 | 1057208.9774 | 78.553 | 189.6436
0.66917 | 995030.4119 | 67.2041 | 162.245
0.69071 | 1043779.7594 | 77.2052 | 186.3898
0.69623 | 1039603.0242 | 71.1365 | 171.7387
0.71048 | 1090718.0189 | 86.5398 | 208.9256
0.69288 | 1048067.9732 | 89.7206 | 216.6047
0.66721 | 981914.461 | 76.8038 | 185.4207
0.70033 | 1067501.7346 | 84.376 | 203.7016
0.70909 | 1090784.5263 | 65.5475 | 158.2456
0.6993 | 1066589.8307 | 77.9222 | 188.1209
0.69267 | 1046294.5978 | 66.3456 | 160.1726
0.67902 | 1020473.5582 | 67.9227 | 163.9799
0.66006 | 973227.7534 | 71.979 | 173.7726
0.69496 | 1047314.8428 | 78.6485 | 189.8743
0.68086 | 1031950.8646 | 79.0995 | 190.9631
0.67985 | 1009039.0205 | 80.6213 | 194.6371
0.68284 | 1025173.4584 | 75.0166 | 181.106
0.67442 | 1005007.1798 | 84.0024 | 202.7997
0.68937 | 1028158.1839 | 82.5585 | 199.3139
0.68262 | 1009543.6588 | 80.335 | 193.9459
0.69004 | 1035804.1541 | 67.8422 | 163.7856
0.69666 | 1063400.7529 | 74.5155 | 179.8964
0.69653 | 1034367.2144 | 77.5008 | 187.1034
0.69361 | 1046281.6674 | 89.4731 | 216.0073
0.67581 | 994227.0002 | 51.9767 | 125.4828
0.66118 | 977427.1869 | 68.6443 | 165.722
0.69482 | 1049028.4563 | 70.1054 | 169.2493
0.69781 | 1063735.1247 | 68.8954 | 166.3283
0.70117 | 1067957.85 | 81.1275 | 195.8591
0.70305 | 1080935.5256 | 61.2816 | 147.9469
0.68869 | 1033234.7642 | 69.6646 | 168.1852
0.69509 | 1047046.9191 | 76.9588 | 185.795
0.68889 | 1026270.4266 | 70.9725 | 171.3428
0.67678 | 1009676.6524 | 63.5695 | 153.4702
0.69438 | 1057667.483 | 76.1076 | 183.7401
0.68725 | 1024198.8704 | 72.7513 | 175.6372
0.67857 | 989516.5711 | 86.4936 | 208.814
0.67514 | 1002501.1208 | 77.4932 | 187.0852
0.6721 | 1009799.5843 | 90.543 | 218.5901
0.69682 | 1054412.7715 | 66.462 | 160.4533
0.68853 | 1026529.3181 | 77.8076 | 187.8441
0.69557 | 1050928.7625 | 73.7573 | 178.0659
0.70786 | 1081427.8584 | 78.051 | 188.4319
0.67743 | 1016007.0408 | 71.4513 | 172.4986
0.69205 | 1046709.6394 | 66.4311 | 160.3788
0.69899 | 1046507.8274 | 59.6742 | 144.0663
0.70136 | 1079047.4757 | 83.5718 | 201.7603
0.67558 | 997178.8169 | 78.6209 | 189.8076
0.67759 | 998965.4716 | 62.8204 | 151.662
0.69734 | 1058275.1151 | 77.5322 | 187.1793
0.69553 | 1048155.8624 | 78.6477 | 189.8723
0.68581 | 1037877.0632 | 73.8781 | 178.3574
0.70999 | 1092795.3963 | 71.4701 | 172.5442
0.67753 | 1006096.8828 | 80.6614 | 194.734
0.69984 | 1057437.7739 | 77.1942 | 186.3633
0.65836 | 976208.1788 | 77.3848 | 186.8235
0.69056 | 1053739.3571 | 77.9587 | 188.2089
0.69757 | 1058498.6849 | 79.1905 | 191.1828
0.71001 | 1073232.6501 | 64.4089 | 155.4968
0.6991 | 1059726.5188 | 85.4485 | 206.291
0.682 | 1027158.3316 | 68.0694 | 164.3341
0.69425 | 1042102.1239 | 79.9525 | 193.0224
0.69449 | 1067757.4271 | 85.1059 | 205.4638
0.67709 | 1022103.8215 | 77.1225 | 186.1903
0.66715 | 978406.1554 | 73.2781 | 176.909
0.68088 | 1009795.8023 | 61.188 | 147.7209
0.67319 | 1006837.9757 | 71.252 | 172.0175
0.71874 | 1118091.0753 | 77.7289 | 187.6541
0.70279 | 1069844.4464 | 62.8056 | 151.6261
0.69645 | 1066036.3722 | 65.2229 | 157.4621
0.703 | 1061195.0115 | 92.3793 | 223.0233
0.70173 | 1068411.774 | 70.6979 | 170.6797
0.68072 | 1029084.5423 | 70.197 | 169.4706
0.68099 | 1023978.3869 | 67.9587 | 164.0668
0.67725 | 1016844.1166 | 76.7067 | 185.1865
0.66862 | 1010976.6589 | 73.8135 | 178.2016
0.67786 | 1013395.7749 | 79.6304 | 192.2447
0.6954 | 1057537.7578 | 74.2396 | 179.2303
0.67716 | 997754.4722 | 54.3719 | 131.2653
0.67356 | 996175.1501 | 72.1998 | 174.3057
0.69152 | 1049817.4035 | 60.5095 | 146.083
0.69584 | 1060113.2286 | 81.1484 | 195.9096
0.70352 | 1083754.2752 | 65.6404 | 158.4701
0.68702 | 1033488.0864 | 66.0718 | 159.5114
0.68598 | 1030713.9937 | 71.5825 | 172.8154
0.68418 | 1024476.9325 | 65.7678 | 158.7775
0.69882 | 1046021.5401 | 73.0712 | 176.4094
0.67412 | 990833.6071 | 70.0229 | 169.0502
0.6682 | 984735.0094 | 84.5443 | 204.108
0.6701 | 1001144.6659 | 61.6391 | 148.81
0.67631 | 1007319.1874 | 74.4216 | 179.6697
0.70716 | 1082497.4679 | 79.8883 | 192.8674
0.69444 | 1051129.6438 | 64.4967 | 155.7089
0.70793 | 1083424.566 | 76.4686 | 184.6116
0.7327 | 1146937.4808 | 85.0303 | 205.2812
0.69018 | 1033958.8134 | 79.7957 | 192.644
0.69307 | 1050243.9301 | 85.421 | 206.2246
0.67285 | 1008091.6012 | 78.1905 | 188.7685
0.68558 | 1029627.4977 | 71.5488 | 172.7341
0.67981 | 1032517.3908 | 85.6174 | 206.6986
0.68205 | 1023585.8049 | 65.9156 | 159.1344
0.67479 | 1003581.3346 | 74.3123 | 179.4058
0.70046 | 1071824.6571 | 79.8571 | 192.7921
0.68798 | 1033835.4294 | 81.6267 | 197.0642
0.68724 | 1016959.1413 | 81.9127 | 197.7547
0.69983 | 1057144.538 | 62.3218 | 150.4582
0.70109 | 1067395.1906 | 79.4554 | 191.8222
0.6978 | 1077074.0471 | 67.7947 | 163.6709
0.68568 | 1022658.1278 | 77.8368 | 187.9147
0.71619 | 1101467.4918 | 78.5735 | 189.6933
0.69616 | 1050297.701 | 66.2652 | 159.9783
0.65955 | 951215.2791 | 76.8193 | 185.4581
0.6894 | 1038217.1824 | 71.7457 | 173.2095
0.69872 | 1044116.5932 | 72.9489 | 176.1143
0.68984 | 1042289.1071 | 88.1614 | 212.8404
0.67857 | 1016997.3153 | 76.5695 | 184.8552
0.67856 | 993581.3277 | 73.1633 | 176.6318
0.68602 | 1055032.2421 | 56.7509 | 137.0088
0.6791 | 1000330.4778 | 64.6986 | 156.1962
0.69145 | 1034596.7089 | 57.4688 | 138.7421
0.69017 | 1060718.3301 | 71.3647 | 172.2895
0.68068 | 1003590.8806 | 68.4624 | 165.2829
0.70052 | 1067453.8439 | 74.3883 | 179.5893
0.69257 | 1050425.7659 | 77.4881 | 187.0728
0.70169 | 1066823.2944 | 66.7773 | 161.2148
0.70186 | 1067237.3057 | 75.9778 | 183.4266
0.67916 | 1016731.2705 | 82.8678 | 200.0605
0.6964 | 1068115.4913 | 67.3663 | 162.6366
0.67505 | 1003769.4971 | 75.9677 | 183.4022
0.68841 | 1053936.3297 | 61.0218 | 147.3196
0.67179 | 981954.4803 | 71.3566 | 172.2701
0.66856 | 1003599.7967 | 78.3415 | 189.133
0.68193 | 1016447.0622 | 57.2494 | 138.2122
0.68978 | 1056912.2236 | 71.1487 | 171.7681
0.67667 | 1012864.7909 | 81.308 | 196.2949
0.66062 | 959464.3865 | 63.7911 | 154.0054
0.69048 | 1043996.6465 | 78.908 | 190.5007
0.69791 | 1043051.8308 | 74.6765 | 180.285
0.67793 | 1003947.6416 | 65.044 | 157.0302
0.7016 | 1066654.3763 | 79.3433 | 191.5516
0.69331 | 1048105.9735 | 73.799 | 178.1664
0.67321 | 1010039.836 | 63.9301 | 154.3408
0.69507 | 1047909.4743 | 86.4912 | 208.8083
0.69406 | 1042551.0358 | 73.3288 | 177.0315
0.6928 | 1038915.6523 | 78.261 | 188.9387
0.66032 | 968809.2389 | 64.4416 | 155.5758
0.6953 | 1061151.3004 | 65.0311 | 156.9988
0.70922 | 1082456.2032 | 79.6532 | 192.2997
0.68348 | 1025949.2246 | 76.6891 | 185.1438
0.68278 | 997591.1074 | 66.0336 | 159.4192
0.67991 | 1010752.6776 | 77.5828 | 187.3015
0.71125 | 1077537.597 | 73.9576 | 178.5494
0.69306 | 1064088.347 | 71.2928 | 172.1161
0.67186 | 1011007.7011 | 80.1517 | 193.5034
0.69104 | 1044594.5031 | 75.5305 | 182.3468
0.71675 | 1092846.6213 | 74.4331 | 179.6973
0.69568 | 1041956.0114 | 78.1656 | 188.7085
0.7017 | 1070872.5604 | 74.5005 | 179.8602
0.67217 | 996778.9084 | 83.593 | 201.8114
0.71428 | 1084185.4083 | 80.8489 | 195.1865
0.68983 | 1034550.4993 | 69.6377 | 168.1203
0.6887 | 1035757.6621 | 71.6486 | 172.975
0.68737 | 1045870.1567 | 87.2899 | 210.7365
0.67417 | 1005020.2774 | 82.5197 | 199.2201
0.70394 | 1070682.2295 | 66.8048 | 161.2811
0.67558 | 1017202.1695 | 68.0493 | 164.2856
0.70339 | 1089930.6252 | 71.9518 | 173.707
0.69882 | 1069722.7911 | 83.15 | 200.7419
0.65593 | 943128.252 | 81.0519 | 195.6766
0.6836 | 1022019.492 | 79.634 | 192.2535
0.68525 | 1023249.4614 | 67.4884 | 162.9314
0.72749 | 1128715.7275 | 85.884 | 207.3423
0.68117 | 1031662.4386 | 58.171 | 140.4371
0.68002 | 1029944.8859 | 72.1554 | 174.1985
0.6983 | 1056252.3955 | 70.2364 | 169.5656
0.68602 | 1027453.6589 | 73.1142 | 176.5134
0.67811 | 998349.9151 | 73.5633 | 177.5976
0.67448 | 1013574.0837 | 66.0038 | 159.3472
0.68213 | 1029934.311 | 66.317 | 160.1035
0.70124 | 1061132.9994 | 80.6015 | 194.5892
0.68686 | 1026265.9869 | 61.6445 | 148.8231
0.68472 | 1030630.8248 | 73.4789 | 177.3938
0.6899 | 1043413.1187 | 66.4286 | 160.3729
0.68121 | 1018111.3426 | 69.3984 | 167.5426
0.69411 | 1059280.7463 | 68.9725 | 166.5143
0.68468 | 1031952.759 | 77.0018 | 185.8988
0.68124 | 1020518.0132 | 71.3666 | 172.2942
0.69631 | 1050856.1547 | 68.2078 | 164.6682
0.68963 | 1044661.6029 | 80.686 | 194.7933
0.68201 | 1022367.1631 | 61.8472 | 149.3124
0.68793 | 1025306.1865 | 77.7832 | 187.7852
0.69437 | 1058443.9911 | 68.8929 | 166.3221
0.70108 | 1070273.8659 | 86.0553 | 207.7559
0.68934 | 1038239.4364 | 50.7352 | 122.4856
0.6797 | 1017417.009 | 73.0726 | 176.4128
0.7052 | 1084451.4249 | 76.8382 | 185.5038
0.70074 | 1054699.2759 | 75.2053 | 181.5617
0.68274 | 1011742.8445 | 78.9324 | 190.5597
0.69389 | 1053571.4672 | 64.6979 | 156.1946
0.69253 | 1041306.9795 | 71.6342 | 172.9402
0.70594 | 1073687.0726 | 82.6161 | 199.4529
0.70532 | 1078385.2501 | 84.633 | 204.3221
0.70132 | 1056018.7195 | 73.5065 | 177.4604
0.69561 | 1048847.1988 | 75.3391 | 181.8848
0.685 | 1031907.4447 | 78.6181 | 189.8008
0.66949 | 996001.4964 | 77.6409 | 187.4417
0.71137 | 1073387.3216 | 83.5246 | 201.6463
0.71274 | 1093502.4857 | 83.3381 | 201.1961
0.71436 | 1096483.6961 | 73.1242 | 176.5374
0.69088 | 1037771.9696 | 76.742 | 185.2715
0.68756 | 1029090.6533 | 73.066 | 176.397
0.69563 | 1071805.6668 | 65.6534 | 158.5014
0.68989 | 1041773.3606 | 80.4324 | 194.1809
0.67564 | 1013382.1496 | 74.9435 | 180.9296
0.69193 | 1055918.0799 | 60.8255 | 146.8458
0.69318 | 1058785.395 | 78.0361 | 188.3958
0.70808 | 1083593.1814 | 69.164 | 166.9767
0.68564 | 1033484.5865 | 60.0797 | 145.0451
0.70635 | 1065959.1918 | 75.4047 | 182.043
0.68392 | 1023832.5139 | 78.1262 | 188.6134
0.69057 | 1058369.6105 | 74.8179 | 180.6263
0.69531 | 1049305.2832 | 50.836 | 122.729
0.71851 | 1102332.9927 | 76.6579 | 185.0685
0.67962 | 1008344.0728 | 52.4189 | 126.5504
0.68461 | 1015355.4421 | 83.6041 | 201.8383
0.70455 | 1068847.7628 | 68.5644 | 165.5291
0.68371 | 1023253.0131 | 81.2962 | 196.2665
0.70339 | 1081814.165 | 54.3379 | 131.1833
0.68193 | 1024033.163 | 61.9547 | 149.5718
0.68802 | 1027902.1844 | 71.1753 | 171.8324
0.67266 | 996969.7192 | 77.3495 | 186.7381
0.68495 | 1031963.392 | 84.93 | 205.039
0.69767 | 1078917.6272 | 75.4935 | 182.2575
0.69003 | 1041363.8963 | 70.0765 | 169.1795
0.69229 | 1057837.9089 | 60.9046 | 147.0366
0.67286 | 999833.2137 | 74.2471 | 179.2483
0.711 | 1092461.65 | 71.3465 | 172.2457
0.69192 | 1041455.8521 | 83.3177 | 201.1466
0.70808 | 1071227.7616 | 73.854 | 178.2994
0.66748 | 991911.6974 | 77.5542 | 187.2323
0.69109 | 1037475.5425 | 81.5459 | 196.8693
0.68417 | 1028321.8109 | 62.011 | 149.7077
0.69427 | 1055467.7585 | 72.4258 | 174.8512
0.68072 | 1024279.8194 | 83.292 | 201.0847
0.70103 | 1073005.3525 | 72.9189 | 176.0419
0.67268 | 1000662.4052 | 74.3284 | 179.4446
0.66006 | 965445.3132 | 79.4509 | 191.8114
0.71663 | 1099911.644 | 69.3679 | 167.469
0.69757 | 1065469.4916 | 84.5484 | 204.118
0.67785 | 1011968.4103 | 71.1232 | 171.7066
0.68035 | 1034984.2813 | 76.0107 | 183.5062
0.70166 | 1069628.1186 | 70.7854 | 170.891
0.69942 | 1053250.9679 | 73.5132 | 177.4765
0.71484 | 1117184.9232 | 62.2922 | 150.3866
0.70122 | 1066761.0858 | 60.6363 | 146.3891
0.68707 | 1022392.7794 | 70.8912 | 171.1466
0.70156 | 1069881.6349 | 70.8121 | 170.9555
0.69951 | 1052487.3662 | 75.1792 | 181.4986
0.6826 | 1018145.6776 | 57.2883 | 138.3061
0.68768 | 1035702.1972 | 66.1503 | 159.7009
0.67536 | 1007825.1364 | 65.5435 | 158.236
0.69859 | 1074003.8093 | 73.2418 | 176.8214
0.6987 | 1054919.1845 | 68.2381 | 164.7413
0.71005 | 1068394.6164 | 64.1153 | 154.788
0.69717 | 1062870.8825 | 76.1757 | 183.9045
0.70413 | 1093650.992 | 81.0591 | 195.694
0.72407 | 1119607.2217 | 78.9984 | 190.7189
0.70819 | 1096160.1544 | 69.955 | 168.8862
0.68143 | 1017088.6735 | 81.0767 | 195.7366
0.71336 | 1088087.3197 | 76.2134 | 183.9954
0.68815 | 1029915.973 | 74.7815 | 180.5386
0.69785 | 1059991.2538 | 66.7148 | 161.0639
0.7005 | 1050359.2938 | 67.7684 | 163.6075
0.71236 | 1084479.4974 | 74.318 | 179.4195
0.69973 | 1066334.9661 | 68.8749 | 166.2786
0.68575 | 1041642.3988 | 77.6041 | 187.3529
0.7088 | 1091560.7225 | 71.5093 | 172.6388
0.69522 | 1073397.1102 | 65.7896 | 158.8301
0.69281 | 1042609.2162 | 64.1249 | 154.8111
0.68538 | 1031275.7996 | 64.1155 | 154.7884
0.6934 | 1055242.6776 | 72.5016 | 175.0343
0.66648 | 991031.8208 | 50.09 | 120.928
0.68685 | 1050110.2321 | 74.6356 | 180.1862
0.6861 | 1020853.5271 | 67.2469 | 162.3485
0.68959 | 1037861.2664 | 73.4724 | 177.378
0.69646 | 1060262.1502 | 85.8164 | 207.1791
0.71184 | 1112550.7566 | 84.897 | 204.9595
0.68637 | 1014557.6005 | 73.8577 | 178.3082
0.68019 | 1020653.3481 | 71.4869 | 172.5847
0.6649 | 994495.0652 | 76.4143 | 184.4804
0.66456 | 975516.9041 | 82.7471 | 199.7691
0.68357 | 1020380.4877 | 72.871 | 175.9262
0.68218 | 1001344.8718 | 60.5599 | 146.2046
0.70904 | 1084772.2762 | 79.8436 | 192.7595
0.69013 | 1042857.3528 | 53.3877 | 128.8892
0.67395 | 995102.9076 | 77.0571 | 186.0324
0.70915 | 1085506.5233 | 63.1533 | 152.4655
0.6967 | 1063191.7089 | 62.9607 | 152.0006
0.66488 | 991798.4252 | 77.849 | 187.9442
0.67626 | 1001507.7707 | 75.661 | 182.6618
0.69082 | 1047093.2889 | 75.7356 | 182.8418
0.70253 | 1065876.9218 | 80.8135 | 195.1011
0.68834 | 1023886.9517 | 80.7989 | 195.0658
0.67575 | 1001813.1856 | 70.6278 | 170.5106
0.68166 | 1022107.114 | 93.1425 | 224.866
0.67829 | 992182.0838 | 67.9225 | 163.9795
0.66899 | 1000180.0607 | 65.6663 | 158.5325
0.68438 | 1031854.1611 | 81.8993 | 197.7224
0.68397 | 1017427.4639 | 70.8674 | 171.0889
0.6865 | 1040703.3983 | 82.346 | 198.8009
0.7138 | 1106067.5773 | 61.7618 | 149.1062
0.68678 | 1038883.6132 | 62.4097 | 150.6703
0.70989 | 1083823.0593 | 58.6401 | 141.5696
0.71359 | 1086595.1154 | 76.0839 | 183.6828
0.67986 | 1010205.839 | 79.1459 | 191.0751
0.69597 | 1066609.1583 | 73.0118 | 176.2661
0.69185 | 1054903.9734 | 62.9079 | 151.873
0.70555 | 1076788.7447 | 70.8661 | 171.086
0.70095 | 1073944.9189 | 71.1431 | 171.7545
0.69059 | 1049472.3588 | 77.6396 | 187.4385
0.69307 | 1030079.4345 | 75.9314 | 183.3147
0.69779 | 1045674.7234 | 61.8266 | 149.2627
0.68338 | 1029676.7787 | 71.7822 | 173.2975
0.68094 | 1017225.7664 | 76.3198 | 184.2522
0.71232 | 1099594.018 | 69.7678 | 168.4343
0.68459 | 1025725.3228 | 71.3802 | 172.3271
0.69386 | 1032562.6024 | 74.3572 | 179.5142
0.68912 | 1037741.3853 | 71.7412 | 173.1986
0.70538 | 1085040.3351 | 59.081 | 142.6341
0.69604 | 1058362.1885 | 66.786 | 161.2356
0.68205 | 1024219.5381 | 78.3701 | 189.2021
0.70984 | 1087292.557 | 69.2282 | 167.1317
0.6569 | 971746.4236 | 78.7944 | 190.2264
0.67743 | 1012219.2871 | 71.1847 | 171.8551
0.70415 | 1075054.1942 | 74.5312 | 179.9343
0.69591 | 1054998.0895 | 79.1173 | 191.006
0.7075 | 1094745.1472 | 71.8497 | 173.4605
0.72578 | 1141091.4262 | 71.058 | 171.5491
0.67124 | 1015890.9634 | 62.4414 | 150.7469
0.69471 | 1047773.3086 | 77.0538 | 186.0243
0.6865 | 1037549.965 | 72.0019 | 173.8279
0.69434 | 1040476.0383 | 79.8179 | 192.6976
0.68199 | 1020970.7419 | 71.3202 | 172.1822
0.70864 | 1071938.2882 | 72.0648 | 173.9797
0.65389 | 959391.8864 | 66.311 | 160.0889
0.6823 | 1020655.8987 | 65.7037 | 158.6227
0.68309 | 1013854.2026 | 79.0563 | 190.8589
0.68629 | 1033993.1514 | 62.1602 | 150.0679
0.70455 | 1076635.0347 | 84.5886 | 204.215
0.6977 | 1052902.97 | 57.9622 | 139.9331
0.70217 | 1054939.8547 | 73.3432 | 177.0662
0.67079 | 1007649.2703 | 59.8747 | 144.5503
0.69964 | 1076170.6036 | 80.2746 | 193.8
0.69022 | 1041426.0049 | 66.2708 | 159.9919
0.70968 | 1085169.1859 | 76.4026 | 184.4523
0.68505 | 1019419.5531 | 82.3725 | 198.8648
0.70163 | 1065281.6954 | 77.5686 | 187.2671
0.73803 | 1167578.5286 | 69.4852 | 167.7521
0.6799 | 1019191.8097 | 79.8458 | 192.7647
0.67792 | 1006767.1793 | 70.7499 | 170.8053
0.66719 | 976832.7992 | 54.027 | 130.4327
0.6938 | 1047385.9222 | 73.6964 | 177.9189
0.68782 | 1033227.8819 | 80.0133 | 193.1691
0.69309 | 1042206.9399 | 71.6373 | 172.9478
0.70225 | 1050028.9935 | 78.0343 | 188.3914
0.68145 | 1010102.3107 | 75.7286 | 182.8251
0.69964 | 1054900.7909 | 71.8313 | 173.4162
0.68334 | 1019233.2568 | 69.2586 | 167.2049
0.7239 | 1130199.1756 | 81.4793 | 196.7083
0.70033 | 1066745.4066 | 70.7288 | 170.7544
0.68095 | 1009177.1916 | 65.0934 | 157.1493
0.69263 | 1040146.7328 | 69.7322 | 168.3485
0.67588 | 1006571.5554 | 80.8966 | 195.3016
0.69356 | 1037982.2828 | 67.6399 | 163.2972
0.67693 | 1016592.5227 | 67.8734 | 163.8609
0.70009 | 1067886.9781 | 64.996 | 156.9141
0.67546 | 1025512.9387 | 60.5141 | 146.094
0.68431 | 1026735.5162 | 88.9954 | 214.8538
0.69219 | 1045545.3996 | 73.7375 | 178.0181
0.68155 | 1011643.0386 | 77.4646 | 187.0161
0.68109 | 1004410.0139 | 76.9623 | 185.8033
0.67789 | 1019632.5688 | 69.8426 | 168.615
0.67783 | 1013174.8749 | 79.6164 | 192.2109
0.68537 | 1025601.5725 | 85.848 | 207.2554
0.70492 | 1081104.7763 | 80.0386 | 193.2302
0.6862 | 1040269.7862 | 77.4506 | 186.9822
0.69128 | 1045692.8321 | 72.066 | 173.9827
0.68415 | 1031768.1789 | 81.9198 | 197.7718
0.67925 | 1005317.2037 | 61.0625 | 147.4178
0.68051 | 1015619.3162 | 77.8876 | 188.0372
0.67675 | 1021780.074 | 70.108 | 169.2557
0.71629 | 1120099.897 | 78.1349 | 188.6343
0.69381 | 1042296.9363 | 83.1582 | 200.7617
0.69881 | 1053090.2723 | 75.2982 | 181.7859
0.68557 | 1020139.2767 | 57.152 | 137.9771
0.67124 | 1000730.7296 | 77.8064 | 187.8412
0.69229 | 1027282.0132 | 83.6483 | 201.9447
0.7087 | 1087434.8038 | 78.3919 | 189.2549
0.68747 | 1043627.8585 | 82.4607 | 199.0778
0.73391 | 1145260.5254 | 83.1187 | 200.6664
0.67908 | 1010618.241 | 80.8393 | 195.1632
0.70088 | 1052648.7708 | 69.7391 | 168.365
0.67645 | 1002757.2695 | 72.3659 | 174.7066
0.67063 | 1006365.1243 | 81.5136 | 196.7913
0.67168 | 995175.597 | 82.0147 | 198.0009
0.69811 | 1055981.6294 | 64.4541 | 155.6061
0.66336 | 973189.9356 | 64.436 | 155.5622
0.67387 | 1005315.4949 | 69.2732 | 167.2402
0.69623 | 1054176.6488 | 79.0191 | 190.7689
0.67333 | 999817.334 | 80.6589 | 194.7277
0.6984 | 1054246.8361 | 75.4248 | 182.0916
0.68673 | 1027893.1129 | 63.7792 | 153.9766
0.67449 | 1026128.7273 | 70.6423 | 170.5455
0.67756 | 1026886.0467 | 82.325 | 198.7502
0.69786 | 1067429.6734 | 71.0245 | 171.4684
0.66867 | 1000688.7878 | 61.7433 | 149.0615
0.67206 | 996562.9559 | 68.8547 | 166.23
0.68023 | 1007145.6888 | 78.863 | 190.3921
0.67025 | 983301.2638 | 71.6184 | 172.9022
0.6696 | 999915.8558 | 65.5887 | 158.3452
0.66976 | 1009760.4582 | 70.006 | 169.0094
0.67885 | 1000899.9064 | 84.5843 | 204.2045
0.68681 | 1015192.1681 | 70.4189 | 170.0062
0.69221 | 1042446.3588 | 71.7455 | 173.209
0.68318 | 1016203.8195 | 63.3819 | 153.0175
0.69949 | 1069760.4894 | 71.0143 | 171.4436
0.67352 | 1000037.5723 | 79.0867 | 190.9321
0.65897 | 961340.5218 | 77.7057 | 187.5981
0.67937 | 1017668.8082 | 71.3634 | 172.2864
0.69086 | 1055775.9571 | 84.5935 | 204.2268
0.70825 | 1066519.9259 | 69.9235 | 168.8102
0.69223 | 1044096.9419 | 81.9307 | 197.7981
0.6785 | 1000736.8704 | 56.6496 | 136.7641
0.68135 | 1023176.6939 | 69.7294 | 168.3417
0.68235 | 1031731.7623 | 82.1329 | 198.2864
0.68993 | 1041285.206 | 83.545 | 201.6955
0.68377 | 1024440.4422 | 78.19 | 188.7675
0.71009 | 1087679.1494 | 78.2556 | 188.9259
0.69496 | 1047987.6475 | 80.3223 | 193.9152
0.69367 | 1051926.104 | 77.4747 | 187.0406
0.69118 | 1052156.7767 | 68.7995 | 166.0967
0.68429 | 1004482.9033 | 69.9326 | 168.8322
0.71703 | 1085294.7068 | 77.3285 | 186.6876
0.69453 | 1049092.505 | 75.4744 | 182.2114
0.68529 | 1037199.1305 | 83.7804 | 202.2637
0.69204 | 1040291.7762 | 73.3837 | 177.1638
0.69119 | 1039225.4195 | 66.0379 | 159.4297
0.70793 | 1095890.9416 | 67.9921 | 164.1474
0.67983 | 1029225.9639 | 82.2377 | 198.5393
0.7119 | 1080246.7737 | 53.6105 | 129.4272
0.70181 | 1074647.2622 | 83.1777 | 200.8086
0.67706 | 1011417.5029 | 59.5414 | 143.7457
0.68536 | 1037486.3156 | 72.1178 | 174.1077
0.69321 | 1059265.6549 | 81.3916 | 196.4966
0.66544 | 970620.9865 | 69.4896 | 167.7628
0.69723 | 1045695.1478 | 79.8854 | 192.8605
0.70224 | 1075687.2613 | 63.1978 | 152.573
0.71596 | 1098911.1006 | 74.9283 | 180.8929
0.67457 | 1004662.7599 | 74.4164 | 179.6571
0.68268 | 1010600.8152 | 81.9437 | 197.8297
0.68354 | 1042768.8049 | 62.9399 | 151.9503
0.70044 | 1062479.1442 | 70.4346 | 170.0443
0.68417 | 1017932.9349 | 76.0131 | 183.5118
0.68478 | 1011081.1713 | 87.5775 | 211.4307
0.72355 | 1129624.2328 | 84.8055 | 204.7386
0.70519 | 1073695.9434 | 82.7053 | 199.6684
0.67589 | 1011652.8289 | 73.0681 | 176.4019
0.70561 | 1085549.8704 | 76.8706 | 185.5822
0.69278 | 1051961.4893 | 82.4829 | 199.1314
0.69426 | 1037213.3596 | 81.1531 | 195.921
0.70667 | 1059311.056 | 79.8443 | 192.7613
0.68606 | 1034877.633 | 80.3468 | 193.9743
0.69998 | 1064969.7625 | 82.643 | 199.5178
0.6999 | 1078286.0875 | 65.6467 | 158.485
0.70884 | 1085483.6129 | 69.2104 | 167.0886
0.70745 | 1087385.1394 | 91.1694 | 220.1025
0.68112 | 999543.0867 | 71.6395 | 172.9529
0.67617 | 1021771.0095 | 66.3711 | 160.2341
0.68984 | 1034119.8782 | 73.181 | 176.6746
0.69181 | 1040274.3169 | 76.823 | 185.4671
0.6606 | 975819.3631 | 78.6809 | 189.9526
0.67138 | 1014636.6882 | 80.4979 | 194.3391
0.67155 | 994042.0928 | 58.999 | 142.4361
0.68281 | 1045465.2415 | 72.9435 | 176.1011
0.70064 | 1066073.24 | 56.1481 | 135.5534
0.70926 | 1071601.2101 | 83.6335 | 201.9091
0.7046 | 1069317.8975 | 66.6568 | 160.9238
0.6968 | 1064026.666 | 76.2997 | 184.2038
0.68708 | 1043326.0372 | 75.873 | 183.1737
0.68714 | 1035516.6409 | 75.2988 | 181.7875
0.68836 | 1055593.1546 | 76.6229 | 184.9839
0.6931 | 1057620.5396 | 77.9589 | 188.2095
0.70716 | 1081881.6761 | 76.4434 | 184.5507
0.66042 | 974042.8094 | 60.8493 | 146.9032
0.71256 | 1087723.9477 | 70.8819 | 171.1241
0.66749 | 985973.937 | 72.1101 | 174.0893
0.6958 | 1048005.9379 | 69.9385 | 168.8466
0.69022 | 1055140.8713 | 79.781 | 192.6083
0.70131 | 1063545.534 | 70.4183 | 170.0048
0.68431 | 1016974.3873 | 79.1586 | 191.1057
0.70291 | 1073579.5456 | 65.0804 | 157.1179
0.69158 | 1037178.3793 | 78.2682 | 188.9562
0.68256 | 1024455.7897 | 65.9226 | 159.1513
0.7156 | 1113847.4336 | 75.9855 | 183.4453
0.69547 | 1062545.3202 | 70.943 | 171.2716
0.66148 | 969896.1271 | 73.263 | 176.8726
0.69491 | 1066733.2887 | 67.4236 | 162.7749
0.69619 | 1075075.8309 | 82.2527 | 198.5755
0.69803 | 1055056.8999 | 75.5231 | 182.3289
0.65651 | 961617.3129 | 84.0303 | 202.867
0.66753 | 985040.3135 | 75.3901 | 182.0077
0.69937 | 1036297.7539 | 74.0376 | 178.7425
0.6747 | 1006478.7205 | 79.4571 | 191.8265
0.68965 | 1045764.2785 | 66.6022 | 160.792
0.69186 | 1036596.4262 | 74.8632 | 180.7359
0.68523 | 1027241.2417 | 71.419 | 172.4206
0.68025 | 1014028.6152 | 77.796 | 187.8163
0.69787 | 1055451.7047 | 55.8096 | 134.7362
0.7037 | 1070354.038 | 72.4634 | 174.9421
0.68583 | 1023450.0697 | 76.2401 | 184.0598
0.7041 | 1084372.7643 | 73.3707 | 177.1325
0.68952 | 1047142.988 | 68.6362 | 165.7025
0.69444 | 1046601.5941 | 71.7998 | 173.34
0.71782 | 1100098.5225 | 83.248 | 200.9785
0.70124 | 1072051.7075 | 67.1476 | 162.1086
0.69416 | 1041534.1596 | 69.6712 | 168.2012
0.72833 | 1143610.0365 | 94.6244 | 228.4436
0.67495 | 1012470.395 | 83.0361 | 200.4668
0.70043 | 1077344.603 | 59.0768 | 142.624
0.68033 | 1022118.5178 | 73.3423 | 177.0639
0.69869 | 1061527.0887 | 60.5572 | 146.1981
0.66012 | 982209.5437 | 83.8758 | 202.4941
0.68988 | 1042511.0055 | 71.9917 | 173.8034
0.71758 | 1123338.5284 | 74.5795 | 180.0508
0.67846 | 1015425.5824 | 67.6428 | 163.3042
0.69802 | 1061794.796 | 76.0059 | 183.4945
0.66859 | 991170.0942 | 63.1693 | 152.5041
0.6518 | 954107.0353 | 77.9696 | 188.2352
0.70387 | 1071900.1211 | 74.3418 | 179.4769
0.6973 | 1067703.291 | 76.5084 | 184.7077
0.66914 | 1004957.415 | 81.5814 | 196.9549
0.69308 | 1053231.0222 | 66.7452 | 161.1372
0.66787 | 1000271.1862 | 65.7822 | 158.8122
0.70537 | 1085934.0625 | 73.3439 | 177.0678
0.70834 | 1081327.1358 | 81.3116 | 196.3036
0.71003 | 1085420.8944 | 60.9309 | 147.1003
0.71372 | 1089153.7305 | 81.2836 | 196.2361
0.70163 | 1073261.9416 | 74.674 | 180.279
0.69389 | 1034973.4646 | 74.5268 | 179.9237
0.65379 | 947573.0105 | 72.4183 | 174.8332
0.70689 | 1089445.8247 | 61.4288 | 148.3022
0.69209 | 1038424.0914 | 67.812 | 163.7127
0.68618 | 1025065.9407 | 77.9623 | 188.2177
0.68977 | 1054504.2151 | 72.8321 | 175.8324
0.68631 | 1037576.2387 | 75.5402 | 182.3701
0.69664 | 1059157.4953 | 87.2216 | 210.5715
0.70292 | 1065524.7703 | 60.4191 | 145.8647
0.6975 | 1053068.175 | 71.927 | 173.6472
0.67381 | 1006065.1142 | 69.3308 | 167.3795
0.71747 | 1128456.6162 | 64.8903 | 156.659
0.67769 | 1020996.3626 | 86.4308 | 208.6623
0.69635 | 1042510.4143 | 85.8864 | 207.3482
0.70359 | 1092257.3058 | 72.0846 | 174.0276
0.71706 | 1094970.8331 | 87.604 | 211.4947
0.69975 | 1052014.0363 | 57.7438 | 139.4059
0.68065 | 1023373.3303 | 89.1807 | 215.3012
0.6857 | 1024416.8871 | 79.5759 | 192.1131
0.68077 | 1034784.6626 | 66.7623 | 161.1785
0.70685 | 1094719.9304 | 77.3837 | 186.8209
0.67884 | 1012182.145 | 79.4199 | 191.7366
0.68643 | 1035430.45 | 70.0426 | 169.0979
0.7082 | 1070778.1456 | 70.2067 | 169.4939
0.68063 | 1006023.5978 | 73.8535 | 178.298
0.70666 | 1077928.1448 | 64.997 | 156.9166
0.68718 | 1029166.7051 | 63.8167 | 154.067
0.69723 | 1068460.756 | 69.2596 | 167.2074
0.69667 | 1050000.9649 | 78.6564 | 189.8934
0.68409 | 1026763.0609 | 69.6936 | 168.2553
0.6881 | 1040455.6652 | 73.8148 | 178.2046
0.6893 | 1047314.4115 | 66.7349 | 161.1122
0.70642 | 1094847.9221 | 68.0731 | 164.3431
0.70731 | 1068488.4838 | 77.1775 | 186.323
0.70184 | 1067367.6352 | 62.7445 | 151.4787
0.70284 | 1063850.8019 | 77.3311 | 186.6938
0.67817 | 1003428.0008 | 70.4897 | 170.1772
0.72304 | 1121438.9752 | 79.2394 | 191.3009
0.66463 | 968649.1779 | 67.6746 | 163.3811
0.70358 | 1071229.6223 | 82.7039 | 199.6649
0.66949 | 986814.0466 | 58.7504 | 141.836
0.70315 | 1077715.2253 | 84.8418 | 204.8262
0.68437 | 1031509.1209 | 68.1204 | 164.4573
0.70636 | 1076682.135 | 58.5441 | 141.338
0.67308 | 1002686.9718 | 74.4346 | 179.7011
0.67869 | 1012285.5688 | 65.5109 | 158.1572
0.67434 | 1004741.3134 | 70.9824 | 171.3667
0.69499 | 1060744.2226 | 81.2966 | 196.2674
0.69769 | 1075457.3307 | 82.1055 | 198.2201
0.67661 | 1004931.0552 | 74.2119 | 179.1633
0.69196 | 1047162.3463 | 89.7616 | 216.7036
0.68846 | 1049780.4152 | 72.6432 | 175.3761
0.70158 | 1055464.4323 | 57.7814 | 139.4966
0.69125 | 1051354.9515 | 81.0127 | 195.582
0.6754 | 1005710.0109 | 65.402 | 157.8944
0.69878 | 1062631.1048 | 66.7919 | 161.2498
0.70554 | 1057524.3828 | 59.7431 | 144.2327
0.7003 | 1071911.4978 | 71.0747 | 171.5894
0.69441 | 1049051.7782 | 70.1488 | 169.3543
0.67711 | 999956.429 | 68.5745 | 165.5535
0.70101 | 1071951.2921 | 76.4062 | 184.4608
0.7018 | 1063669.4618 | 79.098 | 190.9594
0.67136 | 984928.3715 | 82.6992 | 199.6535
0.69881 | 1067423.6525 | 63.0204 | 152.1447
0.69429 | 1038546.8294 | 63.7381 | 153.8774
0.66742 | 977846.7003 | 73.206 | 176.7348
0.69195 | 1050079.3576 | 62.6103 | 151.1546
0.66129 | 966392.4186 | 67.8573 | 163.8221
0.70129 | 1074309.2677 | 76.5437 | 184.7927
0.66082 | 970102.5003 | 76.6138 | 184.9621
0.69264 | 1051373.6571 | 66.67 | 160.9557
0.72115 | 1113858.6634 | 78.9102 | 190.5061
0.71677 | 1105220.264 | 78.7773 | 190.1853
0.70446 | 1051624.4522 | 81.0889 | 195.766
0.68552 | 1035644.4458 | 70.1499 | 169.3568
0.69677 | 1043210.5765 | 78.1372 | 188.6398
0.66713 | 1001633.2913 | 78.3555 | 189.167
0.68516 | 1025534.4376 | 67.6475 | 163.3154
0.6933 | 1053450.6057 | 75.9218 | 183.2915
0.68727 | 1038293.4948 | 77.183 | 186.3363
0.68554 | 1022858.0055 | 67.4336 | 162.7991
0.70214 | 1068789.3649 | 72.5184 | 175.075
0.68759 | 1020200.8303 | 70.686 | 170.651
0.69305 | 1048270.0722 | 78.9375 | 190.5719
0.69262 | 1061138.9485 | 46.1603 | 111.4408
0.67161 | 999122.6884 | 72.6227 | 175.3267
0.66308 | 982307.5811 | 65.0672 | 157.0861
0.68921 | 1034441.573 | 76.6805 | 185.1231
0.69421 | 1055234.6292 | 77.658 | 187.4829
0.69926 | 1049517.2081 | 73.27 | 176.8895
0.69219 | 1056389.6584 | 74.5031 | 179.8664
0.72252 | 1138305.8051 | 73.1553 | 176.6124
0.68242 | 1019854.1741 | 69.394 | 167.5319
0.68307 | 1032742.184 | 78.9047 | 190.4928
0.68663 | 1041579.0533 | 64.638 | 156.05
0.67748 | 1011418.9318 | 77.9126 | 188.0977
0.70572 | 1079668.5331 | 78.1611 | 188.6976
0.71756 | 1115476.1598 | 86.0516 | 207.7468
0.69737 | 1053531.8444 | 68.2774 | 164.8363
0.68417 | 1014488.7871 | 72.3859 | 174.7551
0.68913 | 1035554.6402 | 57.7158 | 139.3381
0.69147 | 1047391.2195 | 83.8473 | 202.4253
0.66366 | 979406.1908 | 75.3954 | 182.0205
0.66724 | 997278.8161 | 74.6815 | 180.2972
0.68419 | 1009183.4323 | 76.9784 | 185.8424
0.69618 | 1045954.4916 | 67.9026 | 163.9314
0.65354 | 942495.5825 | 72.0577 | 173.9627
0.71903 | 1086663.7659 | 80.4984 | 194.3403
0.70402 | 1077550.3457 | 63.147 | 152.4502
0.67767 | 1000355.2397 | 84.6862 | 204.4505
0.67375 | 999985.8969 | 82.1579 | 198.3468
0.69607 | 1057901.4296 | 88.6807 | 214.0941
0.67009 | 992022.4643 | 78.855 | 190.3729
0.70419 | 1080959.4177 | 69.3774 | 167.4917
0.6846 | 1025671.5128 | 74.8599 | 180.7279
0.68194 | 999581.5187 | 62.4582 | 150.7874
0.6826 | 1029888.3282 | 78.7408 | 190.097
0.71458 | 1088318.4162 | 64.8968 | 156.6747
0.67272 | 1008885.2589 | 81.4633 | 196.6697
0.6729 | 992144.2652 | 66.6373 | 160.8767
0.68601 | 1034163.5814 | 64.2353 | 155.0777
0.71355 | 1102233.7251 | 61.5695 | 148.6418
0.68573 | 1037380.2743 | 77.2885 | 186.591
0.69424 | 1063772.8949 | 67.3012 | 162.4795
0.68386 | 1041875.0189 | 66.9166 | 161.5509
0.67863 | 1015224.7324 | 65.7071 | 158.631
0.71025 | 1079023.6697 | 70.5432 | 170.3063
0.66537 | 987922.8115 | 83.6981 | 202.065
0.69492 | 1049282.6312 | 73.7113 | 177.9547
0.69042 | 1029203.1628 | 67.3728 | 162.6524
0.69458 | 1058409.0085 | 66.9291 | 161.5812
0.68042 | 1013631.6964 | 54.9727 | 132.7158
0.66492 | 988347.8565 | 69.0204 | 166.6299
0.70421 | 1064250.5762 | 62.3509 | 150.5283
0.68274 | 1018613.1444 | 57.3032 | 138.3422
0.69811 | 1063547.2813 | 62.854 | 151.743
0.67759 | 1017434.4668 | 69.7236 | 168.3277
0.69543 | 1032588.9045 | 77.4904 | 187.0785
0.67876 | 1005483.6155 | 59.0439 | 142.5446
0.68579 | 1029290.1073 | 76.6789 | 185.1192
0.65361 | 969579.1382 | 74.1207 | 178.9432
0.70001 | 1053060.5476 | 81.716 | 197.2799
0.70688 | 1086318.6645 | 67.4982 | 162.955
0.70048 | 1056511.0139 | 80.3371 | 193.9508
0.69582 | 1059985.3337 | 82.7483 | 199.772
0.6922 | 1041842.4683 | 89.2754 | 215.5299
0.69521 | 1052475.7181 | 80.6376 | 194.6765
0.68586 | 1033282.4906 | 64.195 | 154.9805
0.71268 | 1100081.2059 | 88.5536 | 213.7874
0.69769 | 1050466.3519 | 80.1971 | 193.6129
0.69171 | 1056438.6489 | 67.0812 | 161.9483
0.67827 | 1012662.1541 | 65.5755 | 158.3132
0.70628 | 1068634.5188 | 85.2909 | 205.9105
0.70299 | 1073842.2549 | 65.0888 | 157.1383
0.69111 | 1037512.2363 | 78.3467 | 189.1457
0.67653 | 1011019.5874 | 71.1945 | 171.8788
0.72747 | 1137656.9447 | 81.6151 | 197.0364
0.67871 | 1014735.1261 | 61.6831 | 148.9162
0.6824 | 1019526.1591 | 77.9708 | 188.2382
0.69078 | 1042503.8112 | 73.6732 | 177.8629
0.6888 | 1032446.4722 | 82.9716 | 200.3111
0.67076 | 1006300.2939 | 65.8932 | 159.0804
0.70052 | 1056052.0978 | 62.1962 | 150.1549
0.70763 | 1089558.132 | 67.6281 | 163.2687
0.68045 | 1020988.614 | 61.9049 | 149.4516
0.68444 | 1021728.456 | 88.5435 | 213.7629
0.71957 | 1110990.9262 | 65.7896 | 158.8301
0.66631 | 984685.2666 | 64.9089 | 156.7039
0.68133 | 1027857.088 | 80.3552 | 193.9946
0.70094 | 1070728.1447 | 69.0386 | 166.6739
0.68522 | 1020233.1819 | 65.3961 | 157.8801
0.68109 | 1030941.8906 | 77.3325 | 186.6973
0.70025 | 1050010.0745 | 71.436 | 172.4618
0.68918 | 1032649.6433 | 79.8109 | 192.6806
0.68958 | 1031885.9457 | 74.7696 | 180.5098
0.68547 | 1041323.3918 | 68.6586 | 165.7565
0.69478 | 1037109.9945 | 74.5859 | 180.0663
0.68446 | 1027318.5639 | 81.5292 | 196.8289
0.68406 | 1035361.0708 | 70.2803 | 169.6717
0.69794 | 1058118.9522 | 79.1212 | 191.0154
0.70372 | 1077488.4536 | 72.9438 | 176.102
0.68814 | 1033842.74 | 73.3416 | 177.0623
0.68226 | 1011947.9069 | 77.4075 | 186.8783
0.6863 | 1037424.3682 | 60.2508 | 145.4583
0.70126 | 1055914.6366 | 71.734 | 173.1813
0.71339 | 1101789.3614 | 63.7197 | 153.8331
0.66814 | 981253.691 | 76.6458 | 185.0393
0.69024 | 1018595.8919 | 73.7357 | 178.0136
0.68759 | 1024620.1914 | 69.1948 | 167.0509
0.69504 | 1054387.5178 | 62.5881 | 151.1012
0.70428 | 1108592.8846 | 88.7323 | 214.2188
0.69653 | 1063305.0877 | 70.885 | 171.1315
0.66841 | 990106.7279 | 66.9475 | 161.6255
0.70681 | 1084457.393 | 63.3006 | 152.8212
0.71686 | 1105757.2248 | 77.0728 | 186.0702
0.67055 | 979895.2047 | 66.0455 | 159.448
0.69609 | 1074433.3706 | 83.8439 | 202.417
0.6947 | 1044158.9153 | 76.7888 | 185.3846
0.68564 | 1025625.4913 | 50.3899 | 121.652
0.73279 | 1136041.2204 | 77.4628 | 187.0118
0.66072 | 971576.1737 | 66.459 | 160.4463
0.67241 | 987869.7332 | 54.2136 | 130.8833
0.69099 | 1041827.9454 | 75.3358 | 181.8767
0.69684 | 1060202.236 | 66.8252 | 161.3302
0.70203 | 1083276.0871 | 85.4151 | 206.2104
0.68373 | 1023674.967 | 74.8086 | 180.604
0.68873 | 1034912.7134 | 65.9987 | 159.3351
0.71713 | 1099744.9578 | 90.9608 | 219.5989
0.69326 | 1045747.3567 | 73.6384 | 177.7789
0.67921 | 1019316.2598 | 65.9154 | 159.1339
0.70033 | 1057624.1872 | 73.7229 | 177.9829
0.67876 | 1014724.6504 | 85.7642 | 207.053
0.70173 | 1059343.0197 | 71.9188 | 173.6273
0.69814 | 1065269.0562 | 72.6935 | 175.4976
0.66716 | 982448.0766 | 70.1551 | 169.3694
0.69536 | 1075551.1492 | 69.9068 | 168.77
0.69721 | 1051379.0807 | 64.2561 | 155.128
0.69675 | 1053500.6558 | 75.5498 | 182.3933
0.67681 | 992739.8564 | 81.1933 | 196.0181
0.68152 | 1019710.3675 | 78.0856 | 188.5152
0.69852 | 1059120.8461 | 62.5559 | 151.0233
0.69301 | 1049365.2313 | 72.1383 | 174.1572
0.69626 | 1056073.0959 | 51.0938 | 123.3513
0.6945 | 1073604.9982 | 71.7808 | 173.2941
0.67445 | 1006815.7726 | 76.2659 | 184.1221
0.69365 | 1058100.9213 | 68.9415 | 166.4394
0.68525 | 1030212.8874 | 81.6653 | 197.1575
0.69645 | 1059860.4618 | 74.2899 | 179.3517
0.66641 | 983259.1615 | 75.3804 | 181.9845
0.69316 | 1049852.8428 | 83.8397 | 202.4069
0.69067 | 1037410.3921 | 76.7382 | 185.2623
0.69792 | 1068662.8592 | 84.7592 | 204.6269
0.66694 | 1004558.7417 | 80.228 | 193.6876
0.68305 | 1024713.746 | 74.9594 | 180.9681
0.67977 | 1007013.8314 | 66.9531 | 161.639
0.67796 | 1020682.3212 | 68.0428 | 164.2699
0.67102 | 1009216.8284 | 81.6966 | 197.233
0.69955 | 1060201.7955 | 60.4067 | 145.8348
0.69591 | 1045029.2144 | 84.6507 | 204.3649
0.6804 | 1023353.4097 | 69.8395 | 168.6075
0.68153 | 1021241.8983 | 78.2016 | 188.7954
0.67426 | 1007146.3651 | 69.359 | 167.4473
0.70664 | 1053989.3646 | 85.5687 | 206.5811
0.69427 | 1052071.5265 | 58.2022 | 140.5126
0.68959 | 1038220.4212 | 69.9099 | 168.7773
0.69969 | 1057090.5475 | 82.0662 | 198.1254
0.70308 | 1078254.151 | 69.3124 | 167.3349
0.67815 | 999877.1029 | 66.6371 | 160.8762
0.68839 | 1044492.3145 | 83.9892 | 202.7679
0.69038 | 1048797.3097 | 80.9179 | 195.3531
0.67377 | 1004248.0303 | 61.494 | 148.4598
0.69299 | 1061229.2056 | 76.5171 | 184.7286
0.69425 | 1049044.4198 | 66.3687 | 160.2281
0.68343 | 1021717.8899 | 65.663 | 158.5245
0.69393 | 1047233.1186 | 75.4054 | 182.0448
0.69172 | 1047176.1694 | 64.5069 | 155.7334
0.71814 | 1111424.7535 | 73.7303 | 178.0006
0.71603 | 1102853.584 | 69.612 | 168.0583
0.68595 | 1027744.8345 | 81.3859 | 196.483
0.67904 | 1016890.411 | 73.1063 | 176.4941
0.71703 | 1121852.7019 | 68.7197 | 165.9041
0.68786 | 1030195.8122 | 84.4257 | 203.8216
0.70226 | 1051520.1009 | 79.7547 | 192.5449
0.6849 | 1034386.4944 | 68.2582 | 164.7898
0.70189 | 1062042.1915 | 86.8058 | 209.5677
0.68445 | 1038458.2722 | 68.571 | 165.545
0.67241 | 1004895.0553 | 67.5436 | 163.0647
0.65961 | 974201.7869 | 70.3616 | 169.8679
0.67354 | 1000853.3425 | 65.2665 | 157.5673
0.69141 | 1055458.9371 | 77.9896 | 188.2835
0.70238 | 1076343.4919 | 74.8092 | 180.6054
0.69733 | 1062701.7571 | 84.3585 | 203.6593
0.68556 | 1031411.7397 | 67.9959 | 164.1566
0.6885 | 1042309.1044 | 66.4927 | 160.5277
0.68401 | 1012898.6087 | 68.7604 | 166.0023
0.69155 | 1040854.2598 | 64.1329 | 154.8304
0.6882 | 1032509.416 | 86.07 | 207.7913
0.67235 | 996496.0574 | 77.3489 | 186.7367
0.67872 | 1004038.7803 | 78.7946 | 190.2271
0.69541 | 1048621.1275 | 78.6953 | 189.9874
0.68757 | 1023675.248 | 67.2883 | 162.4483
0.71585 | 1100183.8965 | 65.4877 | 158.1012
0.6842 | 1018862.848 | 70.943 | 171.2715
0.68785 | 1029448.7566 | 79.7165 | 192.4526
0.7122 | 1105804.9718 | 72.0966 | 174.0567
0.68085 | 1017759.0996 | 52.8132 | 127.5024
0.68973 | 1038304.294 | 76.909 | 185.6747
0.70059 | 1080541.1497 | 67.7421 | 163.5439
0.69014 | 1044781.7619 | 73.9032 | 178.4182
0.6955 | 1059078.263 | 56.1921 | 135.6598
0.67348 | 1001435.615 | 72.5894 | 175.2463
0.68229 | 1030003.4696 | 71.9197 | 173.6295
0.67706 | 1018474.8758 | 81.7114 | 197.2689
0.68439 | 1029972.6692 | 78.5412 | 189.6153
0.684 | 1034176.6944 | 58.0713 | 140.1965
0.6943 | 1056918.7821 | 81.0833 | 195.7523
0.65743 | 963824.0267 | 78.0886 | 188.5225
0.70751 | 1082302.1274 | 78.1536 | 188.6794
0.68299 | 1017166.0837 | 83.9521 | 202.6783
0.7075 | 1071490.7992 | 80.6684 | 194.7508
0.68814 | 1051785.0278 | 67.1326 | 162.0725
0.69473 | 1061355.8737 | 70.9388 | 171.2614
0.69475 | 1049145.5321 | 70.1208 | 169.2865
0.68683 | 1018570.4659 | 72.396 | 174.7795
0.66536 | 984828.6522 | 84.9134 | 204.999
0.67592 | 1003554.6421 | 74.1857 | 179.1002
0.68026 | 1025446.8706 | 69.9961 | 168.9856
0.69347 | 1041282.1379 | 84.6433 | 204.3471
0.70426 | 1072746.7555 | 77.8211 | 187.8768
0.68704 | 1032008.7942 | 67.8889 | 163.8984
0.6947 | 1059352.7404 | 84.4108 | 203.7857
0.69791 | 1055044.1257 | 70.1955 | 169.467
0.67524 | 1009338.1488 | 72.9717 | 176.1693
0.71091 | 1080561.2944 | 74.2011 | 179.1374
0.70548 | 1079737.8079 | 86.2751 | 208.2865
0.66953 | 1003455.7594 | 67.7367 | 163.5309
0.67408 | 1019162.8166 | 70.2052 | 169.4904
0.68385 | 1036109.5154 | 77.4066 | 186.8761
0.7022 | 1056478.0817 | 83.2827 | 201.0623
0.69442 | 1050864.5471 | 64.4821 | 155.6735
0.65457 | 953712.8754 | 63.9947 | 154.4969
0.70135 | 1079102.7399 | 88.2671 | 213.0957
0.70003 | 1067624.4607 | 79.5932 | 192.155
0.68017 | 1007785.5684 | 68.7999 | 166.0977
0.66582 | 991301.9633 | 69.1337 | 166.9034
0.67997 | 1032461.2069 | 72.3371 | 174.6372
0.69287 | 1055841.7418 | 77.8539 | 187.9559
0.68764 | 1029014.2898 | 74.7713 | 180.514
0.69255 | 1067106.1195 | 80.4141 | 194.1367
0.67267 | 1006109.4762 | 59.6157 | 143.9251
0.69653 | 1064281.8017 | 81.3874 | 196.4867
0.69013 | 1057674.2981 | 61.086 | 147.4745
0.6705 | 1000998.9827 | 81.0947 | 195.7798
0.67866 | 1011892.4279 | 57.2673 | 138.2555
0.67871 | 1027293.422 | 64.4952 | 155.7053
0.6856 | 1022026.6936 | 65.4751 | 158.0708
0.6743 | 995681.0512 | 64.8562 | 156.5768
0.69149 | 1040159.6038 | 71.6625 | 173.0087
0.72665 | 1136223.6915 | 79.8857 | 192.8611
0.6592 | 967781.0299 | 81.2674 | 196.1968
0.64998 | 941299.8143 | 70.6975 | 170.6788
0.68695 | 1032348.9662 | 73.3626 | 177.1129
0.68787 | 1039163.2771 | 70.1481 | 169.3526
0.67125 | 986781.969 | 78.3022 | 189.0381
0.6879 | 1041595.0376 | 62.4704 | 150.8169
0.67802 | 991246.6811 | 66.0751 | 159.5194
0.70944 | 1090533.4052 | 63.4197 | 153.1086
0.67561 | 1000720.4749 | 65.0582 | 157.0644
0.67502 | 1022123.6013 | 75.6599 | 182.6591
0.69024 | 1039388.854 | 64.1644 | 154.9065
0.68269 | 1024404.0642 | 61.9086 | 149.4606
0.68669 | 1030982.7234 | 79.97 | 193.0648
0.69618 | 1046790.3439 | 64.0697 | 154.678
0.68412 | 1016503.8857 | 74.2429 | 179.2382
0.70468 | 1066186.6635 | 70.2556 | 169.612
0.69424 | 1051872.0519 | 83.8697 | 202.4794
0.70604 | 1078823.6897 | 78.0639 | 188.4628
0.69189 | 1030721.8243 | 59.3776 | 143.3502
0.69532 | 1071968.7157 | 85.3401 | 206.0293
0.6977 | 1067270.9279 | 72.8754 | 175.9369
0.6963 | 1058758.2257 | 68.7216 | 165.9086
0.70641 | 1063641.8077 | 76.925 | 185.7134
0.69468 | 1057017.6995 | 79.4087 | 191.7096
0.67994 | 1017225.1462 | 69.4678 | 167.7101
0.69807 | 1070296.7027 | 71.1935 | 171.8762
0.70179 | 1064177.9109 | 64.6139 | 155.9917
0.70479 | 1074085.7276 | 76.148 | 183.8376
0.69306 | 1051479.0648 | 87.037 | 210.1258
0.71023 | 1087532.7019 | 72.949 | 176.1145
0.6998 | 1069840.9942 | 94.9654 | 229.2667
0.68841 | 1027208.643 | 65.5384 | 158.2237
0.68305 | 1023119.9501 | 79.7836 | 192.6147
0.68054 | 1011346.6968 | 83.772 | 202.2435
0.67916 | 1022038.626 | 76.8827 | 185.6112
0.68294 | 1043265.6576 | 63.0255 | 152.1569
0.67804 | 1027701.7775 | 70.3801 | 169.9126
0.65988 | 951832.1296 | 59.5112 | 143.6728
0.67243 | 1002074.2874 | 74.4608 | 179.7642
0.6882 | 1030182.5463 | 94.3445 | 227.7679
0.70532 | 1069228.1401 | 70.3916 | 169.9405
0.65807 | 966533.9213 | 64.5096 | 155.7401
0.71203 | 1086186.7283 | 68.9444 | 166.4466
0.71515 | 1116903.9605 | 78.036 | 188.3955
0.68769 | 1040043.4695 | 75.5677 | 182.4365
0.68706 | 1042662.2007 | 68.5875 | 165.5848
0.69448 | 1055071.6596 | 79.2862 | 191.4139
0.68917 | 1027864.1667 | 61.5332 | 148.5544
0.70078 | 1065148.4885 | 75.9458 | 183.3494
0.68142 | 1032664.1535 | 72.1974 | 174.2999
0.68838 | 1036067.5097 | 64.1326 | 154.8298
0.68587 | 1029146.9065 | 74.9745 | 181.0045
0.68582 | 1044995.9377 | 77.8453 | 187.9351
0.69213 | 1048837.313 | 72.1933 | 174.2901
0.69524 | 1066428.8298 | 81.7025 | 197.2473
0.7028 | 1072523.4864 | 74.4987 | 179.8557
0.67854 | 1004387.3228 | 66.8991 | 161.5087
0.6834 | 1005351.0102 | 84.3954 | 203.7484
0.70544 | 1078996.3142 | 87.7476 | 211.8414
0.65651 | 959692.3345 | 60.3105 | 145.6025
0.68422 | 1037423.0353 | 86.9828 | 209.9951
0.71384 | 1071061.3704 | 78.3051 | 189.0453
0.66431 | 982857.2085 | 61.0406 | 147.365
0.69583 | 1056175.7301 | 74.0979 | 178.888
0.67556 | 1003846.2573 | 57.0726 | 137.7854
0.68206 | 1040444.6145 | 76.2208 | 184.0133
0.70585 | 1086136.2909 | 86.7409 | 209.411
0.70255 | 1088132.079 | 71.9187 | 173.6272
0.68758 | 1028396.668 | 73.0147 | 176.2731
0.71204 | 1093241.9197 | 69.3917 | 167.5263
0.68393 | 1036792.4482 | 74.0037 | 178.6608
0.70756 | 1077129.8711 | 81.0858 | 195.7583
0.69892 | 1055862.3314 | 66.3531 | 160.1905
0.69965 | 1056045.0087 | 73.684 | 177.8888
0.68529 | 1024626.9904 | 69.4274 | 167.6125
0.68874 | 1040728.1779 | 78.8595 | 190.3837
0.70758 | 1091314.0978 | 78.7966 | 190.2317
0.68624 | 1033561.2923 | 76.9469 | 185.7663
0.69237 | 1044316.5656 | 70.0378 | 169.0862
0.68873 | 1038693.7782 | 75.3559 | 181.9252
0.70192 | 1086634.2613 | 79.0908 | 190.9421
0.69934 | 1065134.9039 | 77.8117 | 187.8539
0.6746 | 999516.3421 | 74.0324 | 178.7301
0.69258 | 1026553.0916 | 80.513 | 194.3756
0.71807 | 1111962.1457 | 69.9877 | 168.9653
0.68402 | 1036525.8844 | 60.3987 | 145.8155
0.68305 | 1008737.2733 | 84.7225 | 204.5383
0.67552 | 993990.89 | 58.6495 | 141.5923
0.67904 | 1009462.3479 | 82.2898 | 198.6651
0.6544 | 948152.6983 | 72.3951 | 174.7771
0.69034 | 1043328.8865 | 85.1948 | 205.6785
0.69596 | 1046294.3277 | 75.1187 | 181.3526
0.66127 | 964562.3944 | 77.694 | 187.57
0.69915 | 1062913.3596 | 79.3421 | 191.5487
0.68897 | 1031861.0693 | 73.1133 | 176.5111
0.67614 | 1007877.4775 | 69.2169 | 167.1043
0.6879 | 1041049.5715 | 50.9791 | 123.0745
0.70337 | 1068917.473 | 68.6839 | 165.8176
0.68745 | 1043475.2993 | 87.0793 | 210.2279
0.69405 | 1048085.4101 | 68.2494 | 164.7686
0.72187 | 1114882.6731 | 80.6021 | 194.5906
0.71726 | 1095946.3471 | 74.5905 | 180.0773
0.67548 | 1035993.6562 | 71.2166 | 171.932
0.68611 | 1034399.9709 | 92.2641 | 222.7452
0.70344 | 1064547.3523 | 60.3191 | 145.6231
0.66378 | 968912.181 | 72.9346 | 176.0798
0.71754 | 1113376.9239 | 73.9222 | 178.4641
0.69409 | 1057017.1578 | 62.8298 | 151.6847
0.66719 | 987797.2698 | 65.0279 | 156.9913
0.68475 | 1034090.5639 | 77.1543 | 186.2669
0.64696 | 955189.8694 | 60.9132 | 147.0574
0.6915 | 1053891.959 | 81.1791 | 195.9836
0.66337 | 981345.3382 | 73.9166 | 178.4504
0.69207 | 1041385.0676 | 62.0813 | 149.8776
0.69157 | 1033998.0103 | 67.5743 | 163.1389
0.67353 | 1012915.5591 | 64.0812 | 154.7057
0.69659 | 1053150.5578 | 80.5294 | 194.4151
0.70332 | 1079178.9739 | 74.9055 | 180.8378
0.69396 | 1037041.9246 | 50.3399 | 121.5312
0.65493 | 939090.3041 | 69.9986 | 168.9915
0.70745 | 1094458.7244 | 75.523 | 182.3287
0.68261 | 1015816.1312 | 76.0282 | 183.5483
0.69941 | 1069134.026 | 58.712 | 141.7433
0.68645 | 1025417.636 | 57.3745 | 138.5143
0.67443 | 1007542.8931 | 69.3686 | 167.4705
0.67146 | 994485.8301 | 75.667 | 182.6762
0.67915 | 1011802.4307 | 76.1132 | 183.7534
0.68124 | 1022466.8198 | 67.8338 | 163.7653
0.67657 | 1012364.5035 | 65.7213 | 158.6653
0.71847 | 1106862.1349 | 78.7076 | 190.017
0.68728 | 1038610.3173 | 58.042 | 140.1258
0.67681 | 1013228.2849 | 73.0827 | 176.4373
0.69705 | 1067677.4201 | 72.9712 | 176.168
0.69755 | 1072336.4039 | 69.4788 | 167.7366
0.67144 | 994800.3179 | 72.3268 | 174.6124
0.70084 | 1061443.1089 | 66.5215 | 160.5972
0.70785 | 1076885.4507 | 76.0163 | 183.5196
0.69385 | 1045232.1178 | 64.1898 | 154.9678
0.71401 | 1096911.0216 | 73.1892 | 176.6944
0.6894 | 1036226.1257 | 73.3765 | 177.1466
0.70922 | 1065861.7264 | 57.7224 | 139.3542
0.67404 | 1000920.8936 | 74.1978 | 179.1294
0.68478 | 1029626.8908 | 78.4701 | 189.4435
0.68217 | 1023677.2507 | 86.0889 | 207.8369
0.68624 | 1038825.9243 | 66.4892 | 160.5191
0.67243 | 1010088.6236 | 65.9774 | 159.2835
0.70075 | 1070501.9387 | 70.9349 | 171.2519
0.71393 | 1096749.7944 | 85.0362 | 205.2954
0.69042 | 1049146.7821 | 72.3936 | 174.7735
0.68951 | 1046182.146 | 70.4543 | 170.0918
0.69622 | 1050988.6392 | 78.7821 | 190.1969
0.69908 | 1070586.5286 | 58.8195 | 142.0028
0.70194 | 1047415.9775 | 72.8438 | 175.8604
0.6835 | 1016011.3309 | 70.817 | 170.9675
0.69021 | 1030741.7568 | 74.6179 | 180.1437
0.69763 | 1032536.7154 | 77.2419 | 186.4784
0.69268 | 1039282.1557 | 69.6692 | 168.1964
0.70549 | 1087677.2345 | 77.0348 | 185.9784
0.69209 | 1043867.9214 | 73.6301 | 177.7587
0.6974 | 1068320.3834 | 65.2737 | 157.5846
0.69165 | 1059504.129 | 83.2097 | 200.8859
0.6689 | 1001677.8898 | 78.1524 | 188.6767
0.67802 | 1005515.6997 | 83.598 | 201.8234
0.68958 | 1037981.9392 | 78.5556 | 189.65
0.67817 | 1011389.6008 | 67.4028 | 162.7248
0.6885 | 1053141.6472 | 72.8723 | 175.9293
0.6961 | 1050243.9387 | 71.8781 | 173.5291
0.69236 | 1057258.1521 | 58.6505 | 141.5949
0.68585 | 1034942.689 | 71.8895 | 173.5566
0.6926 | 1026582.7049 | 68.8252 | 166.1587
0.70564 | 1077670.6106 | 51.4462 | 124.202
0.68811 | 1026900.5343 | 81.0639 | 195.7056
0.70497 | 1075638.959 | 65.0152 | 156.9605
0.67737 | 989719.6777 | 71.3281 | 172.2013
0.72268 | 1142204.3504 | 84.1855 | 203.2417
0.71263 | 1094190.848 | 76.7194 | 185.2171
0.71191 | 1093376.2563 | 67.4575 | 162.8567
0.68714 | 1049232.9747 | 76.2966 | 184.1963
0.68331 | 1012236.3592 | 77.3107 | 186.6446
0.71546 | 1110631.3822 | 84.3218 | 203.5708
0.69909 | 1056798.9216 | 70.1931 | 169.4612
0.67822 | 995405.9627 | 56.3709 | 136.0914
0.6931 | 1054854.1267 | 76.0452 | 183.5893
0.70271 | 1085901.8697 | 67.8727 | 163.8591
0.70422 | 1076149.8417 | 79.7509 | 192.5357
0.70733 | 1057755.7086 | 76.3998 | 184.4455
0.69341 | 1057199.075 | 79.5684 | 192.0952
0.70615 | 1075508.9566 | 88.8698 | 214.5508
0.69548 | 1059948.3114 | 76.1511 | 183.845
0.69472 | 1053054.0702 | 83.9059 | 202.5668
0.69071 | 1037734.1425 | 76.7578 | 185.3096
0.66933 | 980401.3396 | 64.3351 | 155.3186
0.67648 | 1019202.5693 | 80.2501 | 193.7408
0.69151 | 1045337.077 | 77.2305 | 186.4509
0.68341 | 1021011.8189 | 88.6155 | 213.9367
0.71747 | 1108154.4344 | 69.8057 | 168.5258
0.6656 | 985685.5377 | 68.7648 | 166.0128
0.68656 | 1039558.3715 | 80.6515 | 194.71
0.69168 | 1045471.0831 | 60.2453 | 145.445
0.69445 | 1056762.8487 | 76.0353 | 183.5655
0.69817 | 1058442.2669 | 74.0082 | 178.6715
0.67811 | 1009548.6214 | 65.3895 | 157.8643
0.6801 | 1022361.5694 | 66.0257 | 159.4002
0.6918 | 1042126.7567 | 80.2378 | 193.7113
0.71159 | 1084631.4365 | 56.2461 | 135.7901
0.67966 | 1020711.7988 | 78.3767 | 189.2182
0.70258 | 1076590.4427 | 79.9704 | 193.0655
0.68412 | 1035157.0431 | 77.304 | 186.6283
0.65861 | 968622.0728 | 70.0652 | 169.1525
0.69735 | 1066722.2729 | 58.7521 | 141.8401
0.69812 | 1083305.7604 | 82.653 | 199.5421
0.70583 | 1085300.5084 | 65.5188 | 158.1763
0.69088 | 1039213.2755 | 67.9826 | 164.1245
0.69259 | 1060670.012 | 87.8572 | 212.106
0.6823 | 1020319.9946 | 69.4381 | 167.6385
0.66629 | 981173.2509 | 66.5545 | 160.6768
0.68327 | 1011611.1983 | 86.3891 | 208.5616
0.70592 | 1088236.3672 | 63.6018 | 153.5482
0.6795 | 1035032.8079 | 64.9054 | 156.6954
0.69775 | 1067066.7568 | 72.8567 | 175.8916
0.6908 | 1051144.592 | 81.0492 | 195.6701
0.67917 | 1013709.6641 | 82.6212 | 199.4653
0.70113 | 1065680.8334 | 79.3231 | 191.503
0.70442 | 1080966.7479 | 91.0205 | 219.7428
0.73184 | 1141772.468 | 57.5708 | 138.9883
0.67481 | 1011351.8577 | 74.885 | 180.7884
0.71095 | 1099858.802 | 65.9197 | 159.1442
0.70696 | 1080033.5414 | 84.5664 | 204.1614
0.68183 | 1011458.0527 | 83.6232 | 201.8842
0.6694 | 992209.8676 | 68.8265 | 166.1618
0.68033 | 1026971.9284 | 80.4044 | 194.1134
0.67521 | 1011624.9516 | 74.8611 | 180.7308
0.70216 | 1060712.0177 | 62.3252 | 150.4663
0.69236 | 1051474.7186 | 75.8824 | 183.1963
0.70626 | 1087847.688 | 91.3608 | 220.5645
0.68504 | 1026854.9295 | 69.6034 | 168.0374
0.6887 | 1038346.5659 | 48.1812 | 116.3197
0.68153 | 1026869.036 | 89.3689 | 215.7555
0.65033 | 944957.6285 | 69.2234 | 167.1201
0.70424 | 1058164.5298 | 71.5833 | 172.8174
0.71188 | 1088756.7416 | 78.1362 | 188.6375
0.68128 | 1015505.4852 | 74.4794 | 179.8092
0.68613 | 1024361.0874 | 69.3245 | 167.3641
0.68099 | 1026708.1603 | 62.202 | 150.1689
0.69416 | 1055622.2897 | 81.0008 | 195.5532
0.69942 | 1052675.9464 | 73.4736 | 177.3809
0.68091 | 1022170.5304 | 76.6943 | 185.1564
0.67886 | 1022161.5462 | 81.7888 | 197.4557
0.69372 | 1048623.3954 | 55.7795 | 134.6637
0.69276 | 1051234.5007 | 82.4275 | 198.9975
0.69133 | 1045519.2571 | 78.9257 | 190.5435
0.70281 | 1063175.3513 | 86.7837 | 209.5143
0.70137 | 1073710.0832 | 69.189 | 167.037
0.69173 | 1040574.3965 | 69.8512 | 168.6356
0.67652 | 1021016.616 | 77.1048 | 186.1475
0.69977 | 1054984.191 | 84.9854 | 205.173
0.67663 | 1007822.3124 | 91.1532 | 220.0634
0.68945 | 1049377.699 | 72.511 | 175.057
0.6923 | 1050578.2414 | 74.0806 | 178.8464
0.6785 | 1042988.6916 | 67.5165 | 162.9992
0.689 | 1047439.3868 | 77.6681 | 187.5075
0.69199 | 1050294.467 | 80.8447 | 195.1763
0.69025 | 1049251.4349 | 80.9674 | 195.4726
0.68967 | 1044971.2402 | 70.3559 | 169.8541
0.70266 | 1065041.9761 | 88.4422 | 213.5184
0.71548 | 1109300.8541 | 66.0648 | 159.4945
0.69266 | 1069016.2356 | 75.3418 | 181.8913
0.69286 | 1060217.782 | 74.7008 | 180.3438
0.67462 | 1000451.5276 | 72.3954 | 174.7779
0.7005 | 1066078.0737 | 79.4248 | 191.7485
0.66784 | 1006989.9161 | 75.3405 | 181.8879
0.70717 | 1084482.1208 | 63.2702 | 152.7478
0.69927 | 1060758.8076 | 66.6277 | 160.8535
0.6684 | 980446.4989 | 71.608 | 172.877
0.71198 | 1098660.5326 | 78.836 | 190.3269
0.68564 | 1026077.4229 | 73.0566 | 176.3743
0.69887 | 1064863.164 | 70.7511 | 170.8082
0.70765 | 1089773.3851 | 61.3815 | 148.188
0.69071 | 1056147.4835 | 81.4451 | 196.6259
0.68574 | 1028863.9336 | 73.244 | 176.8266
0.69843 | 1043448.9112 | 72.9326 | 176.075
0.69205 | 1044163.2333 | 71.8797 | 173.533
0.68956 | 1038050.4274 | 68.0364 | 164.2545
0.67081 | 990887.8738 | 66.7603 | 161.1737
0.7064 | 1065928.3664 | 60.1445 | 145.2018
0.69897 | 1053347.6586 | 76.9916 | 185.8741
0.70901 | 1085918.6714 | 69.2211 | 167.1146
0.67147 | 1003552.7071 | 67.5474 | 163.0738
0.70778 | 1086263.1062 | 93.9233 | 226.7509
0.68224 | 1027100.5905 | 71.7224 | 173.1532
0.67089 | 982892.2866 | 71.5204 | 172.6654
0.70412 | 1070053.3344 | 72.3991 | 174.7869
0.69226 | 1020769.5685 | 66.6426 | 160.8895
0.69799 | 1043799.2073 | 50.3754 | 121.6171
0.684 | 1045697.4194 | 83.044 | 200.4859
0.68066 | 1006256.2549 | 72.1693 | 174.2322
0.69814 | 1052786.6068 | 78.6996 | 189.9976
0.67794 | 990656.9978 | 83.8315 | 202.3872
0.69788 | 1054014.4504 | 50.876 | 122.8256
0.70158 | 1086600.7869 | 79.1263 | 191.0278
0.71544 | 1093395.4357 | 65.1212 | 157.2165
0.66809 | 985739.3534 | 70.2922 | 169.7004
0.69396 | 1044889.6737 | 81.7883 | 197.4545
0.69098 | 1039315.5313 | 59.9514 | 144.7356
0.70408 | 1062702.3035 | 90.9583 | 219.5926
0.69086 | 1037407.073 | 75.0303 | 181.1392
0.69447 | 1058582.6332 | 72.6192 | 175.3183
0.69716 | 1057891.6589 | 55.9954 | 135.1848
0.66572 | 984315.0522 | 73.8786 | 178.3588
0.70044 | 1056547.4647 | 73.6359 | 177.7728
0.68939 | 1046998.4723 | 66.7058 | 161.0421
0.69816 | 1061071.1833 | 74.1929 | 179.1176
0.69956 | 1061033.9847 | 53.5861 | 129.3684
0.70722 | 1086556.1775 | 100.8973 | 243.5876
0.68463 | 1041004.0404 | 70.6046 | 170.4546
0.69132 | 1036839.9216 | 73.2156 | 176.7582
0.66983 | 993650.1247 | 79.0148 | 190.7587
0.68825 | 1035816.1829 | 76.8898 | 185.6285
0.69117 | 1034997.6628 | 82.891 | 200.1166
0.6927 | 1044954.3754 | 85.0359 | 205.2949
0.68286 | 1033030.8918 | 66.445 | 160.4125
0.67615 | 1010578.8074 | 71.2041 | 171.9019
0.71675 | 1103689.0003 | 78.6702 | 189.9268
0.71377 | 1111470.6611 | 62.067 | 149.843
0.69304 | 1046497.566 | 78.8047 | 190.2514
0.69701 | 1070271.1647 | 51.7252 | 124.8758
0.68984 | 1038790.1937 | 83.3556 | 201.2382
0.68862 | 1038768.0409 | 78.7248 | 190.0584
0.68928 | 1050612.8581 | 76.1596 | 183.8654
0.67338 | 999866.2735 | 70.7393 | 170.7797
0.70097 | 1071122.074 | 92.0857 | 222.3146
0.69527 | 1067436.2833 | 80.1239 | 193.4362
0.66734 | 990726.2613 | 79.0758 | 190.9059
0.7087 | 1082882.0192 | 81.3876 | 196.4871
0.69611 | 1047504.1761 | 55.5292 | 134.0593
0.69451 | 1053878.1132 | 60.4822 | 146.017
0.68001 | 1027503.9632 | 66.367 | 160.224
0.69206 | 1042575.3335 | 76.9657 | 185.8116
0.70816 | 1089656.4772 | 76.8834 | 185.6129
0.69785 | 1058021.1774 | 70.3991 | 169.9585
0.70038 | 1061178.5088 | 88.4078 | 213.4354
0.68107 | 1017531.1766 | 61.6389 | 148.8095
0.67242 | 986474.233 | 76.5632 | 184.8399
0.67496 | 998386.1793 | 72.2099 | 174.3302
0.69026 | 1045725.8556 | 72.174 | 174.2434
0.67374 | 1009773.1983 | 57.5649 | 138.9739
0.67696 | 1005953.3758 | 77.5512 | 187.2251
0.66431 | 982328.0595 | 77.1724 | 186.3107
0.73311 | 1151117.1355 | 72.9099 | 176.0201
0.69666 | 1054274.0202 | 73.8764 | 178.3534
0.68073 | 1020342.8994 | 74.8943 | 180.8108
0.69568 | 1069521.1231 | 71.3872 | 172.344
0.71253 | 1084040.1265 | 76.0187 | 183.5253
0.7045 | 1081554.4331 | 73.1386 | 176.5722
0.68757 | 1026795.5103 | 81.1199 | 195.8407
0.6988 | 1062679.0062 | 79.3638 | 191.6011
0.70323 | 1081853.3226 | 63.4703 | 153.2308
0.66934 | 987752.8026 | 78.6028 | 189.764
0.68528 | 1029827.2154 | 69.7802 | 168.4642
0.66628 | 999543.4154 | 84.5663 | 204.1611
0.69534 | 1054310.0194 | 75.7613 | 182.9039
0.67757 | 1012842.8057 | 57.4135 | 138.6085
0.70847 | 1069308.6405 | 78.7692 | 190.1656
0.69788 | 1063605.9691 | 64.9405 | 156.7802
0.69893 | 1052792.2928 | 81.8211 | 197.5335
0.69339 | 1066913.8503 | 88.2723 | 213.1081
0.70727 | 1081075.6366 | 78.0906 | 188.5273
0.69214 | 1064939.9638 | 75.1113 | 181.3348
0.71358 | 1092392.9979 | 76.4914 | 184.6665
0.68903 | 1039777.5207 | 71.8919 | 173.5623
0.6721 | 1005520.8749 | 61.7249 | 149.0171
0.69941 | 1067344.933 | 86.0254 | 207.6836
0.6941 | 1046478.7229 | 70.0077 | 169.0135
0.67562 | 1007202.8835 | 69.2228 | 167.1186
0.66716 | 988336.1283 | 65.0668 | 157.0853
0.68506 | 1002904.3174 | 89.5922 | 216.2946
0.68791 | 1039503.7911 | 85.1824 | 205.6486
0.67878 | 1027177.7998 | 70.2676 | 169.6411
0.6851 | 1016992.222 | 67.0904 | 161.9706
0.6805 | 1024815.8254 | 80.2159 | 193.6582
0.70004 | 1061906.6773 | 72.2668 | 174.4675
0.67935 | 1014168.8181 | 87.1069 | 210.2947
0.67461 | 1004143.7994 | 82.3941 | 198.917
0.72124 | 1120716.1184 | 81.0941 | 195.7784
0.69404 | 1056831.4988 | 62.8231 | 151.6684
0.68692 | 1024104.3108 | 76.1671 | 183.8836
0.68592 | 1035606.1795 | 90.9866 | 219.6612
0.68129 | 1028186.9481 | 70.0777 | 169.1826
0.70506 | 1077169.6507 | 82.9881 | 200.3509
0.71857 | 1125413.5814 | 74.6138 | 180.1337
0.68871 | 1036673.821 | 87.8299 | 212.0401
0.69019 | 1046050.1817 | 77.4311 | 186.9352
0.67616 | 1003835.3718 | 61.251 | 147.8729
0.68294 | 1027101.9257 | 73.4179 | 177.2464
0.72323 | 1118224.9001 | 61.6273 | 148.7814
0.68027 | 1028306.3142 | 72.0143 | 173.858
0.67437 | 990877.2657 | 81.3347 | 196.3594
0.67309 | 1008492.8044 | 67.7163 | 163.4815
0.68964 | 1020931.3863 | 84.927 | 205.032
0.69698 | 1059575.5956 | 62.7207 | 151.4212
0.69918 | 1043067.7114 | 76.2168 | 184.0037
0.69603 | 1047125.4069 | 87.7821 | 211.9247
0.68021 | 1021062.6544 | 76.6897 | 185.1454
0.6835 | 1021266.7007 | 82.709 | 199.6771
0.6897 | 1039408.5215 | 92.2285 | 222.6593
0.67857 | 1007536.212 | 75.8571 | 183.1351
0.69767 | 1069677.9154 | 81.666 | 197.1591
0.67709 | 1013159.2983 | 57.4718 | 138.7492
0.66239 | 991571.8083 | 71.3164 | 172.1729
0.67482 | 996329.3127 | 66.7289 | 161.0978
0.67793 | 1011929.3418 | 79.5068 | 191.9465
0.69975 | 1058270.573 | 77.2653 | 186.5349
0.68372 | 1033015.9978 | 75.4926 | 182.2554
0.66891 | 994815.1125 | 83.8353 | 202.3963
0.68235 | 1019929.7526 | 61.6689 | 148.882
0.69091 | 1029011.5861 | 85.4419 | 206.2749
0.6807 | 1021820.8248 | 61.9327 | 149.5188
0.66363 | 975373.9025 | 68.2472 | 164.7634
0.6746 | 998316.4011 | 73.9427 | 178.5135
0.66253 | 986111.8477 | 69.1128 | 166.8532
0.67395 | 1001620.0689 | 75.9287 | 183.3081
0.6939 | 1047184.6631 | 78.7587 | 190.1403
0.70213 | 1076158.0825 | 81.6388 | 197.0936
0.69215 | 1050040.4886 | 66.6332 | 160.8668
0.69038 | 1031691.4347 | 65.6359 | 158.4592
0.67482 | 1004091.2672 | 59.3069 | 143.1796
0.68103 | 1007720.2654 | 72.2354 | 174.3917
0.70162 | 1066380.2627 | 64.7632 | 156.3523
0.68665 | 1031605.9109 | 69.7615 | 168.4192
0.66813 | 996194.7537 | 80.9913 | 195.5303
0.6917 | 1073937.4216 | 89.071 | 215.0364
0.68731 | 1046605.6037 | 81.1085 | 195.8134
0.69197 | 1051942.4375 | 45.8029 | 110.5781
0.68075 | 1003919.795 | 73.6861 | 177.8939
0.66789 | 979098.5211 | 64.388 | 155.4463
0.69635 | 1051059.4665 | 70.087 | 169.205
0.69919 | 1059018.5465 | 72.1581 | 174.205
0.70444 | 1084974.9477 | 86.8212 | 209.605
0.68852 | 1035956.285 | 75.1793 | 181.4989
0.67086 | 1009737.9599 | 68.0252 | 164.2274
0.7072 | 1100735.3037 | 79.5188 | 191.9753
0.69387 | 1037705.6284 | 85.7967 | 207.1316
0.68433 | 1025698.1902 | 59.7409 | 144.2273
0.67743 | 1011014.7928 | 79.4937 | 191.9148
0.68599 | 1026828.194 | 65.7448 | 158.7219
0.70214 | 1078899.5833 | 65.4746 | 158.0697
0.69598 | 1059904.5802 | 84.2995 | 203.517
0.66974 | 989236.3068 | 47.2025 | 113.9569
0.68012 | 1016047.6359 | 77.9654 | 188.2251
0.69928 | 1049148.5993 | 61.899 | 149.4375
0.69076 | 1041292.116 | 75.3223 | 181.8441
0.70785 | 1065926.1566 | 75.2426 | 181.6516
0.70832 | 1105256.3916 | 84.7606 | 204.6301
0.69325 | 1035363.559 | 79.5501 | 192.0508
0.69818 | 1061638.8429 | 77.6856 | 187.5496
0.68415 | 1024160.7818 | 69.5109 | 167.8141
0.65654 | 976897.8143 | 65.3537 | 157.7778
0.68544 | 1026097.7107 | 74.7994 | 180.5817
0.68813 | 1023696.4367 | 72.6231 | 175.3276
0.67545 | 1006366.682 | 61.4698 | 148.4011
0.70849 | 1081882.1719 | 87.6028 | 211.4918
0.68641 | 1016309.6101 | 68.751 | 165.9795
0.70367 | 1070845.5553 | 48.6323 | 117.4089
0.67156 | 1010998.0039 | 83.3333 | 201.1843
0.69332 | 1019250.7064 | 73.4549 | 177.3359
0.69557 | 1089427.7375 | 81.7394 | 197.3364
0.69487 | 1028258.2299 | 63.0541 | 152.2261
0.66097 | 957405.2864 | 69.0072 | 166.598
0.68497 | 1017600.3733 | 75.8995 | 183.2376
0.66347 | 962723.3119 | 65.5806 | 158.3255
0.695 | 1063739.9889 | 78.1953 | 188.7802
0.69614 | 1061127.0835 | 68.8919 | 166.3197
0.67775 | 1012108.6315 | 65.5902 | 158.3488
0.7077 | 1077862.3156 | 53.8095 | 129.9077
0.69346 | 1055349.2741 | 68.7808 | 166.0517
0.67794 | 1012824.9568 | 67.5215 | 163.0113
0.68442 | 1012102.1942 | 77.1747 | 186.3162
0.68578 | 1029804.0325 | 75.0319 | 181.143
0.68078 | 1025033.9426 | 66.1325 | 159.6581
0.7049 | 1077473.8189 | 76.3758 | 184.3874
0.67053 | 995414.4012 | 65.2035 | 157.4151
0.69458 | 1040738.0501 | 71.3836 | 172.3352
0.68607 | 1018041.7393 | 82.1392 | 198.3017
0.69363 | 1045334.3524 | 78.6509 | 189.88
0.68594 | 1025153.7666 | 76.3544 | 184.3357
0.69692 | 1049544.8504 | 68.063 | 164.3187
0.69262 | 1049611.3025 | 64.5729 | 155.8927
0.6955 | 1051749.3325 | 78.9202 | 190.5302
0.68466 | 1037128.8808 | 75.6893 | 182.7301
0.68301 | 1011395.4751 | 73.9901 | 178.6278
0.67534 | 1017225.9068 | 87.8879 | 212.1801
0.68666 | 1009948.5813 | 68.2958 | 164.8807
0.68221 | 1034606.8167 | 80.7786 | 195.0168
0.68788 | 1063669.452 | 77.2623 | 186.5278
0.72045 | 1119879.1652 | 70.2523 | 169.6041
0.67924 | 1015755.2869 | 59.6052 | 143.8996
0.69051 | 1032326.0376 | 74.7004 | 180.3426
0.6962 | 1062146.3301 | 68.9515 | 166.4637
0.71727 | 1119454.504 | 79.7362 | 192.5003
0.68979 | 1045741.7595 | 77.4529 | 186.9879
0.66782 | 1020410.0125 | 72.7933 | 175.7386
0.67377 | 989612.7092 | 71.7018 | 173.1034
0.66752 | 960980.1992 | 71.9324 | 173.6601
0.69548 | 1056421.7452 | 72.3651 | 174.7049
0.69349 | 1052454.7684 | 82.4365 | 199.0194
0.6885 | 1049547.3038 | 62.8923 | 151.8355
0.69684 | 1057042.1168 | 89.9739 | 217.2162
0.66283 | 982487.1597 | 56.9313 | 137.4443
0.6908 | 1047502.3409 | 66.6021 | 160.7918
0.66783 | 973528.6353 | 74.4759 | 179.8008
0.70675 | 1077566.8708 | 62.9039 | 151.8634
0.68402 | 1035210.8262 | 68.3998 | 165.1318
0.7032 | 1064981.8338 | 64.5328 | 155.796
0.71103 | 1084744.7883 | 79.6183 | 192.2156
0.70514 | 1057912.511 | 74.9637 | 180.9784
0.70061 | 1057437.3432 | 60.1581 | 145.2345
0.68728 | 1050437.9797 | 72.814 | 175.7885
0.70944 | 1100329.5726 | 68.1227 | 164.4627
0.66423 | 985915.3226 | 86.608 | 209.0902
0.68133 | 1031733.1088 | 62.4755 | 150.8293
0.67789 | 1021194.3704 | 75.2701 | 181.7182
0.69106 | 1033502.0805 | 73.0075 | 176.2558
0.66657 | 1006811.6169 | 75.4705 | 182.2019
0.69445 | 1052368.1061 | 86.9358 | 209.8815
0.6923 | 1032700.5647 | 69.2652 | 167.221
0.70083 | 1071235.3881 | 79.7713 | 192.5849
0.69101 | 1043664.0436 | 71.9536 | 173.7115
0.6897 | 1039108.5192 | 76.9716 | 185.8258
0.69929 | 1068606.56 | 70.9577 | 171.3071
0.68763 | 1026686.519 | 67.4808 | 162.913
0.68571 | 1024357.5169 | 61.6326 | 148.7943
0.6808 | 1017623.3113 | 61.669 | 148.882
0.68377 | 1027072.9099 | 86.1715 | 208.0365
0.68962 | 1043138.503 | 80.5702 | 194.5136
0.67775 | 1011579.7765 | 67.4568 | 162.8552
0.68928 | 1035792.4045 | 73.3583 | 177.1027
0.71552 | 1106039.427 | 78.9726 | 190.6567
0.70929 | 1084090.8058 | 70.9705 | 171.338
0.704 | 1069196.9263 | 72.354 | 174.6781
0.69758 | 1064773.2726 | 79.6593 | 192.3145
0.6848 | 1029033.8723 | 67.9139 | 163.9587
0.66474 | 1006370.4363 | 73.9057 | 178.4242
0.68018 | 1025003.4443 | 66.1888 | 159.794
0.68404 | 1032204.9432 | 50.8317 | 122.7185
0.72069 | 1091494.3622 | 77.3754 | 186.8007
0.72175 | 1115281.2171 | 78.9461 | 190.5927
0.67022 | 994212.4285 | 92.1476 | 222.4639
0.69362 | 1055559.7145 | 66.8414 | 161.3695
0.68861 | 1024645.8036 | 77.6554 | 187.4768
0.68092 | 1026143.1537 | 86.4542 | 208.7188
0.66123 | 969177.8462 | 74.8826 | 180.7826
0.66759 | 981754.4714 | 85.809 | 207.1613
0.6596 | 960638.1596 | 89.129 | 215.1765
0.68003 | 1018802.4139 | 79.7992 | 192.6522
0.66557 | 992779.0112 | 96.6595 | 233.3567
0.70543 | 1076564.1758 | 76.2545 | 184.0947
0.66041 | 968769.9342 | 71.6112 | 172.8847
0.68143 | 1009592.2481 | 83.5177 | 201.6296
0.69519 | 1047152.4171 | 80.2318 | 193.6967
0.67491 | 1010255.0387 | 61.7574 | 149.0955
0.67584 | 995741.4776 | 68.4134 | 165.1645
0.68207 | 1021361.5569 | 67.678 | 163.3891
0.72339 | 1130759.4655 | 84.9093 | 204.9893
0.66766 | 999932.5085 | 67.6882 | 163.4139
0.68792 | 1043685.7298 | 70.4485 | 170.0778
0.68428 | 1019525.7835 | 72.734 | 175.5953
0.70252 | 1072055.211 | 62.9591 | 151.9967
0.6928 | 1048839.1409 | 86.8512 | 209.6775
0.69501 | 1055598.6255 | 74.1454 | 179.0028
0.70045 | 1082346.8058 | 67.8688 | 163.8499
0.71588 | 1105774.2836 | 69.8111 | 168.5388
0.69298 | 1047131.6758 | 80.4863 | 194.3111
0.70596 | 1066331.2003 | 75.4727 | 182.2071
0.68697 | 1031438.7815 | 71.5293 | 172.687
0.689 | 1022827.0004 | 68.9318 | 166.4162
0.6664 | 981782.2147 | 81.5523 | 196.8847
0.73301 | 1152916.2743 | 82.0851 | 198.171
0.69958 | 1054046.7318 | 78.631 | 189.8319
0.67021 | 1000380.9992 | 71.11 | 171.6746
0.68477 | 1028636.2408 | 89.0991 | 215.1042
0.68712 | 1018874.7673 | 75.756 | 182.8912
0.68363 | 1036782.1459 | 72.2514 | 174.4303
0.67826 | 1001991.9374 | 84.7504 | 204.6057
0.6903 | 1031022.2436 | 82.6767 | 199.5992
0.68966 | 1032947.7117 | 78.1657 | 188.7088
0.70098 | 1072159.4131 | 79.8223 | 192.7081
0.70222 | 1063964.4252 | 72.1821 | 174.2631
0.68899 | 1024380.0348 | 72.581 | 175.226
0.69861 | 1047710.083 | 65.8703 | 159.0249
0.70557 | 1091960.7294 | 77.3292 | 186.6893
0.69762 | 1053150.439 | 70.0989 | 169.2337
0.67352 | 1014987.8049 | 69.9755 | 168.9358
0.6834 | 1034144.9872 | 86.8123 | 209.5834
0.69734 | 1059952.2811 | 66.7612 | 161.1759
0.68874 | 1040899.7756 | 80.4999 | 194.3439
0.68045 | 1003507.4784 | 66.6526 | 160.9137
0.70368 | 1070672.6105 | 58.4395 | 141.0855
0.70277 | 1075438.8981 | 58.7752 | 141.8959
0.68655 | 1032902.6644 | 71.6742 | 173.0368
0.69159 | 1058000.4638 | 72.8805 | 175.9492
0.68715 | 1032400.4604 | 69.7649 | 168.4273
0.68674 | 1039898.541 | 77.3627 | 186.77
0.68918 | 1030690.9458 | 73.0077 | 176.2563
0.66431 | 981556.6135 | 60.9603 | 147.1711
0.71126 | 1076632.3802 | 70.2876 | 169.6893
0.67663 | 1019729.7303 | 62.178 | 150.1109
0.68176 | 1026174.5927 | 59.9939 | 144.838
0.68088 | 1018263.7861 | 65.6859 | 158.5798
0.71567 | 1098975.211 | 78.4526 | 189.4013
0.67712 | 1014353.8042 | 76.5338 | 184.7691
0.69403 | 1045794.5906 | 81.8515 | 197.607
0.68828 | 1033398.037 | 78.3566 | 189.1696
0.69049 | 1043455.1686 | 73.7721 | 178.1016
0.69841 | 1069975.1823 | 71.1376 | 171.7413
0.69068 | 1047358.8047 | 51.0819 | 123.3226
0.68456 | 1042038.8067 | 77.7689 | 187.7507
0.68107 | 1022528.505 | 67.8542 | 163.8144
0.67881 | 1032675.8042 | 68.7567 | 165.9933
0.6899 | 1042294.6512 | 88.1535 | 212.8214
0.69381 | 1031694.9774 | 65.3125 | 157.6783
0.68561 | 1039664.9698 | 58.2426 | 140.61
0.67841 | 1026264.139 | 86.7838 | 209.5147
0.67748 | 1014452.4936 | 69.5305 | 167.8615
0.67644 | 1010595.3936 | 71.1749 | 171.8314
0.69404 | 1049865.9256 | 62.4095 | 150.6699
0.71917 | 1131547.7809 | 74.958 | 180.9647
0.69546 | 1044620.5464 | 87.6437 | 211.5907
0.68994 | 1057766.3904 | 78.3269 | 189.0979
0.68802 | 1038473.6769 | 73.6827 | 177.8857
0.65897 | 971828.3346 | 68.4795 | 165.3242
0.69859 | 1046101.8479 | 69.5325 | 167.8662
0.68276 | 1016523.8579 | 74.3923 | 179.5989
0.68072 | 1017377.1394 | 67.5472 | 163.0734
0.68157 | 1028379.9956 | 64.2921 | 155.2149
0.71395 | 1102974.3539 | 78.927 | 190.5465
0.68919 | 1041701.7932 | 82.1032 | 198.2147
0.68231 | 1023213.6294 | 83.011 | 200.4063
0.70207 | 1060311.6173 | 82.0371 | 198.0552
0.71194 | 1097711.8222 | 75.2697 | 181.7171
0.68338 | 1021210.297 | 79.5965 | 192.163
0.67855 | 1011385.0674 | 84.0203 | 202.843
0.69499 | 1039665.5671 | 55.202 | 133.2693
0.69337 | 1058341.8367 | 72.7177 | 175.556
0.67102 | 992104.3319 | 64.9391 | 156.7769
0.67385 | 1013934.9604 | 93.4686 | 225.6531
0.70206 | 1074484.4729 | 86.274 | 208.2838
0.6841 | 1036416.8253 | 80.6968 | 194.8193
0.70061 | 1070790.9896 | 69.0755 | 166.763
0.69217 | 1050500.3902 | 87.0682 | 210.2012
0.66887 | 1002080.9327 | 57.5424 | 138.9195
0.67549 | 995752.957 | 64.7465 | 156.3119
0.69354 | 1059617.0632 | 76.3402 | 184.3016
0.68208 | 1018728.81 | 79.5545 | 192.0615
0.71391 | 1062684.9647 | 82.3973 | 198.9247
0.71326 | 1096808.4136 | 69.4989 | 167.7851
0.69265 | 1051594.7106 | 72.3285 | 174.6164
0.67517 | 1025145.7772 | 72.3357 | 174.6339
0.70616 | 1062530.5619 | 57.7912 | 139.5203
0.69807 | 1077757.6561 | 83.8648 | 202.4675
0.68673 | 1042460.3864 | 70.4869 | 170.1705
0.69908 | 1058911.3681 | 66.53 | 160.6177
0.71507 | 1101176.652 | 73.3644 | 177.1173
0.68717 | 1042161.6546 | 77.4825 | 187.0592
0.67647 | 1025862.8164 | 82.3758 | 198.8728
0.69122 | 1046348.1181 | 66.4883 | 160.5168
0.69192 | 1039108.0714 | 65.9614 | 159.245
0.68481 | 1025404.5906 | 85.9007 | 207.3825
0.66904 | 981839.9147 | 71.6406 | 172.9557
0.69803 | 1049090.449 | 69.7875 | 168.4819
0.68874 | 1025223.274 | 65.38 | 157.8412
0.69314 | 1037642.2384 | 92.9919 | 224.5024
0.70581 | 1073309.6291 | 79.3612 | 191.595
0.67989 | 1014883.9535 | 65.563 | 158.2831
0.68988 | 1046074.1551 | 62.5587 | 151.03
0.70464 | 1058578.6196 | 65.9844 | 159.3005
0.70609 | 1078743.6026 | 69.0487 | 166.6982
0.68886 | 1043471.4942 | 64.5452 | 155.8259
0.68892 | 1048030.5151 | 68.7541 | 165.987
0.68864 | 1024223.2904 | 63.3748 | 153.0002
0.72924 | 1130591.3603 | 71.9901 | 173.7995
0.68679 | 1036460.8269 | 73.6723 | 177.8608
0.69088 | 1044334.4069 | 72.0133 | 173.8555
0.67151 | 995174.1378 | 78.9384 | 190.5742
0.68045 | 1032187.5606 | 68.5584 | 165.5146
0.71056 | 1110311.0541 | 86.4962 | 208.8203
0.72888 | 1131543.9151 | 85.9352 | 207.4659
0.70225 | 1055850.0304 | 86.2554 | 208.2391
0.6922 | 1045984.4066 | 66.371 | 160.2338
0.67753 | 1016518.6993 | 67.8079 | 163.7027
0.68421 | 1037111.9536 | 75.738 | 182.8476
0.6817 | 1017201.1728 | 63.5818 | 153.5001
0.68986 | 1034562.7001 | 61.872 | 149.3721
0.70585 | 1085239.3301 | 63.4329 | 153.1406
0.71155 | 1093615.4573 | 76.796 | 185.4019
0.72239 | 1124115.7486 | 82.885 | 200.102
0.69836 | 1067563.2585 | 69.0891 | 166.7959
0.69322 | 1047816.9681 | 85.2054 | 205.704
0.69073 | 1036954.0908 | 79.9631 | 193.0479
0.69931 | 1058875.3811 | 62.4704 | 150.817
0.69306 | 1043296.8047 | 71.7154 | 173.1362
0.68842 | 1020261.1464 | 59.0715 | 142.6112
0.70557 | 1099195.8855 | 86.2361 | 208.1923
0.69871 | 1053013.8557 | 78.6519 | 189.8825
0.68795 | 1043180.4941 | 77.4695 | 187.0279
0.68292 | 1016281.2718 | 76.0278 | 183.5473
0.69789 | 1057559.6139 | 68.4879 | 165.3443
0.68528 | 1029608.5533 | 84.6944 | 204.4703
0.70413 | 1071577.6547 | 85.7873 | 207.1087
0.6893 | 1053469.3434 | 79.7607 | 192.5595
0.68298 | 1055064.6764 | 73.5949 | 177.6739
0.69828 | 1045477.5838 | 59.6767 | 144.0723
0.68195 | 1030754.057 | 58.0534 | 140.1533
0.6826 | 1024334.3556 | 63.9433 | 154.3728
0.69588 | 1052965.9401 | 75.6065 | 182.5303
0.72051 | 1127401.4781 | 83.3442 | 201.2108
0.69183 | 1056642.3174 | 87.6197 | 211.5327
0.69738 | 1058995.108 | 66.6974 | 161.0217
0.66887 | 1000798.0824 | 67.6235 | 163.2576
0.68327 | 1006691.097 | 58.7247 | 141.7739
0.7058 | 1087849.669 | 92.8838 | 224.2414
0.69936 | 1065248.7194 | 79.4919 | 191.9105
0.68614 | 1033747.2918 | 73.5487 | 177.5624
0.70173 | 1080634.0069 | 76.9461 | 185.7643
0.69544 | 1060372.9158 | 76.1859 | 183.9291
0.68726 | 1052541.8625 | 69.2173 | 167.1054
0.66403 | 996050.0587 | 65.8835 | 159.0569
0.66741 | 983759.3433 | 66.6185 | 160.8312
0.70047 | 1056242.0013 | 68.1506 | 164.5302
0.68597 | 1031119.8967 | 73.38 | 177.1551
0.69268 | 1039725.5524 | 69.6988 | 168.2679
0.68579 | 1026817.6332 | 78.1409 | 188.6489
0.69606 | 1057775.6633 | 75.9323 | 183.3168
0.69485 | 1039424.9169 | 73.0256 | 176.2994
0.69932 | 1069652.8655 | 87.7486 | 211.8438
0.69589 | 1037060.4121 | 69.1997 | 167.0628
0.68019 | 1016044.4971 | 74.6378 | 180.1916
0.70666 | 1076385.8687 | 39.7446 | 95.952
0.69364 | 1028989.3664 | 77.4194 | 186.907
0.68116 | 1028202.2609 | 73.6891 | 177.9012
0.67425 | 1012491.2062 | 65.9561 | 159.232
0.69779 | 1070686.219 | 49.0351 | 118.3812
0.68829 | 1028742.1966 | 76.3604 | 184.3504
0.68205 | 1025221.5148 | 76.7954 | 185.4004
0.66531 | 990704.4049 | 76.076 | 183.6636
0.69849 | 1065663.5461 | 72.6836 | 175.4736
0.68445 | 1029483.9124 | 75.1541 | 181.438
0.68908 | 1055575.3184 | 76.5408 | 184.7858
0.69379 | 1066704.0081 | 69.4883 | 167.7597
0.71104 | 1090376.46 | 74.8611 | 180.7308
0.68206 | 1011312.9727 | 72.3829 | 174.7477
0.69857 | 1067764.4554 | 81.4153 | 196.5538
0.71112 | 1105579.8608 | 74.0675 | 178.8147
0.66527 | 993558.8795 | 71.1494 | 171.7699
0.68641 | 1027927.3381 | 88.826 | 214.4448
0.69402 | 1053830.3007 | 84.0534 | 202.9229
0.689 | 1034844.0164 | 67.8813 | 163.88
0.66864 | 985769.5279 | 81.6087 | 197.0209
0.66969 | 991483.5281 | 63.0856 | 152.3021
0.68777 | 1040324.4957 | 82.1646 | 198.363
0.68313 | 1022448.2139 | 65.7388 | 158.7075
0.68 | 1022480.3211 | 71.6347 | 172.9414
0.66684 | 984239.931 | 87.546 | 211.3549
0.70267 | 1066242.5741 | 62.6043 | 151.1402
0.66249 | 956287.9234 | 78.4088 | 189.2955
0.70649 | 1082876.5362 | 74.8245 | 180.6423
0.69607 | 1064208.0153 | 63.0774 | 152.2823
0.68108 | 1013726.2548 | 85.3876 | 206.1438
0.67207 | 986360.9096 | 71.647 | 172.9711
0.69566 | 1070560.2592 | 83.5932 | 201.8119
0.70723 | 1087944.87 | 83.5616 | 201.7354
0.67528 | 980626.5847 | 53.5872 | 129.3709
0.69181 | 1054656.488 | 68.9002 | 166.3398
0.69336 | 1049975.0502 | 70.3165 | 169.7591
0.68625 | 1034565.1157 | 73.5949 | 177.6739
0.69891 | 1070918.4553 | 63.2169 | 152.6191
0.6894 | 1028558.5481 | 71.0783 | 171.5982
0.70029 | 1053102.8572 | 80.0998 | 193.3779
0.6854 | 1034117.5906 | 66.269 | 159.9875
0.68526 | 1034326.1968 | 77.1144 | 186.1706
0.68576 | 1037605.5496 | 72.0841 | 174.0265
0.68956 | 1046612.2098 | 60.4963 | 146.0509
0.71197 | 1082598.7305 | 66.786 | 161.2357
0.68662 | 1039044.5668 | 64.481 | 155.6709
0.69183 | 1051701.1269 | 75.9577 | 183.3782
0.69116 | 1024396.2669 | 61.1489 | 147.6265
0.66142 | 959178.6073 | 77.2435 | 186.4823
0.69979 | 1076941.2487 | 81.9715 | 197.8966
0.69163 | 1037688.9045 | 61.901 | 149.4421
0.69883 | 1054616.9186 | 83.5363 | 201.6746
0.69877 | 1070778.1935 | 87.9102 | 212.2339
0.69977 | 1055143.2566 | 60.0998 | 145.0939
0.70089 | 1066858.9563 | 87.324 | 210.8187
0.70117 | 1066068.7347 | 68.7164 | 165.896
0.68624 | 1027013.4539 | 76.4593 | 184.5892
0.70385 | 1069024.9137 | 66.7099 | 161.0519
0.66467 | 998045.8559 | 79.9804 | 193.0898
0.68707 | 1040563.7839 | 82.2664 | 198.6085
0.70469 | 1072191.1753 | 90.2472 | 217.876
0.69009 | 1018796.71 | 90.3485 | 218.1206
0.6806 | 1012248.0811 | 61.3617 | 148.1403
0.70386 | 1079760.3236 | 78.3744 | 189.2125
0.67074 | 990310.3074 | 66.7285 | 161.0969
0.68441 | 1031304.8601 | 64.2074 | 155.0103
0.68919 | 1029453.8691 | 63.6482 | 153.6603
0.68612 | 1033005.8975 | 84.8706 | 204.8958
0.72641 | 1142752.483 | 80.3711 | 194.033
0.7107 | 1094636.3603 | 61.3904 | 148.2096
0.66953 | 987972.203 | 87.5251 | 211.3042
0.70291 | 1079058.9469 | 81.7677 | 197.4048
0.67347 | 989732.3495 | 68.24 | 164.746
0.72004 | 1129516.1938 | 74.7336 | 180.4228
0.67686 | 1011823.8923 | 78.0156 | 188.3463
0.71361 | 1106864.5957 | 79.2431 | 191.3098
0.66956 | 1010108.632 | 71.8709 | 173.5116
0.69324 | 1050634.1505 | 73.9846 | 178.6146
0.71765 | 1121094.0501 | 60.0004 | 144.8537
0.67557 | 1008812.0671 | 72.8973 | 175.9897
0.6755 | 997687.9107 | 68.5817 | 165.571
0.6933 | 1047532.2686 | 68.2872 | 164.86
0.67149 | 1005331.5058 | 67.6649 | 163.3575
0.68458 | 1029943.5627 | 76.863 | 185.5637
0.68303 | 1020392.8736 | 71.1717 | 171.8236
0.67222 | 992665.4972 | 70.3509 | 169.842
0.68543 | 1023915.3751 | 63.6527 | 153.6713
0.67692 | 1007480.8664 | 85.3834 | 206.1337
0.66182 | 970183.3771 | 64.3872 | 155.4444
0.67851 | 1006492.8685 | 70.6692 | 170.6105
0.69417 | 1051028.3406 | 66.5015 | 160.5489
0.68694 | 1026233.3136 | 76.3238 | 184.262
0.69751 | 1060726.6027 | 72.4485 | 174.9061
0.6843 | 1026384.051 | 71.4563 | 172.5108
0.696 | 1043233.0974 | 81.4195 | 196.5642
0.68046 | 1025083.1263 | 78.6199 | 189.8052
0.67604 | 1007609.9276 | 72.7605 | 175.6593
0.69914 | 1060820.1009 | 72.2863 | 174.5145
0.68765 | 1034899.6597 | 86.339 | 208.4408
0.70647 | 1065927.9375 | 86.5585 | 208.9708
0.69834 | 1059235.1712 | 77.4923 | 187.0829
0.68658 | 1028539.2922 | 78.181 | 188.7456
0.69295 | 1049454.6127 | 83.8328 | 202.3902
0.66866 | 986535.5963 | 86.0886 | 207.8363
0.7137 | 1094116.2841 | 63.248 | 152.6941
0.66108 | 982401.9821 | 93.6592 | 226.1132
0.67482 | 1009703.5835 | 76.4429 | 184.5496
0.68052 | 1015089.7688 | 74.5316 | 179.9352
0.67525 | 1012669.0351 | 71.4469 | 172.488
0.6797 | 1009721.1415 | 68.9796 | 166.5314
0.69611 | 1050103.6449 | 80.3515 | 193.9856
0.67513 | 1006776.3496 | 76.1864 | 183.9303
0.68269 | 1028148.9444 | 66.1476 | 159.6945
0.70721 | 1093282.9491 | 77.2428 | 186.4806
0.69862 | 1058262.1364 | 64.5461 | 155.828
0.70129 | 1070195.107 | 74.4969 | 179.8515
0.6946 | 1039128.1693 | 74.7504 | 180.4634
0.71578 | 1104301.0789 | 77.0872 | 186.105
0.67649 | 1014018.5709 | 63.8703 | 154.1965
0.67645 | 1005418.0106 | 61.3211 | 148.0422
0.69502 | 1052493.678 | 68.4415 | 165.2324
0.67417 | 1016566.5787 | 75.3434 | 181.895
0.69696 | 1055876.2652 | 85.4001 | 206.1742
0.69006 | 1038576.1032 | 78.5464 | 189.6279
0.69826 | 1060521.5492 | 79.2638 | 191.3598
0.69881 | 1070264.5112 | 71.1159 | 171.689
0.67024 | 989335.0231 | 62.8973 | 151.8476
0.6672 | 983482.9039 | 81.5958 | 196.9897
0.69761 | 1066811.1898 | 74.6534 | 180.2293
0.701 | 1074814.3896 | 63.2584 | 152.7194
0.68733 | 1030466.6579 | 73.1878 | 176.6911
0.69937 | 1056049.0705 | 77.1445 | 186.2432
0.6761 | 1004310.236 | 68.9757 | 166.5221
0.68544 | 1029471.8237 | 53.0627 | 128.1046
0.6918 | 1048616.2993 | 75.7382 | 182.8481
0.68225 | 1022918.0994 | 57.5869 | 139.027
0.69545 | 1044180.5315 | 71.5294 | 172.6872
0.6991 | 1060867.4831 | 80.8283 | 195.1367
0.69441 | 1035353.7301 | 95.3031 | 230.0821
0.68026 | 1019683.3494 | 69.6032 | 168.0369
0.70053 | 1058516.0601 | 55.7589 | 134.614
0.70318 | 1061190.8415 | 73.672 | 177.8599
0.69501 | 1051809.3166 | 68.3979 | 165.127
0.67573 | 1009859.8306 | 82.7258 | 199.7178
0.67224 | 997874.6724 | 75.8375 | 183.0879
0.69764 | 1077281.7866 | 73.566 | 177.604
0.69858 | 1047566.1721 | 78.2593 | 188.9347
0.69113 | 1028642.355 | 69.1105 | 166.8475
0.69227 | 1042083.7184 | 73.4858 | 177.4105
0.67983 | 1024557.7008 | 84.955 | 205.0995
0.69542 | 1050737.1402 | 66.3639 | 160.2167
0.68136 | 1012239.9396 | 68.9837 | 166.5414
0.68733 | 1039454.9415 | 83.7131 | 202.1013
0.70173 | 1067577.2579 | 72.4776 | 174.9763
0.68462 | 1040189.3418 | 70.5501 | 170.323
0.70728 | 1083124.1689 | 58.8183 | 142
0.66845 | 996395.7788 | 71.0135 | 171.4417
0.69381 | 1049834.323 | 73.8679 | 178.3328
0.68203 | 1025585.3536 | 71.6388 | 172.9513
0.66768 | 995968.8568 | 67.4171 | 162.7594
0.71394 | 1085589.1613 | 70.606 | 170.4579
0.70265 | 1054734.876 | 76.689 | 185.1436
0.67575 | 991773.984 | 81.2419 | 196.1353
0.6964 | 1048613.1935 | 76.9286 | 185.722
0.70886 | 1100543.4241 | 79.3185 | 191.4918
0.69246 | 1065669.004 | 81.9722 | 197.8984
0.6704 | 992296.6551 | 65.0115 | 156.9515
0.72064 | 1115352.2622 | 69.9505 | 168.8755
0.6574 | 964361.4723 | 58.4694 | 141.1577
0.69381 | 1044636.8762 | 69.9139 | 168.787
0.67645 | 1010663.6996 | 95.5013 | 230.5604
0.68263 | 1033315.3342 | 70.2884 | 169.6913
0.67359 | 1004216.734 | 80.9818 | 195.5073
0.71512 | 1094724.342 | 63.1577 | 152.4762
0.69852 | 1062032.2256 | 72.4331 | 174.8691
0.66356 | 994167.3799 | 69.5689 | 167.9542
0.68277 | 1028101.1615 | 66.1371 | 159.6692
0.7124 | 1093439.0117 | 71.2699 | 172.0607
0.69932 | 1058214.8852 | 77.3084 | 186.639
0.68295 | 1019119.8112 | 72.8471 | 175.8684
0.71428 | 1111084.3044 | 84.4958 | 203.9909
0.67235 | 999770.5525 | 82.3112 | 198.7168
0.69463 | 1043932.9589 | 73.3652 | 177.1192
0.69849 | 1072850.832 | 91.1036 | 219.9435
0.69457 | 1072115.3246 | 63.6831 | 153.7446
0.67567 | 1008470.036 | 67.4561 | 162.8534
0.64948 | 952532.8378 | 69.8706 | 168.6825
0.69869 | 1052859.1571 | 75.2948 | 181.7778
0.69598 | 1055006.7317 | 73.9996 | 178.6508
0.70363 | 1064271.6105 | 75.5763 | 182.4573
0.68399 | 1034373.0315 | 83.7106 | 202.0954
0.66407 | 977508.0771 | 66.5274 | 160.6114
0.68653 | 1014855.3259 | 78.7382 | 190.0909
0.68064 | 1008227.7837 | 74.0145 | 178.6868
0.68562 | 1029613.0053 | 73.31 | 176.9859
0.69733 | 1045026.2458 | 67.731 | 163.5172
0.70312 | 1064666.6116 | 70.3104 | 169.7443
0.69475 | 1064401.7841 | 72.9888 | 176.2105
0.69659 | 1051012.6449 | 78.7413 | 190.0983
0.67983 | 1018665.6661 | 73.1485 | 176.596
0.65922 | 974010.4188 | 62.9579 | 151.9937
0.71575 | 1114490.6517 | 81.2625 | 196.185
0.67566 | 1020737.2056 | 77.1826 | 186.3353
0.68944 | 1050113.1058 | 87.3797 | 210.9533
0.6939 | 1053561.4534 | 74.1204 | 178.9426
0.70859 | 1074283.1275 | 77.4562 | 186.9957
0.6777 | 1010355.5851 | 72.7802 | 175.7069
0.70553 | 1088288.0707 | 78.8059 | 190.2543
0.73144 | 1138325.0189 | 74.1422 | 178.9951
0.70987 | 1103322.1893 | 66.2279 | 159.8882
0.69561 | 1055815.2977 | 91.4701 | 220.8284
0.68326 | 1031782.1522 | 78.8089 | 190.2615
0.67789 | 1018281.2058 | 75.0559 | 181.201
0.69106 | 1051856.0652 | 74.3039 | 179.3854
0.7001 | 1062470.9012 | 77.5326 | 187.1804
0.70672 | 1100016.4325 | 66.3239 | 160.1201
0.69974 | 1079404.253 | 71.6282 | 172.9258
0.68173 | 1022200.0313 | 63.8912 | 154.2469
0.68761 | 1024894.6054 | 63.5118 | 153.331
0.66866 | 984775.1135 | 54.3466 | 131.2042
0.67724 | 1004153.6593 | 77.381 | 186.8142
0.70433 | 1096948.9088 | 70.3891 | 169.9342
0.69692 | 1062861.5516 | 81.2334 | 196.1148
0.73357 | 1151645.4881 | 86.902 | 209.7999
0.68097 | 1019766.015 | 65.7746 | 158.794
0.69171 | 1043227.2381 | 77.0523 | 186.0208
0.68126 | 1007120.3842 | 70.4276 | 170.0272
0.68903 | 1052335.3964 | 86.0931 | 207.8471
0.71108 | 1093236.9594 | 75.0534 | 181.195
0.70447 | 1080453.1651 | 71.2168 | 171.9326
0.69193 | 1036778.1066 | 67.72 | 163.4905
0.68922 | 1039650.7359 | 83.4991 | 201.5847
0.69949 | 1060785.6818 | 66.6082 | 160.8063
0.70537 | 1057335.0422 | 79.8432 | 192.7586
0.67364 | 1011354.8029 | 65.9414 | 159.1967
0.71113 | 1106671.5762 | 79.5059 | 191.9443
0.68723 | 1023614.1397 | 70.9594 | 171.3112
0.69466 | 1045869.6091 | 62.1161 | 149.9614
0.6795 | 1006261.0723 | 63.1014 | 152.3403
0.65762 | 971396.5597 | 66.7605 | 161.174
0.70238 | 1088439.6367 | 80.8738 | 195.2466
0.68004 | 1014733.1168 | 77.8734 | 188.0031
0.67782 | 1019259.9503 | 66.0471 | 159.4518
0.69805 | 1069387.194 | 74.7866 | 180.5508
0.69352 | 1057687.5688 | 62.9218 | 151.9066
0.69449 | 1042541.6396 | 76.6932 | 185.1538
0.68205 | 1017570.3307 | 79.4674 | 191.8513
0.68645 | 1035974.5931 | 83.5401 | 201.6837
0.67071 | 989119.9199 | 81.6332 | 197.08
0.69702 | 1074752.0776 | 71.5984 | 172.8537
0.69123 | 1038833.0679 | 67.9232 | 163.9811
0.68566 | 1026868.3049 | 82.901 | 200.1407
0.68628 | 1056375.0146 | 65.5894 | 158.3469
0.70781 | 1071701.0873 | 60.3496 | 145.6969
0.67388 | 997399.1198 | 68.7269 | 165.9214
0.68353 | 1030835.7552 | 78.7059 | 190.0127
0.71132 | 1106474.1918 | 79.4706 | 191.859
0.69227 | 1065096.92 | 92.7191 | 223.8436
0.69651 | 1072669.0511 | 71.5114 | 172.6438
0.6671 | 979804.5868 | 66.3811 | 160.2581
0.69329 | 1046206.5225 | 70.071 | 169.1662
0.66835 | 988933.4303 | 78.8133 | 190.272
0.68274 | 1021499.5434 | 65.7149 | 158.6498
0.68169 | 1025291.2078 | 86.2844 | 208.3091
0.69312 | 1066036.851 | 75.0821 | 181.2643
0.6897 | 1050126.7795 | 81.0996 | 195.7917
0.70563 | 1087245.4981 | 82.0707 | 198.1362
0.6743 | 1004387.389 | 72.7144 | 175.5481
0.67682 | 1001626.3839 | 77.2332 | 186.4575
0.68741 | 1020125.8227 | 68.9087 | 166.3602
0.68989 | 1033139.6767 | 79.7657 | 192.5715
0.68685 | 1046265.7743 | 81.1376 | 195.8835
0.66937 | 978197.261 | 73.0187 | 176.2828
0.66876 | 1001084.792 | 65.7486 | 158.731
0.68145 | 1032460.2527 | 88.4044 | 213.427
0.69702 | 1053625.9485 | 86.7191 | 209.3583
0.69488 | 1060164.979 | 76.2324 | 184.0412
0.69312 | 1041912.6966 | 89.1625 | 215.2572
0.69914 | 1063966.7981 | 74.2917 | 179.3561
0.69981 | 1072223.2272 | 71.1451 | 171.7595
0.67059 | 1001732.7006 | 71.1656 | 171.8089
0.64917 | 944256.3846 | 57.9758 | 139.966
0.66597 | 983121.6334 | 72.0652 | 173.9807
0.69331 | 1050424.5176 | 77.4048 | 186.8717
0.68777 | 1034213.702 | 62.0393 | 149.7762
0.70785 | 1076320.1163 | 68.7225 | 165.9107
0.67182 | 1003477.2241 | 77.9526 | 188.1943
0.66641 | 978981.669 | 78.2484 | 188.9084
0.69983 | 1054506.1426 | 90.8053 | 219.2233
0.67761 | 1001845.2979 | 66.822 | 161.3226
0.6893 | 1042210.0914 | 82.1306 | 198.2809
0.70969 | 1098583.3727 | 67.9904 | 164.1433
0.67502 | 1006267.173 | 63.5204 | 153.3517
0.69105 | 1043290.2899 | 78.7897 | 190.2151
0.67032 | 992607.2211 | 70.5464 | 170.314
0.69196 | 1039289.6453 | 80.0014 | 193.1404
0.72673 | 1146092.8958 | 71.2761 | 172.0756
0.69224 | 1038205.8421 | 65.8728 | 159.0311
0.65971 | 970099.9258 | 71.7854 | 173.3054
0.68977 | 1050178.002 | 64.8748 | 156.6216
0.68366 | 1028414.2774 | 67.987 | 164.1352
0.69633 | 1063657.8022 | 67.0921 | 161.9746
0.6784 | 1006461.3127 | 68.8891 | 166.313
0.71397 | 1111367.3553 | 81.237 | 196.1234
0.6787 | 1020675.6914 | 64.081 | 154.7052
0.68641 | 1041167.1897 | 75.7903 | 182.9739
0.69946 | 1062940.4181 | 75.932 | 183.3161
0.70122 | 1080669.9938 | 75.2338 | 181.6304
0.68975 | 1042559.8185 | 73.7876 | 178.139
0.68438 | 1024683.0178 | 76.4928 | 184.67
0.6738 | 1016066.4797 | 81.9125 | 197.7543
0.68095 | 1002376.6293 | 62.7239 | 151.4289
0.68778 | 1036890.7626 | 80.3014 | 193.8647
0.67411 | 996423.955 | 58.9155 | 142.2345
0.70204 | 1079582.796 | 75.0697 | 181.2344
0.69177 | 1039416.3119 | 73.3506 | 177.084
0.69738 | 1045853.6435 | 65.8697 | 159.0236
0.68822 | 1033627.5899 | 61.9635 | 149.5931
0.68291 | 1021337.5959 | 61.7031 | 148.9645
0.69053 | 1037932.8816 | 57.7315 | 139.3762
0.6965 | 1046380.7897 | 75.0661 | 181.2255
0.67559 | 1004734.3554 | 63.37 | 152.9887
0.72274 | 1107998.3608 | 81.6576 | 197.1389
0.67552 | 1004550.4069 | 82.3162 | 198.729
0.69003 | 1041021.6235 | 68.1846 | 164.6123
0.67874 | 1027259.6136 | 66.4755 | 160.4862
0.6827 | 1025645.2247 | 71.5719 | 172.79
0.70216 | 1056515.5511 | 81.7329 | 197.3207
0.68161 | 1025255.0609 | 71.6594 | 173.0011
0.66988 | 1006770.2614 | 62.0063 | 149.6965
0.7079 | 1080090.0044 | 90.4992 | 218.4843
0.68047 | 1017214.5367 | 72.9754 | 176.1781
0.6704 | 990568.0899 | 68.6097 | 165.6386
0.6622 | 972759.0002 | 81.3145 | 196.3105
0.67394 | 979129.8077 | 81.8975 | 197.718
0.68292 | 1023512.2293 | 73.1637 | 176.6329
0.71966 | 1116180.0639 | 72.0182 | 173.8674
0.70875 | 1083172.3793 | 74.6068 | 180.1168
0.68584 | 1029398.8223 | 65.9241 | 159.1549
0.68435 | 1028334.9179 | 73.3551 | 177.0948
0.65694 | 961745.3747 | 66.1654 | 159.7374
0.687 | 1037470.7938 | 75.4564 | 182.1677
0.69457 | 1037221.1225 | 77.9344 | 188.1502
0.67659 | 1019621.2947 | 87.9326 | 212.2881
0.66649 | 988984.031 | 80.6033 | 194.5935
0.69277 | 1050456.704 | 62.559 | 151.0307
0.66585 | 979135.2366 | 64.588 | 155.9292
0.7003 | 1088809.6429 | 75.8996 | 183.2378
0.70863 | 1072435.4139 | 68.3989 | 165.1295
0.68416 | 1030255.735 | 55.5682 | 134.1534
0.68094 | 1026632.4845 | 78.6825 | 189.9565
0.69242 | 1045626.4468 | 70.1202 | 169.2852
0.69741 | 1060935.9847 | 60.6343 | 146.3841
0.67971 | 1023812.3442 | 73.5312 | 177.52
0.70597 | 1068956.4302 | 89.8119 | 216.825
0.67895 | 1008517.8901 | 75.8883 | 183.2105
0.70896 | 1070518.1331 | 64.3415 | 155.3342
0.66961 | 992907.592 | 72.3865 | 174.7565
0.68096 | 1022514.3835 | 83.1197 | 200.6688
0.68685 | 1032863.1946 | 59.9377 | 144.7024
0.69348 | 1040303.1491 | 74.9557 | 180.9591
0.70836 | 1110051.6438 | 81.7834 | 197.4426
0.69044 | 1037891.4498 | 76.7989 | 185.4089
0.69804 | 1071484.4388 | 65.4066 | 157.9056
0.69014 | 1033094.7728 | 82.8617 | 200.0459
0.69312 | 1051808.1595 | 70.4636 | 170.1142
0.69287 | 1056051.3963 | 60.7508 | 146.6655
0.66184 | 967608.124 | 60.3181 | 145.6208
0.69608 | 1048217.4835 | 71.7412 | 173.1986
0.69743 | 1040710.3377 | 57.3132 | 138.3663
0.67823 | 999302.6697 | 76.8945 | 185.6397
0.70086 | 1069146.9599 | 82.8633 | 200.0498
0.6731 | 991163.7547 | 54.3381 | 131.1838
0.69153 | 1047635.3789 | 80.6932 | 194.8106
0.69953 | 1063781.515 | 90.1992 | 217.7602
0.71385 | 1096076.8595 | 58.9734 | 142.3744
0.67054 | 1000154.5816 | 77.8419 | 187.9269
0.69884 | 1053358.829 | 97.675 | 235.8082
0.69685 | 1044041.1678 | 79.8025 | 192.6603
0.69365 | 1046188.164 | 73.6275 | 177.7524
0.69641 | 1055709.9175 | 83.9136 | 202.5855
0.68518 | 1019125.6439 | 75.99 | 183.4562
0.67341 | 1010137.9469 | 67.7948 | 163.671
0.6589 | 966847.9166 | 64.4498 | 155.5955
0.66682 | 992418.8441 | 72.6538 | 175.4017
0.6905 | 1030615.2549 | 69.3061 | 167.3198
0.67669 | 1010068.4879 | 69.3101 | 167.3293
0.66323 | 981415.8779 | 64.9939 | 156.9092
0.67674 | 1002619.3776 | 80.4766 | 194.2877
0.69172 | 1026962.7122 | 75.8493 | 183.1165
0.67868 | 1024158.6317 | 70.0716 | 169.1679
0.67154 | 991892.2979 | 66.7554 | 161.1619
0.71712 | 1104310.2979 | 65.6173 | 158.4141
0.68969 | 1044102.4505 | 64.5071 | 155.734
0.71899 | 1113529.6635 | 77.3848 | 186.8235
0.70026 | 1081091.6208 | 63.6947 | 153.7725
0.72017 | 1128109.2567 | 89.8645 | 216.9521
0.69776 | 1040701.0852 | 67.4088 | 162.7392
0.66306 | 986522.6341 | 57.145 | 137.9603
0.67225 | 997731.5934 | 80.6195 | 194.6327
0.69901 | 1064264.1178 | 70.6979 | 170.6799
0.68367 | 1020046.0203 | 69.2809 | 167.2588
0.70043 | 1072073.8532 | 65.2197 | 157.4542
0.68444 | 1035861.4712 | 77.3208 | 186.6689
0.66647 | 990460.1945 | 69.0069 | 166.5973
0.70607 | 1081403.8347 | 81.5795 | 196.9503
0.68211 | 1028915.6004 | 75.3701 | 181.9595
0.67802 | 1010594.5477 | 80.9636 | 195.4634
0.69378 | 1060246.5371 | 76.3592 | 184.3475
0.68371 | 1027910.8427 | 78.4618 | 189.4234
0.69371 | 1070768.5381 | 78.8581 | 190.3802
0.68652 | 1031435.3311 | 71.576 | 172.7997
0.68664 | 1031320.0243 | 76.735 | 185.2546
0.70713 | 1087207.2554 | 78.55 | 189.6366
0.71003 | 1092654.4071 | 54.3695 | 131.2595
0.71863 | 1091821.0091 | 77.6208 | 187.3932
0.6843 | 1016354.7273 | 67.1267 | 162.0582
0.68324 | 1025674.9926 | 73.2845 | 176.9243
0.68945 | 1033617.923 | 63.8811 | 154.2227
0.69343 | 1043248.1539 | 55.5175 | 134.031
0.678 | 1006603.9335 | 54.7858 | 132.2647
0.69113 | 1050972.1995 | 64.7961 | 156.4316
0.70359 | 1057658.6205 | 55.4715 | 133.9201
0.70171 | 1080910.6085 | 65.6372 | 158.4622
0.68778 | 1038187.5007 | 75.4233 | 182.0879
0.68066 | 1034554.2735 | 72.4101 | 174.8134
0.71972 | 1115179.8685 | 91.9048 | 221.8777
0.68938 | 1039228.5454 | 80.2953 | 193.8499
0.6764 | 1014778.3766 | 63.9273 | 154.3342
0.69755 | 1043597.7143 | 71.2751 | 172.0734
0.6556 | 956153.4126 | 69.4811 | 167.7421
0.685 | 1040484.7981 | 70.5313 | 170.2777
0.69162 | 1057520.5676 | 74.6439 | 180.2063
0.6889 | 1055828.515 | 78.301 | 189.0354
0.68529 | 1022830.0988 | 76.253 | 184.0911
0.65834 | 968202.7392 | 68.2896 | 164.8657
0.70231 | 1058202.9842 | 76.9789 | 185.8434
0.67626 | 1019657.4866 | 68.1458 | 164.5186
0.68344 | 1027707.0339 | 80.0952 | 193.3669
0.66669 | 988651.313 | 86.5277 | 208.8964
0.70075 | 1062006.2127 | 68.4955 | 165.3628
0.68064 | 1022329.6228 | 80.0044 | 193.1477
0.67371 | 1003860.3089 | 79.3675 | 191.6101
0.67935 | 1017681.4097 | 73.1291 | 176.5492
0.67436 | 1008976.5012 | 61.7291 | 149.0272
0.69389 | 1048925.8437 | 80.5175 | 194.3865
0.68492 | 1029342.5602 | 71.597 | 172.8504
0.67105 | 995867.8248 | 65.182 | 157.3634
0.71334 | 1088191.0023 | 76.2232 | 184.0192
0.69106 | 1050585.5789 | 72.4254 | 174.8503
0.6981 | 1067718.5772 | 63.1015 | 152.3406
0.70009 | 1067931.9771 | 90.4789 | 218.4354
0.68717 | 1023062.1116 | 67.2878 | 162.4471
0.67976 | 1030648.137 | 74.0161 | 178.6906
0.6923 | 1047462.0166 | 77.2996 | 186.6177
0.68596 | 1031732.844 | 64.089 | 154.7245
0.68956 | 1039354.9257 | 78.6622 | 189.9073
0.71068 | 1097858.8603 | 78.9227 | 190.5363
0.70385 | 1079624.0676 | 74.8802 | 180.7769
0.6789 | 1017098.6371 | 73.1362 | 176.5664
0.7013 | 1072539.964 | 87.7371 | 211.8161
0.6944 | 1046321.5758 | 83.4084 | 201.3658
0.66393 | 968449.5314 | 62.1539 | 150.0527
0.67865 | 1026993.3265 | 70.2685 | 169.6432
0.68613 | 1025479.5459 | 72.9612 | 176.144
0.70146 | 1056180.6098 | 84.8765 | 204.9101
0.68808 | 1037396.8886 | 80.3059 | 193.8755
0.68926 | 1052801.8271 | 72.7685 | 175.6787
0.69463 | 1037694.2257 | 82.6932 | 199.6391
0.67202 | 985829.7023 | 68.1178 | 164.4509
0.68026 | 1023047.2167 | 67.9698 | 164.0937
0.68755 | 1044161.2331 | 75.8821 | 183.1956
0.68897 | 1035372.5565 | 69.7801 | 168.4641
0.72382 | 1109700.6893 | 76.6036 | 184.9374
0.64971 | 948756.6958 | 69.5684 | 167.9529
0.66488 | 1002101.736 | 80.2803 | 193.8138
0.69441 | 1051154.8239 | 86.8404 | 209.6513
0.71318 | 1097256.1093 | 69.5349 | 167.8721
0.69446 | 1048849.7474 | 73.7296 | 177.999
0.67246 | 1004422.2901 | 77.9517 | 188.192
0.67957 | 1004084.2736 | 62.9587 | 151.9958
0.68757 | 1038618.3387 | 80.4601 | 194.248
0.66304 | 975782.1862 | 73.5342 | 177.5273
0.68817 | 1047270.3665 | 64.8453 | 156.5504
0.69122 | 1049022.4037 | 82.4376 | 199.0219
0.68426 | 1022180.7199 | 71.168 | 171.8148
0.69887 | 1062982.1932 | 66.8053 | 161.2822
0.70007 | 1051721.2923 | 65.9846 | 159.301
0.67817 | 1016810.0809 | 79.8624 | 192.805
0.70259 | 1064456.3431 | 85.4036 | 206.1826
0.67811 | 1027773.7622 | 70.3779 | 169.9073
0.67234 | 1008181.7339 | 61.8654 | 149.3562
0.69284 | 1052107.9663 | 62.651 | 151.2528
0.69364 | 1043118.6919 | 76.8297 | 185.4833
0.68945 | 1031315.5845 | 69.4591 | 167.6892
0.7015 | 1065569.5727 | 72.3654 | 174.7055
0.69457 | 1052165.6406 | 70.3402 | 169.8164
0.72195 | 1112653.7556 | 73.4226 | 177.2578
0.66817 | 986776.2954 | 75.4502 | 182.1528
0.6808 | 1021914.5992 | 73.2775 | 176.9076
0.69375 | 1041872.6533 | 65.9625 | 159.2476
0.69194 | 1064840.7644 | 75.1259 | 181.37
0.69237 | 1063629.7466 | 73.186 | 176.6865
0.69647 | 1054140.1145 | 73.8868 | 178.3784
0.68917 | 1037213.5903 | 62.0928 | 149.9054
0.70302 | 1073654.8527 | 65.0748 | 157.1044
0.70957 | 1094109.6 | 75.2733 | 181.7258
0.68458 | 1034338.5623 | 68.2861 | 164.8572
0.69463 | 1059773.898 | 76.2314 | 184.039
0.6834 | 1025692.3681 | 69.6913 | 168.2497
0.66141 | 974530.252 | 64.7155 | 156.2371
0.69192 | 1060668.7776 | 78.3165 | 189.0728
0.67774 | 1027094.7014 | 72.1981 | 174.3017
0.71702 | 1106484.0618 | 65.7549 | 158.7463
0.69927 | 1045968.4637 | 79.6735 | 192.3487
0.68847 | 1033602.6316 | 71.5228 | 172.6714
0.72048 | 1110587.0437 | 85.3267 | 205.9968
0.66818 | 995150.0985 | 65.4153 | 157.9265
0.70187 | 1063602.5879 | 82.3153 | 198.7267
0.68326 | 1005794.5128 | 78.4081 | 189.2939
0.6866 | 1015207.9367 | 61.0034 | 147.2752
0.6654 | 986027.7159 | 72.3397 | 174.6435
0.7171 | 1105861.6396 | 78.7866 | 190.2076
0.69088 | 1024296.1038 | 72.3758 | 174.7306
0.689 | 1043084.5761 | 72.1232 | 174.1208
0.71707 | 1117318.0172 | 89.3895 | 215.8052
0.68892 | 1047051.3905 | 77.6481 | 187.4591
0.69601 | 1045858.4954 | 87.6784 | 211.6744
0.70696 | 1080700.3551 | 83.0361 | 200.467
0.6894 | 1049686.0688 | 81.1013 | 195.796
0.69807 | 1046146.6223 | 78.2107 | 188.8174
0.7005 | 1052539.5874 | 76.7649 | 185.3269
0.68326 | 1029620.0305 | 83.4148 | 201.3812
0.68018 | 1013658.5761 | 74.4636 | 179.771
0.70104 | 1063776.1191 | 62.6045 | 151.1407
0.69245 | 1042232.064 | 75.2055 | 181.5622
0.66862 | 984630.4071 | 60.7569 | 146.6801
0.69329 | 1041999.0266 | 86.2236 | 208.1622
0.70363 | 1062828.7986 | 77.1699 | 186.3045
0.69892 | 1081891.6751 | 75.5661 | 182.4328
0.68313 | 1020629.4228 | 65.6216 | 158.4245
0.69548 | 1066778.471 | 74.8798 | 180.7758
0.70639 | 1065532.8246 | 70.0421 | 169.0967
0.68572 | 1021518.1836 | 76.1081 | 183.7413
0.71106 | 1083319.5091 | 76.125 | 183.7821
0.68426 | 1025650.5098 | 73.1728 | 176.6548
0.70229 | 1076587.8374 | 69.283 | 167.264
0.70608 | 1065417.2906 | 73.6612 | 177.8339
0.70819 | 1087169.006 | 73.1344 | 176.5621
0.70048 | 1068385.8228 | 82.8497 | 200.0168
0.67611 | 1000586.9896 | 81.9202 | 197.7729
0.65276 | 951748.488 | 76.0929 | 183.7046
0.69598 | 1043104.6267 | 69.5937 | 168.0141
0.69533 | 1051956.8633 | 58.1071 | 140.2829
0.68086 | 1012609.4605 | 83.837 | 202.4004
0.68362 | 1025418.6072 | 91.7966 | 221.6166
0.67337 | 1009379.6098 | 63.8727 | 154.2023
0.68298 | 1030653.1221 | 77.0971 | 186.1289
0.67276 | 1003072.8739 | 71.0315 | 171.4852
0.68474 | 1031690.5008 | 64.2011 | 154.9952
0.66729 | 985716.344 | 93.1009 | 224.7654
0.68151 | 1013458.831 | 83.8276 | 202.3777
0.66788 | 990211.2281 | 87.941 | 212.3083
0.69478 | 1052303.9818 | 64.5369 | 155.8058
0.69279 | 1048013.1553 | 73.8481 | 178.285
0.68737 | 1032520.8146 | 83.1498 | 200.7414
0.70717 | 1102178.7087 | 79.6267 | 192.236
0.68528 | 1031740.7527 | 71.7265 | 173.1631
0.68787 | 1044331.4109 | 74.1145 | 178.9282
0.68994 | 1037855.8318 | 57.776 | 139.4835
0.68969 | 1052226.9749 | 76.2313 | 184.0386
0.68702 | 1033185.6424 | 81.6892 | 197.2152
0.69676 | 1054850.8396 | 50.9982 | 123.1205
0.67733 | 1000954.5139 | 62.9693 | 152.0212
0.6785 | 1018923.1787 | 73.31 | 176.9859
0.68727 | 1032334.5545 | 75.0532 | 181.1944
0.70884 | 1095275.996 | 71.7481 | 173.2152
0.71191 | 1098110.0166 | 59.2444 | 143.0285
0.68687 | 1040464.4331 | 79.0543 | 190.854
0.70507 | 1069669.3287 | 68.5662 | 165.5335
0.68279 | 1021367.8275 | 76.4229 | 184.5013
0.71852 | 1113336.0599 | 77.4216 | 186.9124
0.66347 | 980283.2101 | 70.3894 | 169.935
0.70516 | 1083895.1429 | 76.8027 | 185.4182
0.68091 | 1002267.804 | 75.201 | 181.5514
0.69012 | 1053783.4535 | 76.2965 | 184.1961
0.71218 | 1094299.457 | 75.0103 | 181.0909
0.69393 | 1046821.7112 | 71.8679 | 173.5045
0.66747 | 996267.0109 | 71.1086 | 171.6713
0.69686 | 1048354.6664 | 73.4407 | 177.3016
0.67008 | 998626.3541 | 70.9994 | 171.4077
0.68815 | 1042589.3715 | 72.1788 | 174.2551
0.68318 | 1015744.7269 | 65.3024 | 157.654
0.69808 | 1061336.8209 | 79.3457 | 191.5576
0.71829 | 1109290.3875 | 65.8058 | 158.8693
0.70041 | 1068874.197 | 84.4745 | 203.9394
0.70443 | 1081624.028 | 76.7215 | 185.2222
0.68693 | 1032317.4666 | 59.899 | 144.609
0.68158 | 1025530.2081 | 76.8709 | 185.5828
0.69298 | 1048779.3524 | 70.2753 | 169.6597
0.70126 | 1062552.1715 | 85.4135 | 206.2064
0.66857 | 985988.6224 | 83.1341 | 200.7035
0.69013 | 1040545.0971 | 81.9011 | 197.7268
0.68609 | 1026046.9663 | 74.7243 | 180.4005
0.69338 | 1058541.409 | 86.0572 | 207.7604
0.7103 | 1094192.065 | 78.7017 | 190.0027
0.70706 | 1064207.6539 | 81.7577 | 197.3805
0.69413 | 1040232.4487 | 63.8567 | 154.1636
0.70119 | 1094977.337 | 86.4422 | 208.6898
0.68065 | 1022954.9877 | 76.7831 | 185.3708
0.69385 | 1039154.707 | 69.5429 | 167.8914
0.69306 | 1045228.7177 | 95.0511 | 229.4737
0.71195 | 1114758.8767 | 79.9946 | 193.124
0.69334 | 1041339.9179 | 83.1383 | 200.7136
0.69189 | 1032541.5929 | 50.2709 | 121.3646
0.69859 | 1077712.4124 | 75.3097 | 181.8137
0.6971 | 1058627.9187 | 77.5857 | 187.3085
0.68879 | 1036274.8417 | 55.5176 | 134.0313
0.69091 | 1065249.2694 | 80.4644 | 194.2582
0.70272 | 1064252.3441 | 77.3733 | 186.7956
0.6865 | 1020093.2776 | 67.1574 | 162.1322
0.67618 | 1004713.665 | 75.912 | 183.2678
0.68712 | 1036576.3443 | 64.2814 | 155.1891
0.69193 | 1054640.3206 | 81.1867 | 196.002
0.6793 | 1019132.4654 | 81.4939 | 196.7437
0.71223 | 1093751.0963 | 74.9668 | 180.986
0.68644 | 1048683.1022 | 74.5791 | 180.0499
0.71163 | 1089296.0273 | 72.9242 | 176.0545
0.67738 | 1000867.6267 | 68.6089 | 165.6366
0.69425 | 1067964.555 | 83.5595 | 201.7305
0.68227 | 1029772.4327 | 70.0844 | 169.1987
0.6794 | 1015647.574 | 71.2185 | 171.9366
0.68599 | 1039389.5892 | 68.4786 | 165.3218
0.69675 | 1074943.21 | 78.8204 | 190.2893
0.69435 | 1061841.0784 | 81.4552 | 196.6502
0.68424 | 1021411.7759 | 65.5658 | 158.2899
0.68267 | 1023868.7795 | 67.7843 | 163.6457
0.71511 | 1103740.1596 | 63.6792 | 153.7352
0.70595 | 1086299.2775 | 65.5678 | 158.2946
0.68926 | 1047542.0061 | 72.4042 | 174.7992
0.70324 | 1061624.8047 | 78.7844 | 190.2023
0.69918 | 1084020.602 | 75.6872 | 182.7251
0.64345 | 944188.0624 | 58.4871 | 141.2003
0.71559 | 1091949.624 | 76.2455 | 184.0729
0.70257 | 1080316.8744 | 81.9031 | 197.7316
0.6585 | 971285.146 | 70.2701 | 169.6471
0.68968 | 1031370.1442 | 67.5865 | 163.1683
0.6935 | 1050601.7275 | 79.0618 | 190.8721
0.6789 | 1008670.4429 | 54.8257 | 132.3609
0.6834 | 1027080.5257 | 80.0514 | 193.2611
0.68945 | 1035823.7209 | 80.0241 | 193.1952
0.68833 | 1035864.7795 | 83.3034 | 201.1122
0.69824 | 1070363.2424 | 63.2453 | 152.6876
0.66764 | 987082.4398 | 60.9979 | 147.2619
0.68826 | 1043878.6813 | 82.386 | 198.8973
0.68398 | 1036758.3656 | 73.9953 | 178.6406
0.68662 | 1038105.5753 | 72.0277 | 173.8903
0.66253 | 984465.7995 | 74.2457 | 179.2451
0.68891 | 1048604.7065 | 60.6695 | 146.4691
0.7054 | 1077839.5356 | 78.0638 | 188.4626
0.674 | 1015366.7421 | 86.425 | 208.6484
0.7004 | 1066885.4504 | 77.822 | 187.8789
0.68996 | 1039990.2971 | 71.8094 | 173.3633
0.70551 | 1085093.4039 | 88.2348 | 213.0175
0.70453 | 1088960.5232 | 78.9665 | 190.6421
0.67289 | 992609.389 | 58.7965 | 141.9473
0.70779 | 1085608.8066 | 78.3604 | 189.1786
0.69764 | 1058138.8399 | 60.4545 | 145.9501
0.69108 | 1034774.1963 | 69.5276 | 167.8546
0.6687 | 982332.9102 | 73.4016 | 177.2072
0.67926 | 1005919.9378 | 65.0449 | 157.0322
0.68763 | 1039331.5471 | 58.0529 | 140.1521
0.68856 | 1034953.2864 | 75.1022 | 181.3128
0.69782 | 1056671.1883 | 68.4373 | 165.2222
0.6983 | 1065091.3748 | 74.4597 | 179.7615
0.66367 | 979054.7346 | 66.6736 | 160.9642
0.67848 | 1022398.539 | 76.986 | 185.8607
0.7012 | 1051577.7785 | 83.0146 | 200.415
0.68731 | 1042407.6475 | 62.5726 | 151.0637
0.66487 | 969991.8987 | 77.7085 | 187.6049
0.67384 | 1003196.6162 | 63.4376 | 153.152
0.69816 | 1051329.6389 | 69.9232 | 168.8096
0.69633 | 1062451.8616 | 70.8491 | 171.0449
0.68037 | 991783.2629 | 64.0259 | 154.5722
0.68422 | 1021120.7208 | 61.5721 | 148.6481
0.69531 | 1058346.5934 | 79.4373 | 191.7786
0.66537 | 973531.1892 | 77.9344 | 188.1502
0.69552 | 1039665.6294 | 94.2107 | 227.4447
0.68557 | 1023372.4763 | 84.1465 | 203.1475
0.6671 | 987496.2801 | 70.5214 | 170.2538
0.69449 | 1049578.4486 | 64.3967 | 155.4675
0.68275 | 1029388.3358 | 53.2677 | 128.5996
0.70669 | 1081646.4639 | 74.7184 | 180.3862
0.68919 | 1040963.5761 | 82.0454 | 198.0751
0.67565 | 992164.9793 | 66.368 | 160.2266
0.67552 | 1008317.0796 | 65.5603 | 158.2767
0.69445 | 1059984.295 | 72.7165 | 175.5531
0.70026 | 1049158.2155 | 82.9349 | 200.2226
0.74308 | 1164211.1217 | 74.6402 | 180.1973
0.6969 | 1059991.3542 | 79.3787 | 191.6371
0.68217 | 1018998.2597 | 71.1639 | 171.8047
0.71082 | 1089839.7777 | 71.1931 | 171.8754
0.68793 | 1052462.0904 | 72.8857 | 175.9616
0.68662 | 1016899.3551 | 77.2922 | 186.6
0.6922 | 1053206.0804 | 79.4066 | 191.7045
0.69402 | 1035473.625 | 71.0797 | 171.6016
0.67741 | 1012989.1718 | 67.5851 | 163.1648
0.69421 | 1043776.8934 | 64.0665 | 154.6703
0.68074 | 1001021.1363 | 46.9929 | 113.4509
0.71393 | 1094187.708 | 65.3059 | 157.6625
0.68935 | 1026378.1829 | 65.396 | 157.8799
0.69837 | 1058680.5818 | 77.4485 | 186.9772
0.68574 | 1027697.169 | 79.8261 | 192.7173
0.69556 | 1048117.0287 | 78.6406 | 189.8552
0.67289 | 1004539.5548 | 77.9102 | 188.0918
0.68324 | 1016795.9768 | 91.0742 | 219.8725
0.69559 | 1059655.9907 | 88.8885 | 214.5958
0.70063 | 1058687.3159 | 68.2923 | 164.8723
0.68938 | 1026986.6399 | 63.478 | 153.2493
0.69017 | 1049974.6719 | 64.8246 | 156.5004
0.69065 | 1032611.1688 | 82.762 | 199.8052
0.68485 | 1022420.8958 | 61.5939 | 148.7008
0.67663 | 1012422.4616 | 69.4779 | 167.7344
0.69673 | 1050617.8689 | 73.6126 | 177.7166
0.69436 | 1077280.199 | 77.5207 | 187.1516
0.72769 | 1118839.1599 | 76.8237 | 185.4689
0.67209 | 1004642.8715 | 79.6202 | 192.2202
0.6903 | 1022801.651 | 74.0341 | 178.7342
0.68498 | 1019202.7115 | 80.8237 | 195.1256
0.68999 | 1044834.6494 | 62.4753 | 150.8288
0.68451 | 1021130.033 | 81.0319 | 195.6284
0.66981 | 998317.2425 | 74.4718 | 179.7908
0.70466 | 1073211.6088 | 66.8941 | 161.4965
0.7006 | 1067052.7574 | 90.3384 | 218.0963
0.68616 | 1013946.4948 | 83.3001 | 201.1041
0.72046 | 1109857.8641 | 83.679 | 202.0189
0.67046 | 994643.0603 | 82.1199 | 198.255
0.66849 | 986212.612 | 68.5039 | 165.3831
0.70641 | 1065830.7933 | 68.1915 | 164.6288
0.7032 | 1068127.2161 | 66.7159 | 161.0664
0.72695 | 1126417.2507 | 84.1693 | 203.2026
0.70166 | 1070551.5745 | 81.2684 | 196.1993
0.68944 | 1042623.3833 | 89.6998 | 216.5544
0.71608 | 1095634.711 | 83.1111 | 200.648
0.67765 | 1007795.6016 | 74.3098 | 179.3997
0.67841 | 1010768.273 | 63.4861 | 153.269
0.68566 | 1035141.9325 | 66.3091 | 160.0842
0.69946 | 1058255.097 | 73.8589 | 178.3112
0.69953 | 1064192.8652 | 79.3939 | 191.6737
0.6829 | 1019419.698 | 65.566 | 158.2904
0.66902 | 1004987.7997 | 83.1521 | 200.7468
0.69874 | 1074646.9269 | 78.574 | 189.6944
0.68893 | 1041928.2119 | 70.2265 | 169.5418
0.6882 | 1048150.3962 | 84.3068 | 203.5347
0.68588 | 1027373.785 | 74.8441 | 180.6895
0.69901 | 1051940.5511 | 78.5388 | 189.6094
0.69432 | 1063696.011 | 72.9848 | 176.201
0.67654 | 1001551.9469 | 70.5272 | 170.2678
0.70253 | 1081983.2436 | 54.4127 | 131.364
0.68929 | 1020477.7707 | 72.2726 | 174.4816
0.65064 | 944330.2022 | 75.7453 | 182.8654
0.66591 | 982885.6459 | 86.1241 | 207.9219
0.70388 | 1077900.334 | 84.7682 | 204.6485
0.71092 | 1092628.46 | 63.4096 | 153.0843
0.7057 | 1078563.6391 | 90.6537 | 218.8573
0.68547 | 1032436.958 | 64.1792 | 154.9422
0.71231 | 1090957.6002 | 58.8255 | 142.0174
0.71609 | 1113477.6112 | 81.1502 | 195.914
0.69021 | 1030292.9591 | 82.6296 | 199.4855
0.66911 | 993421.619 | 85.174 | 205.6282
0.69255 | 1042810.9814 | 80.2038 | 193.6291
0.70095 | 1060827.1778 | 66.4719 | 160.4773
0.67279 | 1000730.2928 | 63.3803 | 153.0135
0.70085 | 1081112.8728 | 61.4842 | 148.436
0.6936 | 1041154.3543 | 75.0042 | 181.076
0.66565 | 970856.0315 | 71.2001 | 171.8922
0.70137 | 1069608.9336 | 62.9186 | 151.8991
0.68736 | 1027223.9593 | 71.1965 | 171.8835
0.72123 | 1120719.4716 | 77.6421 | 187.4447
0.70554 | 1076631.7502 | 79.6377 | 192.2624
0.70425 | 1066237.4261 | 72.1279 | 174.1322
0.69526 | 1048073.3275 | 69.9973 | 168.9883
0.69883 | 1072673.7092 | 78.4197 | 189.322
0.70451 | 1075065.2835 | 77.9615 | 188.2157
0.69503 | 1037342.6477 | 67.4551 | 162.8511
0.70565 | 1080076.1961 | 56.4816 | 136.3587
0.69304 | 1041018.8357 | 86.1738 | 208.0419
0.68936 | 1041925.0719 | 78.8796 | 190.4321
0.67975 | 1021666.3266 | 66.0146 | 159.3734
0.67277 | 999141.0727 | 74.2058 | 179.1487
0.71028 | 1089290.8657 | 76.6291 | 184.9989
0.70528 | 1087940.9308 | 65.73 | 158.6862
0.67085 | 993367.5311 | 65.0383 | 157.0163
0.67071 | 1003131.8883 | 85.8019 | 207.144
0.68405 | 1031089.0544 | 73.5836 | 177.6465
0.68047 | 1004225.479 | 64.8193 | 156.4877
0.68475 | 1033096.9979 | 80.3617 | 194.0103
0.70458 | 1096872.1417 | 74.1658 | 179.052
0.67164 | 1005506.2628 | 74.8038 | 180.5924
0.6867 | 1023743.779 | 67.3783 | 162.6655
0.70798 | 1085408.2593 | 61.107 | 147.5253
0.67858 | 1017585.2331 | 65.8644 | 159.0106
0.66711 | 981437.1926 | 85.8427 | 207.2426
0.69477 | 1053024.2173 | 75.7309 | 182.8306
0.70179 | 1085563.1374 | 82.3925 | 198.9132
0.69253 | 1049767.1481 | 55.9363 | 135.0421
0.68054 | 1010364.657 | 77.4811 | 187.056
0.6859 | 1030652.3192 | 73.3538 | 177.0917
0.67623 | 1006779.8701 | 76.0623 | 183.6305
0.73361 | 1142628.6734 | 72.3233 | 174.6039
0.70547 | 1099321.9977 | 77.8617 | 187.9748
0.67615 | 986791.0685 | 56.0539 | 135.326
0.70798 | 1072652.91 | 60.3887 | 145.7912
0.6894 | 1030372.8332 | 79.6088 | 192.1926
0.69736 | 1046129.3986 | 63.9215 | 154.3202
0.68844 | 1021178.6237 | 54.7362 | 132.1449
0.65166 | 959868.1672 | 81.5355 | 196.844
0.65432 | 958450.8121 | 66.2023 | 159.8265
0.66592 | 989418.0731 | 74.2398 | 179.2307
0.70499 | 1075095.5606 | 82.8369 | 199.9861
0.68065 | 1028496.8071 | 78.8589 | 190.3823
0.68828 | 1041975.5177 | 60.3403 | 145.6743
0.69557 | 1050974.0746 | 75.4963 | 182.2642
0.70933 | 1094296.5073 | 75.3111 | 181.817
0.7003 | 1054373.241 | 69.9111 | 168.7804
0.69727 | 1062025.407 | 79.4892 | 191.904
0.68099 | 1016758.1056 | 87.1196 | 210.3253
0.67091 | 1007289.3575 | 63.9745 | 154.4482
0.69102 | 1052918.7932 | 64.9261 | 156.7455
0.70284 | 1060564.8 | 78.7511 | 190.122
0.70554 | 1071646.9991 | 72.361 | 174.695
0.6903 | 1060894.5581 | 81.86 | 197.6275
0.69817 | 1057838.0477 | 64.562 | 155.8666
0.6913 | 1031630.8742 | 57.3164 | 138.3741
0.68588 | 1036904.1732 | 77.2365 | 186.4654
0.66578 | 992919.4171 | 92.6414 | 223.656
0.68531 | 1027142.9045 | 61.8361 | 149.2855
0.69092 | 1047183.1556 | 82.3293 | 198.7605
0.69065 | 1041732.5614 | 71.8579 | 173.4802
0.70101 | 1075310.7514 | 56.6046 | 136.6555
0.68771 | 1028061.1917 | 72.9772 | 176.1825
0.66869 | 990023.1155 | 66.9138 | 161.5443
0.70721 | 1093895.3542 | 79.0242 | 190.7814
0.70037 | 1063067.1496 | 68.6006 | 165.6166
0.68994 | 1049106.0011 | 74.2307 | 179.2087
0.68882 | 1043932.0135 | 66.5655 | 160.7034
0.71528 | 1093911.529 | 81.4602 | 196.6623
0.69133 | 1037370.6722 | 84.586 | 204.2087
0.69051 | 1049419.7085 | 72.4031 | 174.7964
0.70315 | 1066892.8417 | 72.2849 | 174.5111
0.69419 | 1031119.5467 | 88.1029 | 212.6992
0.67546 | 1005990.2004 | 63.4618 | 153.2104
0.68996 | 1035712.6546 | 79.9569 | 193.0331
0.68063 | 1016098.7486 | 74.5931 | 180.0837
0.67753 | 994133.4091 | 79.7034 | 192.421
0.68874 | 1038228.0143 | 68.1288 | 164.4775
0.68151 | 1033040.5171 | 68.5077 | 165.3922
0.71309 | 1085446.7717 | 70.6807 | 170.6384
0.70006 | 1047892.8293 | 55.2362 | 133.3519
0.65578 | 964884.3612 | 64.6253 | 156.0192
0.68039 | 1024351.6137 | 93.5218 | 225.7815
0.6739 | 995300.6662 | 66.7505 | 161.15
0.69384 | 1053921.7628 | 56.0519 | 135.3212
0.69946 | 1055724.1118 | 68.2158 | 164.6876
0.6961 | 1060013.4947 | 72.5465 | 175.1427
0.68064 | 1005792.8443 | 72.1383 | 174.1572
0.69255 | 1046853.0933 | 60.2491 | 145.4541
0.70815 | 1091341.0648 | 82.0866 | 198.1746
0.69389 | 1041688.9445 | 79.9632 | 193.0481
0.695 | 1058730.5747 | 66.9086 | 161.5316
0.67529 | 1007512.0342 | 79.4695 | 191.8563
0.66063 | 978402.3076 | 68.77 | 166.0255
0.68404 | 1019464.257 | 71.0017 | 171.4133
0.67623 | 999616.5703 | 83.3276 | 201.1705
0.68618 | 1008477.9458 | 81.3552 | 196.4088
0.70065 | 1054139.6875 | 69.8613 | 168.6601
0.69811 | 1046445.2583 | 73.1756 | 176.6616
0.68804 | 1040521.454 | 68.3492 | 165.0096
0.6998 | 1062494.52 | 74.1188 | 178.9387
0.68682 | 1032561.6 | 76.8077 | 185.4301
0.67304 | 992558.1079 | 60.8433 | 146.8888
0.68156 | 1022354.9184 | 84.557 | 204.1386
0.71205 | 1097663.1543 | 80.4516 | 194.2274
0.69234 | 1042118.4954 | 62.1006 | 149.9241
0.66194 | 983155.8056 | 80.6797 | 194.778
0.69867 | 1069019.5728 | 79.8517 | 192.7792
0.67903 | 1014990.4917 | 69.4082 | 167.5661
0.70799 | 1076956.0415 | 74.258 | 179.2748
0.68101 | 1023294.0378 | 81.6218 | 197.0525
0.69004 | 1053751.6933 | 89.1044 | 215.117
0.66518 | 992965.64 | 67.4669 | 162.8795
0.66934 | 989584.2491 | 75.5326 | 182.3518
0.70957 | 1087239.7604 | 83.2309 | 200.9372
0.69312 | 1042748.1414 | 90.6615 | 218.8763
0.69836 | 1050765.1701 | 59.9716 | 144.7842
0.67535 | 1019688.7132 | 73.7072 | 177.9448
0.6566 | 960826.6213 | 79.4879 | 191.9008
0.69225 | 1046258.2966 | 50.9124 | 122.9134
0.66734 | 997056.0355 | 71.1791 | 171.8415
0.68973 | 1051876.7624 | 76.201 | 183.9655
0.68955 | 1032831.3291 | 73.1201 | 176.5275
0.68714 | 1050123.4396 | 79.7574 | 192.5515
0.68114 | 1014126.844 | 59.3436 | 143.2681
0.67482 | 1009811.4099 | 79.7066 | 192.4289
0.6962 | 1049313.4422 | 68.119 | 164.4537
0.71017 | 1075610.0351 | 80.6732 | 194.7623
0.69945 | 1063013.3794 | 72.4034 | 174.7974
0.67012 | 999177.0532 | 57.2685 | 138.2583
0.69226 | 1028750.0374 | 86.7119 | 209.3409
0.69023 | 1025912.0331 | 77.5688 | 187.2676
0.67518 | 1012456.348 | 69.6299 | 168.1015
0.68675 | 1040630.4228 | 68.4846 | 165.3364
0.67542 | 1005715.2792 | 74.4005 | 179.6187
0.69552 | 1042018.7516 | 90.2861 | 217.9699
0.69754 | 1050263.7085 | 68.0492 | 164.2853
0.70605 | 1074225.6701 | 87.2837 | 210.7215
0.71485 | 1113940.9357 | 72.3582 | 174.6882
0.68346 | 1035635.6085 | 77.4147 | 186.8957
0.67726 | 1013885.4497 | 69.5138 | 167.8213
0.67832 | 1013285.5766 | 61.6304 | 148.789
0.67315 | 994417.0775 | 83.2742 | 201.0416
0.69081 | 1040325.0484 | 66.1443 | 159.6866
0.66384 | 970598.8964 | 83.8782 | 202.4999
0.67572 | 992919.8171 | 60.6241 | 146.3595
0.68885 | 1036368.2452 | 67.9958 | 164.1564
0.70606 | 1074919.5466 | 85.8199 | 207.1876
0.68369 | 1022036.6223 | 76.3727 | 184.3801
0.7008 | 1077018.1446 | 67.5014 | 162.9627
0.68911 | 1041917.2126 | 75.5473 | 182.3873
0.67042 | 1006579.8367 | 65.9454 | 159.2064
0.72274 | 1123173.9593 | 84.4151 | 203.796
0.70571 | 1108056.1373 | 62.5827 | 151.088
0.68493 | 1037538.673 | 73.949 | 178.5287
0.69957 | 1061581.1269 | 60.4838 | 146.0209
0.67315 | 1003834.7617 | 76.1872 | 183.9321
0.66897 | 981251.9758 | 82.685 | 199.6192
0.69135 | 1049305.3668 | 80.8426 | 195.1713
0.67935 | 1018344.2354 | 69.6046 | 168.0404
0.7112 | 1085162.4143 | 68.9606 | 166.4855
0.6893 | 1024570.333 | 63.3359 | 152.9063
0.70213 | 1071895.4634 | 82.9266 | 200.2026
0.70002 | 1068441.4666 | 62.9712 | 152.0258
0.69322 | 1041460.2847 | 75.0673 | 181.2285
0.69688 | 1070769.4266 | 71.3472 | 172.2473
0.68371 | 1027600.2887 | 93.3623 | 225.3965
0.69153 | 1036110.0749 | 73.1427 | 176.5821
0.70452 | 1054054.067 | 73.0375 | 176.3281
0.70046 | 1061230.2748 | 85.4046 | 206.185
0.6838 | 1019806.9607 | 87.0214 | 210.0883
0.67957 | 1012551.4912 | 74.4488 | 179.7354
0.68776 | 1035068.504 | 80.1624 | 193.5292
0.67671 | 1019137.9523 | 75.2496 | 181.6686
0.69866 | 1080084.5054 | 73.6497 | 177.8061
0.70234 | 1070260.5027 | 60.7376 | 146.6336
0.69291 | 1062918.1775 | 80.056 | 193.2723
0.70267 | 1080013.1966 | 78.5248 | 189.5756
0.71891 | 1122771.3202 | 66.5477 | 160.6604
0.69348 | 1046086.4529 | 68.1759 | 164.5912
0.69966 | 1064642.3183 | 79.4122 | 191.718
0.69894 | 1048990.177 | 63.9507 | 154.3907
0.69853 | 1070246.6857 | 78.2766 | 188.9764
0.69763 | 1052476.8727 | 78.7336 | 190.0797
0.68773 | 1039948.5713 | 70.2153 | 169.5148
0.67979 | 1018673.8132 | 67.7261 | 163.5053
0.67293 | 999002.0652 | 68.9342 | 166.4218
0.6906 | 1050868.9302 | 70.6582 | 170.584
0.70086 | 1064364.8879 | 75.881 | 183.193
0.70713 | 1102472.436 | 81.366 | 196.4349
0.66169 | 969453.8794 | 79.5843 | 192.1335
0.67724 | 1015697.5502 | 57.6035 | 139.0672
0.68372 | 1036805.6887 | 79.1395 | 191.0596
0.70817 | 1085047.632 | 71.1455 | 171.7603
0.72248 | 1113209.6431 | 83.6964 | 202.061
0.71818 | 1093861.5143 | 62.8394 | 151.7078
0.69567 | 1058754.3032 | 74.2962 | 179.3669
0.70065 | 1062450.455 | 74.0253 | 178.7128
0.69969 | 1066723.9253 | 70.7924 | 170.908
0.7011 | 1059085.9899 | 88.1426 | 212.795
0.67925 | 1006067.9755 | 61.1081 | 147.5281
0.70562 | 1087762.3977 | 65.6874 | 158.5834
0.71011 | 1104259.8012 | 79.4463 | 191.8004
0.704 | 1077066.0433 | 74.6869 | 180.3102
0.70374 | 1075751.5967 | 69.0867 | 166.79
0.69985 | 1056177.4926 | 70.076 | 169.1783
0.71408 | 1096525.664 | 65.4317 | 157.9661
0.68127 | 1025534.1064 | 81.7687 | 197.407
0.7019 | 1063912.6819 | 75.7372 | 182.8457
0.69218 | 1046452.0516 | 68.3275 | 164.9571
0.69101 | 1039341.2657 | 71.6559 | 172.9926
0.69568 | 1052459.1988 | 83.743 | 202.1734
0.72551 | 1127238.2323 | 86.005 | 207.6343
0.67507 | 1016784.9592 | 60.0338 | 144.9345
0.67326 | 1000195.6699 | 75.8983 | 183.2347
0.70813 | 1081319.1029 | 65.0667 | 157.0849
0.70449 | 1063655.4758 | 66.3143 | 160.0968
0.70512 | 1071905.7394 | 82.5758 | 199.3557
0.70682 | 1092784.6195 | 85.5814 | 206.6117
0.68143 | 1019115.6106 | 61.7029 | 148.9639
0.68358 | 1027831.6512 | 81.6747 | 197.1801
0.69497 | 1071018.2883 | 80.4277 | 194.1697
0.70533 | 1063224.8894 | 78.6693 | 189.9245
0.68592 | 1036133.0125 | 80.4601 | 194.2478
0.69047 | 1054153.827 | 74.5299 | 179.931
0.68952 | 1031956.1681 | 85.8591 | 207.2822
0.69767 | 1061700.0668 | 66.8399 | 161.3658
0.68523 | 1040604.3484 | 74.1344 | 178.9764
0.68252 | 1007657.0118 | 68.5545 | 165.5053
0.69561 | 1043073.5533 | 79.8707 | 192.8249
0.67232 | 988173.6904 | 60.6394 | 146.3965
0.68888 | 1031521.5896 | 71.3366 | 172.2217
0.72162 | 1108499.4379 | 80.1683 | 193.5435
0.69164 | 1050554.9682 | 70.5306 | 170.2759
0.68607 | 1018595.5895 | 67.1011 | 161.9963
0.68274 | 1031823.2533 | 68.3037 | 164.8996
0.68363 | 1021642.3734 | 76.3503 | 184.326
0.68865 | 1039038.1108 | 68.1915 | 164.6288
0.66588 | 996725.466 | 71.3106 | 172.1589
0.68065 | 1018144.6201 | 73.023 | 176.2932
0.68998 | 1041604.1159 | 85.0984 | 205.4456
0.68833 | 1031865.7951 | 79.8488 | 192.7721
0.66493 | 971816.1873 | 64.1932 | 154.9762
0.67791 | 990633.8628 | 80.8896 | 195.2848
0.69456 | 1049139.7644 | 82.0506 | 198.0877
0.68758 | 1031439.8906 | 78.2883 | 189.0047
0.67132 | 992363.5351 | 72.1634 | 174.2179
0.70704 | 1074430.7306 | 60.5689 | 146.2262
0.69434 | 1058886.1754 | 77.913 | 188.0986
0.68053 | 1006162.745 | 72.1762 | 174.2488
0.70406 | 1065921.6415 | 70.3001 | 169.7195
0.69104 | 1052842.442 | 53.828 | 129.9522
0.68582 | 1035744.6665 | 68.2556 | 164.7836
0.68101 | 1013069.6156 | 74.3293 | 179.4467
0.70956 | 1084567.0328 | 65.13 | 157.2378
0.67671 | 1013481.8958 | 76.5148 | 184.7231
0.68599 | 1039979.8167 | 80.7506 | 194.9493
0.68864 | 1034407.8033 | 69.7486 | 168.3881
0.70834 | 1100110.6213 | 64.0759 | 154.693
0.69236 | 1042545.5232 | 64.1625 | 154.9019
0.68698 | 1047766.0117 | 70.8199 | 170.9745
0.67701 | 1005288.6942 | 79.0926 | 190.9465
0.67982 | 1027105.2782 | 75.4918 | 182.2534
0.68457 | 1028550.5195 | 81.6136 | 197.0327
0.71594 | 1119504.8305 | 95.6951 | 231.0283
0.70964 | 1072019.037 | 66.3514 | 160.1864
0.71791 | 1103102.0813 | 78.5017 | 189.5199
0.67978 | 1026378.4771 | 71.9314 | 173.6578
0.68018 | 1032823.9898 | 74.1274 | 178.9595
0.69606 | 1053933.4596 | 70.3083 | 169.7393
0.718 | 1106177.5983 | 71.5384 | 172.7089
0.7228 | 1110633.1552 | 78.5026 | 189.5221
0.68569 | 1014320.3239 | 73.914 | 178.4441
0.7174 | 1092697.2059 | 70.7258 | 170.7471
0.67448 | 1017593.6667 | 64.287 | 155.2025
0.68815 | 1045358.9902 | 75.9032 | 183.2465
0.70268 | 1085127.8315 | 80.5932 | 194.5693
0.69792 | 1067417.4236 | 76.4198 | 184.4937
0.69553 | 1036407.7448 | 50.1955 | 121.1827
0.70173 | 1072214.9828 | 69.0567 | 166.7176
0.71521 | 1107904.1738 | 80.843 | 195.1724
0.70826 | 1092526.2883 | 80.5035 | 194.3526
0.70774 | 1090222.549 | 54.4237 | 131.3904
0.6789 | 1011733.1933 | 70.9969 | 171.4018
0.68671 | 1046938.2637 | 72.6318 | 175.3487
0.7023 | 1066702.0985 | 77.5985 | 187.3393
0.67969 | 1019209.4784 | 63.8956 | 154.2577
0.69658 | 1045747.6508 | 75.0125 | 181.0963
0.67969 | 1003772.1754 | 84.7222 | 204.5376
0.66185 | 960785.2259 | 67.4064 | 162.7335
0.70438 | 1080631.6053 | 71.2376 | 171.9827
0.68482 | 1018863.9822 | 57.1428 | 137.9549
0.66561 | 983565.3312 | 58.8978 | 142.1918
0.69213 | 1050032.7026 | 85.5199 | 206.4634
0.66865 | 1002765.7981 | 87.5056 | 211.2572
0.69062 | 1030387.4405 | 50.2582 | 121.334
0.71163 | 1081555.6643 | 84.1294 | 203.1064
0.69867 | 1063153.5601 | 66.8352 | 161.3545
0.68363 | 1035725.2773 | 75.7011 | 182.7587
0.71935 | 1115431.0299 | 79.2192 | 191.252
0.69118 | 1017317.657 | 68.3442 | 164.9974
0.6759 | 1031452.4688 | 74.4975 | 179.8528
0.67818 | 1014386.6728 | 74.7371 | 180.4313
0.69292 | 1050546.4273 | 66.5907 | 160.7641
0.66955 | 987080.0582 | 90.221 | 217.8128
0.672 | 1000534.5212 | 65.3947 | 157.8767
0.66445 | 987579.1788 | 52.5118 | 126.7747
0.68675 | 1034037.6016 | 86.3792 | 208.5378
0.69138 | 1053866.5818 | 72.6178 | 175.3149
0.69772 | 1061482.8516 | 72.4788 | 174.9794
0.69624 | 1043688.907 | 67.75 | 163.5629
0.69094 | 1028800.5286 | 57.189 | 138.0664
0.69415 | 1057857.3695 | 68.8763 | 166.282
0.68696 | 1034825.9311 | 71.7647 | 173.2554
0.69894 | 1073782.4538 | 74.9975 | 181.0601
0.70022 | 1073656.4416 | 73.0477 | 176.3529
0.67603 | 992142.8469 | 79.7194 | 192.4598
0.69686 | 1053364.8799 | 83.6737 | 202.0062
0.68858 | 1028449.9432 | 69.3532 | 167.4335
0.70637 | 1081773.5858 | 87.8571 | 212.1059
0.69804 | 1059771.2321 | 82.4684 | 199.0964
0.68186 | 1005574.195 | 83.1331 | 200.701
0.693 | 1042810.1779 | 76.8781 | 185.6003
0.69106 | 1038011.5045 | 65.9732 | 159.2734
0.7053 | 1083494.7248 | 67.4737 | 162.8959
0.65237 | 947515.3286 | 80.3986 | 194.0993
0.70204 | 1064148.6715 | 73.9968 | 178.6442
0.67923 | 1005975.7338 | 81.9833 | 197.9251
0.67631 | 991859.1538 | 76.5377 | 184.7784
0.68521 | 1035303.6473 | 66.3626 | 160.2136
0.67727 | 1015714.1313 | 71.4476 | 172.4897
0.68283 | 1033980.6655 | 88.3561 | 213.3104
0.70557 | 1073892.3643 | 56.1644 | 135.5928
0.69122 | 1058364.5059 | 78.2259 | 188.8541
0.68137 | 1025858.4253 | 89.302 | 215.594
0.68996 | 1048710.7202 | 62.7096 | 151.3945
0.71295 | 1094977.4475 | 69.4129 | 167.5776
0.69413 | 1046749.0403 | 73.6169 | 177.727
0.69271 | 1053041.0314 | 82.5759 | 199.3559
0.6831 | 1035044.6915 | 73.9685 | 178.5758
0.71517 | 1120018.02 | 62.4212 | 150.6982
0.69515 | 1066743.6176 | 65.3878 | 157.86
0.69456 | 1055997.6072 | 70.5979 | 170.4384
0.70845 | 1075490.4464 | 70.4917 | 170.182
0.6891 | 1013566.4964 | 64.6032 | 155.966
0.70473 | 1085762.9441 | 83.6892 | 202.0436
0.69806 | 1052655.3073 | 57.9188 | 139.8284
0.66372 | 992939.1234 | 79.6846 | 192.3757
0.68618 | 1043564.9786 | 85.7305 | 206.9716
0.66892 | 988546.4939 | 73.8411 | 178.2681
0.69669 | 1058637.3319 | 75.9253 | 183.2998
0.70873 | 1094004.9443 | 65.7745 | 158.7936
0.6942 | 1052067.2594 | 75.7239 | 182.8138
0.69154 | 1036709.9967 | 69.6112 | 168.0564
0.69407 | 1038125.6872 | 67.5993 | 163.1991
0.69968 | 1054468.2898 | 75.3032 | 181.798
0.70857 | 1065555.0625 | 48.0574 | 116.0207
0.6944 | 1056713.6874 | 68.7774 | 166.0434
0.68425 | 1036876.4819 | 77.4186 | 186.9051
0.70455 | 1072945.6004 | 72.5508 | 175.1531
0.67729 | 1008519.3815 | 63.4491 | 153.1796
0.68129 | 1039242.4673 | 48.7478 | 117.6876
0.69467 | 1048020.8781 | 68.1846 | 164.6122
0.70458 | 1074898.1167 | 74.4746 | 179.7975
0.67405 | 1006602.0422 | 72.9023 | 176.0016
0.66904 | 992274.0206 | 63.1969 | 152.5708
0.71992 | 1123489.8311 | 77.9757 | 188.2499
0.68829 | 1036694.5416 | 96.2759 | 232.4307
0.67369 | 997900.1028 | 74.0124 | 178.6817
0.68469 | 1017561.7356 | 72.5509 | 175.1533
0.71302 | 1093227.2544 | 76.6104 | 184.9539
0.6926 | 1052775.3608 | 82.5682 | 199.3372
0.70094 | 1083434.0114 | 88.6738 | 214.0776
0.70362 | 1074437.5688 | 58.65 | 141.5937
0.68875 | 1040686.5037 | 66.3657 | 160.2211
0.67722 | 1030367.3976 | 64.8302 | 156.514
0.69733 | 1062443.8351 | 79.5138 | 191.9633
0.68816 | 1045771.4261 | 60.5712 | 146.2317
0.68289 | 1025295.8978 | 83.1094 | 200.6438
0.71224 | 1092718.2431 | 56.6139 | 136.678
0.7048 | 1073890.3149 | 76.1323 | 183.7997
0.70286 | 1068276.2808 | 68.6929 | 165.8394
0.69662 | 1053795.1747 | 72.0668 | 173.9846
0.69649 | 1048293.5337 | 71.7043 | 173.1095
0.69176 | 1036745.6335 | 78.2151 | 188.8279
0.6983 | 1061365.5882 | 74.1983 | 179.1305
0.71944 | 1141630.842 | 75.5975 | 182.5084
0.68157 | 1022932.9678 | 78.327 | 189.0981
0.69681 | 1058993.7308 | 85.6709 | 206.8278
0.70545 | 1084835.4327 | 81.9103 | 197.749
0.71736 | 1097528.3455 | 79.8457 | 192.7646
0.69545 | 1069413.3594 | 75.0674 | 181.2288
0.70435 | 1080487.8105 | 84.9147 | 205.0023
0.7191 | 1115132.6613 | 75.7029 | 182.763
0.66458 | 982562.3199 | 83.3427 | 201.207
0.71919 | 1115569.6653 | 82.6211 | 199.4651
0.6928 | 1059869.7565 | 65.1874 | 157.3764
0.6952 | 1040169.2748 | 59.6362 | 143.9745
0.69361 | 1043333.1181 | 78.5018 | 189.5201
0.7211 | 1118247.0488 | 79.2269 | 191.2707
0.68862 | 1041146.6989 | 75.546 | 182.3843
0.67645 | 1004643.4486 | 88.0917 | 212.6721
0.71346 | 1091966.6332 | 74.7159 | 180.3802
0.68661 | 1044083.6719 | 70.6099 | 170.4673
0.69006 | 1039856.6034 | 75.2943 | 181.7766
0.69308 | 1048614.4829 | 58.1109 | 140.292
0.6638 | 975391.0518 | 73.4209 | 177.2536
0.69333 | 1049692.7028 | 83.8063 | 202.3262
0.70371 | 1060206.815 | 50.7511 | 122.524
0.7166 | 1101065.1508 | 83.4633 | 201.4981
0.67026 | 1002673.7641 | 79.6818 | 192.3688
0.70597 | 1079700.3476 | 63.0988 | 152.3339
0.68994 | 1033903.6869 | 81.4002 | 196.5175
0.66963 | 997395.7508 | 77.7333 | 187.6647
0.68453 | 1027954.6748 | 73.3083 | 176.9818
0.70801 | 1100127.4141 | 70.2261 | 169.5407
0.68055 | 1021240.3299 | 67.8208 | 163.7338
0.70199 | 1064774.7679 | 75.7889 | 182.9705
0.67982 | 1025935.4482 | 68.2059 | 164.6636
0.66806 | 979935.917 | 68.1111 | 164.4348
0.69788 | 1040557.4616 | 84.0496 | 202.9136
0.72345 | 1131554.155 | 81.6282 | 197.0679
0.70742 | 1065460.4522 | 78.6021 | 189.7622
0.7052 | 1082734.2992 | 88.0813 | 212.6471
0.68897 | 1022878.8309 | 50.0115 | 120.7384
0.67459 | 996261.9781 | 66.7462 | 161.1395
0.68581 | 1030852.3652 | 76.7931 | 185.3949
0.69254 | 1038352.0402 | 71.4295 | 172.446
0.68933 | 1041869.1937 | 64.4025 | 155.4813
0.69225 | 1047587.0387 | 78.9769 | 190.667
0.67582 | 1001331.9485 | 59.0562 | 142.5743
0.68126 | 1017197.6609 | 72.8894 | 175.9706
0.67397 | 988259.3243 | 73.2665 | 176.8809
0.67323 | 994405.9937 | 73.8039 | 178.1785
0.6597 | 970041.7468 | 68.2779 | 164.8375
0.67841 | 1010606.8867 | 54.9704 | 132.7102
0.67842 | 1018480.023 | 83.1008 | 200.6231
0.68959 | 1047283.2399 | 85.61 | 206.6807
0.74026 | 1162986.6371 | 76.6944 | 185.1565
0.68804 | 1040860.6501 | 55.8238 | 134.7706
0.677 | 1006250.7921 | 74.2671 | 179.2966
0.69276 | 1056332.8878 | 70.8034 | 170.9345
0.67182 | 1002626.2267 | 65.549 | 158.2494
0.68476 | 1052242.6205 | 67.4935 | 162.9438
0.68803 | 1038108.507 | 66.2699 | 159.9897
0.66318 | 988429.7482 | 69.2075 | 167.0818
0.68243 | 1013630.4247 | 67.1301 | 162.0664
0.68673 | 1037710.5261 | 80.4879 | 194.3151
0.65511 | 981286.8056 | 73.1232 | 176.5351
0.69575 | 1063719.2323 | 72.837 | 175.8441
0.70158 | 1068913.086 | 64.9205 | 156.732
0.69722 | 1075514.7148 | 83.7918 | 202.2914
0.68561 | 1032784.9818 | 60.0411 | 144.952
0.68801 | 1039566.3876 | 68.2899 | 164.8664
0.68594 | 1024349.6964 | 74.6167 | 180.1407
0.71759 | 1097519.7137 | 78.1387 | 188.6436
0.71205 | 1098763.4549 | 82.1926 | 198.4304
0.69846 | 1071861.4547 | 73.1099 | 176.5029
0.69153 | 1027320.0379 | 68.9813 | 166.5355
0.70153 | 1049326.4628 | 67.6021 | 163.2058
0.70732 | 1079847.4273 | 76.2819 | 184.1607
0.69771 | 1078904.1963 | 71.8033 | 173.3485
0.68027 | 1013646.7398 | 69.1898 | 167.039
0.67276 | 989225.7022 | 54.2597 | 130.9945
0.692 | 1054219.5309 | 87.3981 | 210.9976
0.70354 | 1081917.7715 | 61.2944 | 147.9777
0.68403 | 1031621.7797 | 68.1617 | 164.5569
0.66386 | 980400.8918 | 83.2491 | 200.981
0.69219 | 1032141.4742 | 77.8196 | 187.8732
0.68816 | 1031787.3826 | 74.9169 | 180.8654
0.6928 | 1043438.4669 | 76.9466 | 185.7656
0.68567 | 1020798.217 | 61.4223 | 148.2865
0.6897 | 1030308.909 | 79.569 | 192.0967
0.71633 | 1104282.3626 | 69.6729 | 168.2053
0.68546 | 1022281.1976 | 85.5586 | 206.5567
0.68886 | 1034643.8787 | 67.8826 | 163.883
0.70277 | 1078666.1544 | 76.6923 | 185.1516
0.69784 | 1064346.2147 | 74.4561 | 179.7529
0.71005 | 1093060.1489 | 65.595 | 158.3604
0.68011 | 1019609.2191 | 69.613 | 168.0606
0.67562 | 1011164.2377 | 78.108 | 188.5694
0.68323 | 1032049.9353 | 66.3467 | 160.1751
0.69352 | 1054835.5343 | 88.7495 | 214.2601
0.68428 | 1044288.8926 | 68.974 | 166.518
0.68953 | 1027568.5136 | 74.462 | 179.7671
0.68888 | 1029117.4012 | 74.6446 | 180.2081
0.67759 | 1017009.7949 | 71.5053 | 172.629
0.70126 | 1059166.95 | 77.1644 | 186.2914
0.69043 | 1033407.8139 | 78.1139 | 188.5836
0.703 | 1068759.3134 | 75.9626 | 183.39
0.6852 | 1015023.1289 | 74.0176 | 178.6942
0.69357 | 1073349.3491 | 77.3263 | 186.6822
0.65974 | 977059.7157 | 70.5551 | 170.3351
0.68703 | 1040714.0885 | 60.3772 | 145.7634
0.70282 | 1046925.4334 | 63.4725 | 153.2361
0.6991 | 1059244.487 | 75.7063 | 182.7712
0.688 | 1040573.5432 | 85.2584 | 205.832
0.66698 | 984221.0886 | 56.6766 | 136.8294
0.70447 | 1075524.9799 | 74.5303 | 179.9322
0.68112 | 1017758.788 | 67.5334 | 163.04
0.69957 | 1070200.6034 | 76.4384 | 184.5387
0.70842 | 1078770.6542 | 79.4716 | 191.8614
0.68614 | 1023905.3054 | 77.8785 | 188.0154
0.69464 | 1052780.4909 | 79.0951 | 190.9526
0.69305 | 1039334.6121 | 78.2642 | 188.9465
0.68773 | 1026580.2352 | 61.5854 | 148.6803
0.66816 | 983187.4659 | 71.8339 | 173.4224
0.67928 | 1021109.4855 | 67.9391 | 164.0194
0.69746 | 1050332.7664 | 84.8911 | 204.9453
0.68357 | 1011585.1421 | 78.8241 | 190.2982
0.70597 | 1087333.8479 | 63.5446 | 153.4102
0.69455 | 1045569.4327 | 68.0371 | 164.2561
0.699 | 1055867.8034 | 71.9628 | 173.7335
0.65722 | 969163.8525 | 86.045 | 207.7311
0.70225 | 1068757.0933 | 62.7898 | 151.588
0.68417 | 1017226.4117 | 74.2902 | 179.3524
0.69844 | 1059291.1512 | 70.4245 | 170.0198
0.69905 | 1074509.4658 | 78.53 | 189.5881
0.70013 | 1065323.7573 | 77.7383 | 187.6767
0.6837 | 1040287.7511 | 68.768 | 166.0206
0.68498 | 1048826.3132 | 81.5581 | 196.8986
0.69545 | 1052529.7749 | 62.4405 | 150.7448
0.67152 | 1012809.1481 | 80.3356 | 193.9472
0.68997 | 1052807.2436 | 67.019 | 161.7983
0.68011 | 1009220.3 | 74.1453 | 179.0025
0.69011 | 1040861.1349 | 70.0343 | 169.0776
0.70384 | 1077487.4267 | 76.4922 | 184.6686
0.68758 | 1038487.7823 | 68.2611 | 164.7969
0.69064 | 1036197.2292 | 74.9663 | 180.9847
0.65583 | 958682.0505 | 51.645 | 124.6821
0.67063 | 996899.7412 | 69.0249 | 166.6409
0.67113 | 987043.7621 | 78.3377 | 189.1239
0.69225 | 1042303.1818 | 71.7314 | 173.175
0.69702 | 1061231.0862 | 74.3252 | 179.4369
0.69723 | 1045323.8884 | 74.9124 | 180.8545
0.69508 | 1049984.5705 | 71.966 | 173.7413
0.68155 | 1006398.5723 | 80.2103 | 193.6449
0.69908 | 1056489.7414 | 71.9976 | 173.8176
0.68093 | 1013237.7131 | 83.8811 | 202.5068
0.67542 | 1004253.3493 | 80.766 | 194.9864
0.67988 | 1026109.085 | 71.902 | 173.5868
0.70745 | 1082597.3186 | 71.057 | 171.5467
0.68334 | 1026199.1077 | 76.7231 | 185.2258
0.7061 | 1088567.8084 | 80.4568 | 194.2398
0.68872 | 1038989.0641 | 66.2599 | 159.9655
0.70186 | 1039974.5368 | 66.9678 | 161.6746
0.68054 | 1022892.4666 | 64.0461 | 154.621
0.69987 | 1051354.3276 | 78.3983 | 189.2703
0.69445 | 1061816.257 | 78.1157 | 188.5879
0.6954 | 1054757.1995 | 66.6191 | 160.8327
0.69353 | 1052879.9898 | 74.1116 | 178.9212
0.70897 | 1086901.2531 | 80.0085 | 193.1576
0.67712 | 1000426.3025 | 84.7604 | 204.6297
0.69063 | 1052699.9973 | 66.9483 | 161.6276
0.67815 | 1015655.7224 | 69.5442 | 167.8946
0.6833 | 1027753.0446 | 69.8419 | 168.6132
0.68988 | 1051257.9501 | 76.1404 | 183.8192
0.69914 | 1054284.9211 | 64.2554 | 155.1263
0.71113 | 1088251.214 | 67.2145 | 162.2701
0.70062 | 1060737.6445 | 79.0135 | 190.7555
0.70529 | 1083277.6468 | 74.9799 | 181.0175
0.70153 | 1075421.3179 | 65.3206 | 157.6979
0.67268 | 1021823.5137 | 72.3681 | 174.7121
0.69658 | 1055945.997 | 64.5936 | 155.9427
0.70359 | 1063719.3392 | 70.2011 | 169.4804
0.69784 | 1053136.9731 | 71.8964 | 173.5733
0.70581 | 1083220.3145 | 65.3956 | 157.879
0.70326 | 1080017.3691 | 61.2111 | 147.7767
0.68028 | 1008678.1574 | 80.5418 | 194.4452
0.66939 | 993509.3862 | 82.1571 | 198.3447
0.69578 | 1038354.518 | 47.6918 | 115.1383
0.70955 | 1106281.2096 | 79.6555 | 192.3053
0.66721 | 999074.4286 | 61.7788 | 149.1472
0.66058 | 980632.1629 | 79.0711 | 190.8945
0.68486 | 1033439.472 | 81.9668 | 197.8854
0.69869 | 1065173.4864 | 76.1746 | 183.9018
0.6773 | 1016470.0645 | 74.9876 | 181.0362
0.69723 | 1039553.0486 | 69.2324 | 167.1418
0.70984 | 1095921.7894 | 69.7779 | 168.4587
0.69496 | 1059887.8423 | 74.4516 | 179.742
0.70306 | 1087420.1512 | 65.9057 | 159.1105
0.69876 | 1077545.6373 | 78.7838 | 190.201
0.71002 | 1097674.7825 | 73.6515 | 177.8105
0.70291 | 1073413.3646 | 67.0738 | 161.9304
0.67241 | 1005563.7209 | 74.7231 | 180.3976
0.70309 | 1067929.2443 | 80.9044 | 195.3206
0.71727 | 1092039.3397 | 76.0729 | 183.6563
0.70393 | 1067682.7381 | 74.043 | 178.7556
0.7117 | 1093840.1899 | 83.4857 | 201.5523
0.69024 | 1041432.0362 | 53.3064 | 128.6931
0.70265 | 1067150.7015 | 75.8864 | 183.206
0.69188 | 1053579.7119 | 72.5457 | 175.1408
0.70405 | 1059917.7352 | 66.1226 | 159.6342
0.69548 | 1050506.8086 | 89.593 | 216.2965
0.71384 | 1097327.6925 | 73.2344 | 176.8036
0.70244 | 1071552.7573 | 70.834 | 171.0085
0.6854 | 1024400.6931 | 65.646 | 158.4835
0.68349 | 1021837.8163 | 94.2247 | 227.4786
0.69432 | 1048833.6582 | 82.0546 | 198.0973
0.69634 | 1043616.7557 | 59.7358 | 144.2149
0.71253 | 1109365.0644 | 77.7945 | 187.8125
0.68831 | 1031584.8652 | 63.861 | 154.174
0.66724 | 1010850.1883 | 77.3991 | 186.8579
0.71113 | 1094056.0465 | 75.1039 | 181.3168
0.68501 | 1013199.53 | 61.0234 | 147.3236
0.70865 | 1097694.2352 | 75.6173 | 182.5563
0.69245 | 1040461.9531 | 83.1744 | 200.8009
0.7121 | 1100225.9494 | 71.7425 | 173.2018
0.71255 | 1087823.7337 | 60.8503 | 146.9056
0.70419 | 1073422.2806 | 87.4484 | 211.1191
0.70803 | 1077070.9588 | 64.8203 | 156.4901
0.67621 | 1008203.1965 | 72.7847 | 175.7177
0.68831 | 1040904.1685 | 73.8245 | 178.228
0.70366 | 1079010.1076 | 69.3043 | 167.3153
0.68666 | 1022938.9725 | 65.432 | 157.9668
0.68081 | 1020251.1756 | 63.8556 | 154.1609
0.68146 | 1035407.6793 | 60.5604 | 146.2057
0.69063 | 1033720.3311 | 73.0691 | 176.4043
0.68789 | 1035816.0704 | 81.7938 | 197.4677
0.71123 | 1074519.1181 | 64.3751 | 155.4152
0.68801 | 1045022.6919 | 77.5995 | 187.3418
0.67579 | 1008230.124 | 71.0762 | 171.5931
0.66769 | 987246.7266 | 56.7907 | 137.1049
0.69193 | 1047776.9737 | 73.9229 | 178.4657
0.70368 | 1062994.5537 | 68.2738 | 164.8274
0.70493 | 1073883.0269 | 76.1185 | 183.7664
0.68175 | 1006891.422 | 68.5796 | 165.5658
0.70119 | 1077096.5186 | 78.4785 | 189.4638
0.69531 | 1052377.7643 | 85.3189 | 205.9779
0.68956 | 1058521.938 | 69.378 | 167.4933
0.68127 | 1021541.5712 | 84.5263 | 204.0647
0.68585 | 1034482.2978 | 66.2485 | 159.938
0.66542 | 976742.2589 | 58.5064 | 141.247
0.68731 | 1034519.1022 | 66.11 | 159.6037
0.67256 | 1013418.681 | 80.2591 | 193.7625
0.69673 | 1062295.055 | 77.8956 | 188.0565
0.69119 | 1030665.16 | 87.0076 | 210.0548
0.67204 | 1010796.4988 | 76.8418 | 185.5124
0.68014 | 1010479.4952 | 70.7798 | 170.8775
0.70711 | 1090609.3136 | 85.3755 | 206.1148
0.7078 | 1076207.5137 | 87.2284 | 210.588
0.68478 | 1034895.3112 | 80.4983 | 194.3401
0.69504 | 1051626.8616 | 73.8626 | 178.3202
0.69389 | 1041915.3608 | 79.98 | 193.0889
0.69966 | 1049552.8599 | 65.887 | 159.0652
0.67189 | 1021865.4889 | 84.1877 | 203.2471
0.68558 | 1041262.3754 | 99.7428 | 240.8004
0.69722 | 1046073.0294 | 76.6524 | 185.0552
0.66128 | 962482.3869 | 63.9282 | 154.3364
0.67605 | 1008251.1735 | 63.5488 | 153.4203
0.67354 | 999449.7456 | 85.1273 | 205.5154
0.70811 | 1080099.6873 | 82.8558 | 200.0316
0.65681 | 961957.0982 | 76.435 | 184.5304
0.66098 | 966003.1596 | 83.8426 | 202.414
0.6834 | 1012467.7514 | 72.3231 | 174.6035
0.68313 | 1044960.3901 | 84.6742 | 204.4215
0.70275 | 1066316.1793 | 64.6549 | 156.0907
0.67797 | 1012059.4265 | 65.5663 | 158.291
0.66441 | 975486.9685 | 71.6732 | 173.0344
0.6879 | 1038732.5899 | 70.1159 | 169.2748
0.70867 | 1075235.466 | 72.2834 | 174.5075
0.71904 | 1081319.6547 | 80.1019 | 193.383
0.6946 | 1039994.6134 | 59.6783 | 144.0761
0.69443 | 1062009.8842 | 79.8129 | 192.6854
0.70112 | 1075772.0701 | 63.3039 | 152.8292
0.71416 | 1095275.5616 | 83.3077 | 201.1225
0.66701 | 980138.9179 | 64.5414 | 155.8168
0.66559 | 1001996.1459 | 78.566 | 189.6751
0.69751 | 1053371.8022 | 75.4589 | 182.1738
0.71974 | 1119524.8814 | 74.1033 | 178.9013
0.68233 | 1029085.7997 | 68.1639 | 164.5622
0.67338 | 990283.847 | 58.6158 | 141.511
0.66486 | 976210.6687 | 68.1791 | 164.5988
0.68487 | 1027515.6568 | 65.8964 | 159.0879
0.70403 | 1064543.8891 | 64.4304 | 155.5488
0.68951 | 1030564.6414 | 78.0023 | 188.3142
0.69399 | 1047015.0222 | 66.2644 | 159.9765
0.69992 | 1048455.8356 | 65.7933 | 158.839
0.73293 | 1159634.5539 | 73.4677 | 177.3667
0.68126 | 995428.948 | 69.6106 | 168.0548
0.68333 | 1026348.2142 | 73.323 | 177.0175
0.70591 | 1090096.1937 | 78.8944 | 190.4679
0.67113 | 1003861.2912 | 74.738 | 180.4335
0.6631 | 983445.5062 | 72.4012 | 174.7919
0.70141 | 1069877.7756 | 72.6675 | 175.435
0.68017 | 1030757.9271 | 68.4912 | 165.3523
0.69505 | 1047531.5223 | 76.9988 | 185.8915
0.71071 | 1102190.4224 | 72.0112 | 173.8505
0.68925 | 1030850.3923 | 81.2408 | 196.1326
0.68001 | 1012875.2807 | 74.4252 | 179.6782
0.70468 | 1091276.7649 | 77.3785 | 186.8083
0.71875 | 1115468.6258 | 75.7619 | 182.9054
0.69592 | 1040036.1277 | 84.2451 | 203.3856
0.68715 | 1027263.1335 | 71.221 | 171.9426
0.66391 | 985482.988 | 70.7158 | 170.7231
0.69365 | 1026344.3661 | 81.9039 | 197.7336
0.69935 | 1058315.5501 | 75.6134 | 182.5468
0.67256 | 999269.8141 | 70.7832 | 170.8858
0.68251 | 1031869.1144 | 68.3304 | 164.9643
0.6634 | 988337.5603 | 67.3331 | 162.5565
0.68291 | 1020716.678 | 79.6122 | 192.2008
0.68379 | 1025142.1733 | 67.7576 | 163.5812
0.68042 | 1011784.8091 | 50.0908 | 120.9299
0.67711 | 1009312.2038 | 55.0055 | 132.795
0.68771 | 1023909.8053 | 88.311 | 213.2017
0.69232 | 1051096.739 | 75.8598 | 183.1418
0.69893 | 1040734.0069 | 72.6906 | 175.4906
0.70622 | 1085322.8484 | 86.6309 | 209.1454
0.66968 | 988447.299 | 68.5365 | 165.4617
0.69364 | 1060306.2778 | 81.4215 | 196.5688
0.68751 | 1041843.3282 | 79.0851 | 190.9282
0.69504 | 1049685.2546 | 73.7268 | 177.9923
0.67651 | 1013634.3146 | 78.1955 | 188.7807
0.69742 | 1048727.2146 | 87.7414 | 211.8264
0.71793 | 1096268.5152 | 67.0681 | 161.9168
0.69775 | 1063547.8319 | 56.247 | 135.7922
0.71806 | 1095969.0745 | 81.297 | 196.2683
0.68014 | 1031156.6209 | 72.2275 | 174.3726
0.67977 | 998416.6538 | 71.7003 | 173.0999
0.68164 | 1029120.289 | 66.312 | 160.0913
0.69737 | 1057191.0945 | 82.347 | 198.8033
0.68375 | 1029957.0987 | 78.6136 | 189.7901
0.68249 | 1029042.8101 | 86.4995 | 208.8283
0.68849 | 1031232.8658 | 73.1195 | 176.526
0.69822 | 1054050.5341 | 89.5428 | 216.1755
0.69463 | 1058041.8334 | 74.3573 | 179.5145
0.69758 | 1070522.2402 | 76.6788 | 185.119
0.70311 | 1076025.9795 | 78.1863 | 188.7585
0.67693 | 1012607.2449 | 73.0245 | 176.2967
0.71565 | 1107380.6531 | 77.3163 | 186.6581
0.7039 | 1063021.9641 | 66.3304 | 160.1359
0.69013 | 1029411.1724 | 69.2617 | 167.2125
0.68229 | 1035484.7362 | 77.5369 | 187.1906
0.68619 | 1029223.8866 | 57.6092 | 139.081
0.68943 | 1048311.5934 | 77.684 | 187.5459
0.69534 | 1049881.9729 | 70.11 | 169.2605
0.70544 | 1071084.6881 | 68.6217 | 165.6674
0.67872 | 1008344.1355 | 75.9012 | 183.2416
0.68027 | 1022676.5801 | 66.0294 | 159.4092
0.70551 | 1072234.6893 | 87.1893 | 210.4936
0.68134 | 1020680.2239 | 90.2776 | 217.9493
0.69095 | 1047790.7373 | 72.2445 | 174.4137
0.67585 | 1005409.3189 | 70.8707 | 171.0971
0.69741 | 1060730.6951 | 62.751 | 151.4943
0.70067 | 1069243.822 | 79.6408 | 192.27
0.69201 | 1038189.2929 | 63.9265 | 154.3322
0.69634 | 1053753.2875 | 85.3042 | 205.9425
0.682 | 1012821.4974 | 89.4956 | 216.0614
0.70221 | 1068063.2273 | 64.8106 | 156.4667
0.6799 | 1022604.8776 | 76.841 | 185.5106
0.68411 | 1033015.0461 | 71.9383 | 173.6743
0.67327 | 1000838.3303 | 72.5685 | 175.1958
0.70126 | 1068557.1981 | 66.9275 | 161.5772
0.70789 | 1090133.5434 | 80.3683 | 194.0263
0.69633 | 1057966.4514 | 66.7328 | 161.1071
0.69201 | 1042329.5766 | 83.3771 | 201.2902
0.68695 | 1032160.0907 | 71.5809 | 172.8116
0.70173 | 1059774.0397 | 75.4601 | 182.1769
0.68816 | 1024921.0771 | 83.9567 | 202.6893
0.66166 | 982484.0163 | 77.4939 | 187.0867
0.7105 | 1099155.7191 | 75.5212 | 182.3243
0.67564 | 999822.4595 | 70.4998 | 170.2016
0.68521 | 1027804.4103 | 71.4605 | 172.5208
0.68152 | 1015636.8166 | 69.1984 | 167.0597
0.67612 | 1005101.0839 | 85.2824 | 205.8899
0.69236 | 1053154.5828 | 68.7472 | 165.9704
0.68374 | 1003535.8148 | 69.9228 | 168.8086
0.73051 | 1148022.7409 | 82.0157 | 198.0035
0.70282 | 1060318.9917 | 70.0537 | 169.1246
0.68681 | 1061030.8008 | 85.5151 | 206.4517
0.70603 | 1090284.0466 | 69.793 | 168.4952
0.70461 | 1063528.2988 | 57.9728 | 139.9587
0.70239 | 1058275.6786 | 75.2827 | 181.7485
0.7141 | 1094721.5596 | 74.8372 | 180.6729
0.70556 | 1078798.2278 | 86.1901 | 208.0813
0.6896 | 1032453.511 | 71.3261 | 172.1966
0.68285 | 1016882.0109 | 63.4546 | 153.1929
0.69802 | 1040180.5676 | 69.1955 | 167.0527
0.68233 | 1043196.5533 | 72.837 | 175.844
0.67921 | 1012458.1784 | 74.4821 | 179.8156
0.70144 | 1085977.8251 | 69.9719 | 168.927
0.6939 | 1044256.8266 | 73.4661 | 177.3628
0.693 | 1068673.945 | 85.3625 | 206.0832
0.69013 | 1045626.4244 | 72.1807 | 174.2595
0.69277 | 1047294.9381 | 77.2341 | 186.4595
0.6917 | 1058984.3182 | 69.1942 | 167.0497
0.69309 | 1035218.0829 | 79.5581 | 192.0703
0.68458 | 1026344.951 | 67.7585 | 163.5835
0.68537 | 1026248.8097 | 74.8181 | 180.6269
0.69482 | 1061755.8118 | 69.0639 | 166.735
0.6693 | 972841.2412 | 57.9349 | 139.8673
0.69011 | 1036094.4047 | 69.7134 | 168.303
0.70879 | 1067082.6903 | 66.1249 | 159.6396
0.68607 | 1026569.3225 | 82.8272 | 199.9626
0.69397 | 1062627.4304 | 69.2046 | 167.0747
0.69728 | 1054574.8902 | 88.2491 | 213.0522
0.71128 | 1085724.1982 | 74.5163 | 179.8983
0.6679 | 991753.68 | 65.219 | 157.4525
0.67429 | 994195.5896 | 70.2426 | 169.5807
0.68181 | 1030410.4254 | 66.3789 | 160.2528
0.67734 | 1024212.8071 | 80.5425 | 194.4468
0.68416 | 1026356.1525 | 63.9232 | 154.3241
0.7222 | 1127202.9426 | 84.7812 | 204.6799
0.67828 | 1022498.2468 | 60.086 | 145.0604
0.70103 | 1068417.9638 | 68.8812 | 166.2938
0.717 | 1093150.6242 | 64.9645 | 156.8382
0.67911 | 1025712.7696 | 77.1639 | 186.2902
0.67061 | 996843.5913 | 72.5659 | 175.1897
0.66281 | 970223.7334 | 82.5157 | 199.2106
0.68095 | 1006123.9215 | 68.6079 | 165.634
0.73652 | 1154974.9567 | 60.2298 | 145.4076
0.69077 | 1045160.6722 | 78.9624 | 190.6322
0.69133 | 1047791.3731 | 75.7291 | 182.8263
0.66387 | 982964.3528 | 78.8672 | 190.4023
0.70771 | 1060922.1918 | 71.4172 | 172.4164
0.70109 | 1082193.1141 | 80.5569 | 194.4815
0.6949 | 1059859.3151 | 66.9901 | 161.7284
0.68101 | 1018029.9504 | 98.1712 | 237.0062
0.69283 | 1041677.5184 | 66.0375 | 159.4287
0.66189 | 962998.9729 | 77.4956 | 187.0908
0.67199 | 1003881.8195 | 65.6147 | 158.4079
0.68154 | 1018731.9454 | 79.6175 | 192.2137
0.71405 | 1090346.678 | 82.9453 | 200.2477
0.6996 | 1068761.4595 | 82.9827 | 200.3381
0.67655 | 1010409.4261 | 72.907 | 176.013
0.68323 | 1033919.0377 | 72.094 | 174.0503
0.72605 | 1110071.551 | 81.4375 | 196.6076
0.70903 | 1098400.4494 | 77.4052 | 186.8726
0.71842 | 1113035.478 | 77.4118 | 186.8885
0.70214 | 1073083.2729 | 69.0728 | 166.7566
0.69102 | 1045940.1631 | 72.1093 | 174.0872
0.69335 | 1056218.2165 | 70.736 | 170.7718
0.66617 | 993438.4138 | 76.1879 | 183.9339
0.69727 | 1050861.123 | 66.1949 | 159.8086
0.68152 | 1038518.9243 | 72.5956 | 175.2612
0.68814 | 1044335.4338 | 82.4368 | 199.0201
0.67676 | 1015183.8361 | 73.2293 | 176.7913
0.6758 | 993530.7067 | 81.3798 | 196.4682
0.7032 | 1074708.7893 | 69.0733 | 166.7576
0.66942 | 1001167.7706 | 71.2498 | 172.0123
0.68759 | 1044491.1041 | 74.156 | 179.0284
0.698 | 1053054.2061 | 71.8741 | 173.5194
0.70896 | 1087518.5639 | 78.3696 | 189.2008
0.69178 | 1055116.8942 | 64.9907 | 156.9014
0.68178 | 1020961.1865 | 73.1037 | 176.4878
0.69011 | 1037229.8187 | 64.0435 | 154.6148
0.71849 | 1101286.8915 | 76.5867 | 184.8966
0.67923 | 1012825.329 | 79.4254 | 191.7498
0.70439 | 1090901.1155 | 57.1496 | 137.9714
0.66317 | 985482.8602 | 70.7944 | 170.9128
0.68305 | 1026321.1782 | 73.3503 | 177.0832
0.68963 | 1041608.5665 | 78.8245 | 190.2992
0.67906 | 1020330.6149 | 61.9913 | 149.6603
0.6926 | 1016379.3392 | 69.9123 | 168.7832
0.67331 | 1001313.2368 | 79.2131 | 191.2373
0.7084 | 1089453.4739 | 75.0765 | 181.2506
0.68992 | 1025850.0761 | 77.599 | 187.3405
0.70193 | 1060184.3568 | 70.1337 | 169.3177
0.70185 | 1076457.1859 | 86.458 | 208.7281
0.69823 | 1062571.8151 | 62.786 | 151.5788
0.68072 | 1013764.6219 | 61.4427 | 148.3358
0.67524 | 1016847.4322 | 73.5141 | 177.4787
0.67638 | 1012627.5812 | 71.3246 | 172.1929
0.67683 | 1011839.3135 | 74.6984 | 180.3378
0.68123 | 1008152.1885 | 66.8847 | 161.4739
0.68969 | 1043942.6528 | 66.4825 | 160.5029
0.6838 | 1049364.1391 | 80.0882 | 193.3501
0.69746 | 1042492.44 | 67.5538 | 163.0892
0.68666 | 1040512.1561 | 79.0822 | 190.9212
0.70215 | 1071226.8406 | 77.9438 | 188.1729
0.69094 | 1042743.6935 | 62.2645 | 150.3198
0.69505 | 1064327.6475 | 78.2333 | 188.8719
0.66238 | 965212.1441 | 79.1538 | 191.0941
0.70951 | 1085913.4644 | 81.5228 | 196.8134
0.67107 | 991624.041 | 52.2063 | 126.0372
0.68719 | 1025304.2007 | 72.8363 | 175.8423
0.66607 | 979048.256 | 81.3647 | 196.4318
0.68849 | 1029634.0213 | 74.7238 | 180.3992
0.70233 | 1072009.7305 | 68.9846 | 166.5436
0.68588 | 1038888.0021 | 66.5274 | 160.6113
0.71649 | 1098528.4344 | 78.3304 | 189.1064
0.69907 | 1064968.8304 | 79.504 | 191.9397
0.70598 | 1080358.7858 | 84.7091 | 204.506
0.69918 | 1062745.324 | 84.1383 | 203.1278
0.70354 | 1065150.379 | 70.3009 | 169.7214
0.68268 | 1033215.7488 | 75.6224 | 182.5686
0.67542 | 1015018.1429 | 62.0017 | 149.6854
0.67322 | 997781.9651 | 61.148 | 147.6243
0.66206 | 978699.8237 | 58.9207 | 142.2471
0.70506 | 1072344.5666 | 85.7361 | 206.9852
0.68518 | 1037298.5284 | 64.5079 | 155.7358
0.69569 | 1045880.4949 | 73.3904 | 177.1802
0.68188 | 1010273.7344 | 70.5843 | 170.4056
0.68539 | 1029862.2243 | 69.7713 | 168.4427
0.70365 | 1093120.9367 | 75.8373 | 183.0875
0.69284 | 1036732.5972 | 50.4057 | 121.6901
0.69393 | 1034245.4479 | 69.2062 | 167.0787
0.68233 | 1029942.6332 | 73.6874 | 177.8971
0.65276 | 950195.6053 | 77.5352 | 187.1865
0.69842 | 1064180.8681 | 77.8443 | 187.9326
0.66855 | 992354.7145 | 80.6283 | 194.6538
0.67612 | 1013486.0822 | 69.6036 | 168.0379
0.68661 | 1019263.2454 | 63.2547 | 152.7103
0.71497 | 1088336.7174 | 72.5196 | 175.0779
0.70979 | 1080217.8159 | 88.8008 | 214.3841
0.66616 | 975568.8815 | 64.322 | 155.2872
0.68904 | 1036224.712 | 66.0534 | 159.467
0.68685 | 1033760.5427 | 86.3437 | 208.452
0.7058 | 1086804.3031 | 88.3375 | 213.2656
0.70134 | 1062533.7576 | 64.5553 | 155.8504
0.69176 | 1053192.3989 | 79.4557 | 191.8229
0.69826 | 1062652.7012 | 72.5018 | 175.0348
0.68343 | 1026977.3474 | 73.357 | 177.0995
0.69086 | 1029741.8955 | 72.7633 | 175.6661
0.69054 | 1033585.5299 | 87.3365 | 210.8488
0.71612 | 1102517.4235 | 67.6213 | 163.2522
0.6887 | 1066341.3797 | 80.8059 | 195.0826
0.67409 | 1008665.9342 | 50.4053 | 121.6893
0.69599 | 1054216.7668 | 75.6841 | 182.7175
0.65644 | 952656.6486 | 74.1433 | 178.9978
0.68108 | 1036109.2499 | 79.393 | 191.6717
0.6995 | 1066352.697 | 85.9338 | 207.4625
0.6925 | 1038704.5758 | 78.2795 | 188.9835
0.68209 | 1010805.3621 | 77.3391 | 186.713
0.67781 | 1014603.4377 | 67.6527 | 163.328
0.67785 | 1008905.0946 | 91.0864 | 219.9019
0.68148 | 1029376.5236 | 71.9613 | 173.73
0.7132 | 1087589.8225 | 72.6489 | 175.39
0.67939 | 1022232.298 | 73.4519 | 177.3286
0.68529 | 1037935.6278 | 72.1563 | 174.2008
0.69125 | 1056352.2548 | 76.3567 | 184.3415
0.69295 | 1040889.5634 | 69.7493 | 168.3898
0.66788 | 993663.8705 | 67.2415 | 162.3354
0.69728 | 1055094.0146 | 85.2972 | 205.9257
0.68418 | 1017025.8989 | 83.7942 | 202.2971
0.68575 | 1028752.3501 | 73.2346 | 176.804
0.69242 | 1027747.5934 | 85.1515 | 205.5739
0.70181 | 1073220.6376 | 72.8531 | 175.8829
0.67556 | 1003825.4098 | 70.7901 | 170.9025
0.68383 | 1011680.8684 | 81.8933 | 197.7078
0.69525 | 1056869.0358 | 75.9552 | 183.372
0.68848 | 1044463.3417 | 79.1727 | 191.1398
0.69404 | 1048723.0284 | 71.9869 | 173.7918
0.68216 | 1016374.716 | 70.981 | 171.3633
0.70122 | 1073798.4384 | 63.1789 | 152.5275
0.6993 | 1078338.279 | 89.977 | 217.2236
0.71683 | 1106011.3168 | 69.7331 | 168.3505
0.66206 | 989723.7144 | 72.978 | 176.1845
0.68793 | 1046352.156 | 64.8114 | 156.4687
0.68036 | 1012344.1686 | 76.0185 | 183.5248
0.68565 | 1021955.6641 | 71.0082 | 171.429
0.71046 | 1097796.9009 | 80.6415 | 194.6859
0.68522 | 1017856.3268 | 72.5155 | 175.068
0.69509 | 1042307.0355 | 63.8961 | 154.2588
0.69746 | 1055469.2735 | 53.4657 | 129.0775
0.68086 | 1013489.8486 | 65.3802 | 157.8417
0.70268 | 1089995.3554 | 80.9544 | 195.4412
0.68039 | 1025405.5829 | 80.275 | 193.801
0.6787 | 1021618.989 | 69.8953 | 168.7421
0.68238 | 1026337.1853 | 73.4235 | 177.2599
0.6892 | 1034773.5825 | 67.8573 | 163.8221
0.6782 | 1009113.619 | 80.8239 | 195.1261
0.68707 | 1037159.3343 | 88.081 | 212.6463
0.68625 | 1022008.7623 | 77.7224 | 187.6385
0.6898 | 1030737.4509 | 61.6492 | 148.8344
0.68508 | 1036077.3481 | 72.0489 | 173.9415
0.6635 | 980612.9754 | 89.066 | 215.0245
0.68079 | 1007852.2771 | 66.9086 | 161.5316
0.68537 | 1018886.3187 | 75.9506 | 183.361
0.67608 | 996496.8697 | 83.0869 | 200.5896
0.70138 | 1077074.2147 | 80.1435 | 193.4836
0.66382 | 985739.1867 | 74.1977 | 179.129
0.69751 | 1050243.9873 | 73.5047 | 177.456
0.69539 | 1055733.3768 | 85.5809 | 206.6106
0.69372 | 1040419.5893 | 71.4496 | 172.4946
0.66845 | 999485.9508 | 82.7679 | 199.8195
0.66841 | 972469.2109 | 69.3124 | 167.335
0.70042 | 1069667.6619 | 79.7011 | 192.4155
0.69094 | 1039018.3573 | 83.241 | 200.9615
0.70466 | 1068111.1159 | 73.9962 | 178.6426
0.66587 | 968415.4462 | 67.5323 | 163.0373
0.71529 | 1093093.1083 | 60.9103 | 147.0504
0.68783 | 1032937.8359 | 79.9895 | 193.1117
0.68596 | 1024186.9851 | 74.6031 | 180.1079
0.68828 | 1037469.6501 | 68.125 | 164.4682
0.66328 | 990973.4912 | 81.156 | 195.928
0.68946 | 1047816.1535 | 85.6703 | 206.8264
0.68533 | 1036619.7473 | 83.7284 | 202.1384
0.68235 | 1028804.4621 | 75.3363 | 181.8779
0.67676 | 1011960.0099 | 88.6921 | 214.1218
0.68493 | 1038203.3294 | 82.3369 | 198.7789
0.69325 | 1043580.5337 | 62.1065 | 149.9382
0.71057 | 1097296.0853 | 91.5957 | 221.1316
0.68493 | 1012110.2204 | 72.137 | 174.154
0.69624 | 1063982.7142 | 86.1441 | 207.9703
0.6916 | 1045511.7132 | 78.8946 | 190.4684
0.67741 | 1008618.8106 | 76.0684 | 183.6453
0.68741 | 1024259.0333 | 69.1879 | 167.0345
0.68859 | 1024740.8063 | 57.1581 | 137.9918
0.70955 | 1085366.7953 | 69.1337 | 166.9035
0.68198 | 1021436.8009 | 76.5196 | 184.7346
0.69836 | 1070612.1055 | 83.2737 | 201.0405
0.6824 | 1031874.715 | 48.323 | 116.6621
0.69822 | 1075736.8952 | 67.6702 | 163.3703
0.66553 | 988634.5165 | 75.8926 | 183.2209
0.68159 | 1024645.7217 | 67.9434 | 164.03
0.68898 | 1025374.5875 | 63.4154 | 153.0982
0.69845 | 1061218.6643 | 62.6856 | 151.3365
0.70019 | 1080802.2316 | 77.1266 | 186.2002
0.67786 | 1027444.2358 | 75.7343 | 182.8387
0.68266 | 1009593.6513 | 80.3339 | 193.9432
0.67296 | 994497.0331 | 68.6201 | 165.6636
0.66685 | 997343.6876 | 63.7288 | 153.8549
0.6682 | 996713.5445 | 79.4517 | 191.8135
0.68375 | 1025252.5892 | 90.3627 | 218.155
0.66318 | 977075.6302 | 68.4127 | 165.163
0.6939 | 1055039.41 | 82.5902 | 199.3904
0.67085 | 1003631.4982 | 78.0774 | 188.4956
0.702 | 1081890.1359 | 85.3106 | 205.958
0.68114 | 1015631.7973 | 71.0358 | 171.4956
0.69729 | 1073329.7875 | 63.5066 | 153.3184
0.66447 | 991997.6866 | 63.6144 | 153.5787
0.70886 | 1086555.46 | 79.9951 | 193.1252
0.68476 | 1042165.7515 | 82.672 | 199.5878
0.67917 | 1011902.0027 | 70.9797 | 171.3602
0.6855 | 1036488.3372 | 60.2668 | 145.4969
0.67476 | 1018679.5409 | 60.174 | 145.2728
0.66631 | 993627.9909 | 74.5114 | 179.8864
0.70251 | 1065580.4983 | 85.5043 | 206.4257
0.70252 | 1076693.2398 | 73.0153 | 176.2744
0.69465 | 1058432.6397 | 77.8442 | 187.9325
0.67881 | 1015371.3451 | 73.0219 | 176.2904
0.67889 | 1007556.0375 | 82.1533 | 198.3355
0.71223 | 1090445.5163 | 81.5502 | 196.8796
0.6726 | 1015041.7752 | 66.2846 | 160.0252
0.66942 | 988395.4726 | 75.4333 | 182.1121
0.67972 | 1002792.2769 | 80.1378 | 193.4698
0.66644 | 987710.4269 | 83.5453 | 201.6963
0.68343 | 1007064.7551 | 73.6278 | 177.7533
0.67751 | 1027555.1132 | 66.6149 | 160.8226
0.68513 | 1028360.3234 | 74.9984 | 181.0621
0.69456 | 1045916.2324 | 76.9344 | 185.736
0.69109 | 1032546.7643 | 67.5262 | 163.0227
0.67959 | 1025902.1043 | 88.0837 | 212.6529
0.69385 | 1054825.9272 | 68.7087 | 165.8776
0.68869 | 1041952.0325 | 83.7483 | 202.1862
0.67518 | 1008414.4298 | 71.1531 | 171.7787
0.67557 | 1010374.8618 | 45.0609 | 108.7867
0.6761 | 1009769.9348 | 72.9097 | 176.0196
0.68238 | 1028941.7441 | 75.3427 | 181.8933
0.68866 | 1042389.0229 | 82.2212 | 198.4996
0.70664 | 1065582.1892 | 73.6139 | 177.7198
0.68986 | 1033480.4521 | 61.8073 | 149.2159
0.66467 | 971538.3339 | 82.3958 | 198.9211
0.68034 | 1010680.2786 | 74.2274 | 179.2008
0.69246 | 1036859.3348 | 67.6739 | 163.3793
0.70527 | 1082787.89 | 84.9848 | 205.1714
0.69248 | 1054789.4049 | 64.9046 | 156.6935
0.67619 | 992340.0541 | 60.5471 | 146.1737
0.67729 | 1000436.9208 | 72.1087 | 174.0858
0.69838 | 1048446.2997 | 75.0118 | 181.0946
0.71481 | 1084283.7178 | 74.0503 | 178.7733
0.69525 | 1054737.8399 | 85.5182 | 206.4593
0.70003 | 1048803.2976 | 67.7131 | 163.4739
0.7176 | 1111130.4472 | 93.3118 | 225.2747
0.67813 | 1013703.2522 | 55.162 | 133.1729
0.66424 | 976898.0723 | 62.6668 | 151.2911
0.68541 | 1029300.7561 | 65.9588 | 159.2387
0.70122 | 1067608.0393 | 62.8146 | 151.6478
0.67572 | 1013578.9669 | 71.4611 | 172.5223
0.68038 | 1031884.5038 | 82.3829 | 198.89
0.68248 | 1025267.3808 | 80.018 | 193.1806
0.67103 | 987349.9427 | 76.7901 | 185.3877
0.67274 | 989188.7774 | 51.9489 | 125.4157
0.69867 | 1044973.1905 | 81.2442 | 196.1408
0.68256 | 1034930.5013 | 72.2352 | 174.3913
0.70378 | 1072922.5346 | 79.562 | 192.0797
0.69345 | 1047563.6597 | 78.8381 | 190.3319
0.70443 | 1077745.3427 | 67.1986 | 162.2317
0.66952 | 985617.3509 | 71.8649 | 173.4971
0.68592 | 1032869.406 | 84.884 | 204.9282
0.69451 | 1054984.7319 | 68.6538 | 165.7449
0.69391 | 1054819.1829 | 77.6618 | 187.4921
0.69469 | 1039219.4291 | 85.8204 | 207.1888
0.72158 | 1099450.1077 | 70.7502 | 170.8061
0.69315 | 1056927.5218 | 72.6428 | 175.3752
0.70072 | 1065546.2091 | 72.4442 | 174.8959
0.70132 | 1061289.3928 | 80.6046 | 194.5967
0.6618 | 964168.7377 | 69.4072 | 167.5637
0.71148 | 1086352.0956 | 74.5388 | 179.9526
0.69031 | 1036263.6581 | 65.9339 | 159.1786
0.6878 | 1028535.5902 | 90.1177 | 217.5634
0.70002 | 1059272.15 | 83.7631 | 202.222
0.68902 | 1045358.1714 | 75.8068 | 183.0138
0.67453 | 1015554.1123 | 86.3729 | 208.5225
0.71133 | 1084517.6347 | 77.8931 | 188.0506
0.70271 | 1072964.4248 | 87.5962 | 211.4758
0.70693 | 1076890.4701 | 64.9097 | 156.7058
0.6773 | 1019323.8839 | 66.102 | 159.5843
0.67925 | 1017162.6245 | 82.8927 | 200.1207
0.699 | 1063829.7964 | 72.5054 | 175.0435
0.67219 | 1002799.7745 | 69.2729 | 167.2395
0.67489 | 1012188.3937 | 78.2716 | 188.9644
0.69256 | 1053502.3633 | 76.0076 | 183.4986
0.69339 | 1055239.4863 | 70.666 | 170.6027
0.66859 | 1002089.016 | 71.4046 | 172.386
0.6804 | 1020714.7624 | 71.4691 | 172.5417
0.67299 | 990795.1009 | 62.7322 | 151.4489
0.66949 | 991273.0712 | 66.9186 | 161.5558
0.68908 | 1034371.756 | 73.2789 | 176.9109
0.70212 | 1086103.8122 | 80.73 | 194.8996
0.68515 | 1025365.0222 | 71.2965 | 172.125
0.68581 | 1027310.7193 | 69.556 | 167.923
0.71449 | 1101240.1165 | 78.7436 | 190.104
0.709 | 1091375.2195 | 65.5915 | 158.3518
0.66969 | 984752.1977 | 68.2797 | 164.8417
0.68055 | 1011438.7097 | 65.2769 | 157.5924
0.6925 | 1038811.5923 | 76.6388 | 185.0225
0.67743 | 1022250.4375 | 64.2997 | 155.2331
0.68557 | 1034975.7047 | 71.9211 | 173.6329
0.68302 | 1021204.8632 | 74.7059 | 180.3561
0.67314 | 994320.265 | 78.6794 | 189.9489
0.71473 | 1094354.504 | 79.9074 | 192.9136
0.68227 | 1018045.4285 | 84.113 | 203.0667
0.68912 | 1036484.0669 | 76.8421 | 185.5132
0.68418 | 1036792.4994 | 70.3658 | 169.878
0.6714 | 985961.5448 | 66.3701 | 160.2316
0.67631 | 1008713.7772 | 59.4485 | 143.5215
0.70563 | 1085332.0006 | 69.5158 | 167.8261
0.6664 | 975863.3796 | 62.3981 | 150.6422
0.69749 | 1044251.8039 | 69.5197 | 167.8353
0.69093 | 1044834.3615 | 77.2582 | 186.5178
0.66524 | 971291.9297 | 74.5936 | 180.0849
0.67559 | 1007888.8058 | 69.2739 | 167.242
0.65877 | 965445.6002 | 71.5435 | 172.7212
0.68809 | 1044294.1241 | 66.6589 | 160.9289
0.68036 | 1025712.8087 | 73.5972 | 177.6794
0.68605 | 1017199.0327 | 72.3806 | 174.7421
0.68111 | 1023471.9417 | 67.9134 | 163.9574
0.69933 | 1067291.1447 | 68.9756 | 166.5218
0.71513 | 1107279.1279 | 91.8398 | 221.7209
0.6776 | 1019037.679 | 73.416 | 177.2419
0.68248 | 1013899.1855 | 74.2302 | 179.2077
0.70781 | 1095405.5182 | 54.6771 | 132.0022
0.68274 | 1027500.099 | 66.1009 | 159.5816
0.67103 | 979280.9754 | 67.765 | 163.5992
0.66654 | 977775.6436 | 85.5958 | 206.6467
0.67836 | 1011833.5677 | 65.5141 | 158.165
0.69886 | 1048616.0673 | 69.6725 | 168.2044
0.69788 | 1065342.3622 | 91.9951 | 222.0959
0.68216 | 1030381.6592 | 73.737 | 178.0169
0.69735 | 1066760.749 | 71.0315 | 171.4853
0.68705 | 1039628.2555 | 75.6074 | 182.5324
0.68392 | 1021351.7152 | 72.9022 | 176.0015
0.68495 | 1032709.2965 | 86.4957 | 208.8192
0.67471 | 1016914.6136 | 62.1828 | 150.1226
0.70198 | 1059735.6647 | 80.4099 | 194.1267
0.67553 | 993143.4656 | 66.4453 | 160.4131
0.67796 | 1002557.2139 | 78.7671 | 190.1606
0.69882 | 1062389.9788 | 82.58 | 199.3659
0.71313 | 1089780.182 | 83.01 | 200.4038
0.71158 | 1102354.572 | 84.1502 | 203.1566
0.68987 | 1040773.8341 | 78.7344 | 190.0816
0.69065 | 1042425.917 | 64.3138 | 155.2673
0.69381 | 1059811.3297 | 60.8844 | 146.988
0.71902 | 1099994.0719 | 71.0372 | 171.4989
0.6703 | 994529.8899 | 72.4301 | 174.8617
0.71768 | 1114506.8418 | 72.1088 | 174.0861
0.70855 | 1100956.2659 | 81.0905 | 195.7699
0.67127 | 1013824.4663 | 78.8209 | 190.2904
0.69738 | 1047719.1841 | 87.6628 | 211.6366
0.69881 | 1058075.4564 | 72.1335 | 174.1457
0.6895 | 1022336.3528 | 78.9772 | 190.6679
0.69267 | 1037938.2922 | 78.2021 | 188.7965
0.69304 | 1062023.1007 | 92.3483 | 222.9485
0.67388 | 1007633.1452 | 78.0358 | 188.395
0.68165 | 1017641.6612 | 67.4734 | 162.8951
0.69485 | 1038026.622 | 55.1263 | 133.0866
0.70683 | 1084745.2267 | 89.5463 | 216.1839
0.70359 | 1082082.4557 | 80.2627 | 193.7713
0.69128 | 1041313.038 | 80.236 | 193.7068
0.697 | 1062953.6802 | 81.2312 | 196.1096
0.68278 | 1014696.3124 | 77.5583 | 187.2423
0.69764 | 1058353.0048 | 79.1719 | 191.1378
0.68442 | 1029938.4665 | 84.8285 | 204.7942
0.68919 | 1045724.9769 | 83.9912 | 202.7726
0.69558 | 1041525.9757 | 69.5279 | 167.8553
0.66865 | 996971.3842 | 57.2675 | 138.256
0.70874 | 1090120.7662 | 59.0763 | 142.6229
0.67265 | 995067.7722 | 70.4762 | 170.1445
0.67599 | 1016440.4032 | 80.0901 | 193.3547
0.68006 | 1036412.1206 | 82.7833 | 199.8565
0.68045 | 1018095.7105 | 71.2808 | 172.087
0.69222 | 1046071.5385 | 68.2992 | 164.8888
0.70084 | 1057383.1603 | 73.6524 | 177.8126
0.70283 | 1071896.7827 | 81.2351 | 196.1189
0.67772 | 1013411.819 | 74.7162 | 180.3809
0.68149 | 1039816.2806 | 65.015 | 156.9602
0.66911 | 980895.1118 | 68.0708 | 164.3375
0.68327 | 1013027.5384 | 72.3778 | 174.7354
0.68985 | 1027545.8793 | 72.7138 | 175.5466
0.69651 | 1063408.2345 | 81.323 | 196.3311
0.69918 | 1053568.0057 | 84.9424 | 205.0692
0.7021 | 1062210.2161 | 62.418 | 150.6903
0.68067 | 1029026.1364 | 66.4003 | 160.3046
0.68961 | 1049889.1544 | 72.5298 | 175.1024
0.69809 | 1060320.2492 | 68.6473 | 165.7292
0.66575 | 979326.265 | 70.0795 | 169.187
0.67694 | 1001795.1832 | 56.8399 | 137.2236
0.69244 | 1038852.0722 | 59.7985 | 144.3664
0.68815 | 1038616.8968 | 81.9844 | 197.9278
0.68086 | 1012755.863 | 61.369 | 148.1579
0.68875 | 1049247.0361 | 81.1438 | 195.8985
0.71399 | 1076957.1838 | 70.0398 | 169.0912
0.6972 | 1065036.2715 | 70.9321 | 171.2453
0.69127 | 1035010.3403 | 61.7729 | 149.1329
0.68978 | 1036905.069 | 55.4718 | 133.9208
0.71295 | 1095660.5652 | 71.36 | 172.2783
0.68019 | 1013910.6186 | 79.3979 | 191.6835
0.70624 | 1099269.6216 | 87.742 | 211.828
0.70269 | 1076650.1644 | 61.0698 | 147.4355
0.67702 | 1012563.9458 | 57.4443 | 138.6828
0.69813 | 1043843.096 | 63.7108 | 153.8114
0.67676 | 997260.1595 | 76.9033 | 185.6609
0.67325 | 1009611.1999 | 65.8655 | 159.0135
0.71058 | 1080676.0573 | 68.7351 | 165.9411
0.69267 | 1056552.3272 | 84.4349 | 203.8438
0.69788 | 1061697.8962 | 82.6378 | 199.5052
0.69457 | 1060083.244 | 77.9745 | 188.2471
0.70072 | 1060910.6393 | 66.4991 | 160.5431
0.70737 | 1074289.0426 | 70.52 | 170.2504
0.67602 | 1002883.6948 | 67.0477 | 161.8674
0.70795 | 1069757.3854 | 77.1993 | 186.3755
0.69585 | 1047730.4743 | 78.579 | 189.7064
0.70623 | 1074259.8518 | 64.816 | 156.4797
0.6913 | 1033139.6618 | 82.7267 | 199.72
0.69154 | 1053379.6891 | 72.5678 | 175.194
0.70054 | 1057118.3686 | 80.3764 | 194.0457
0.6754 | 1003841.9551 | 59.2406 | 143.0195
0.69073 | 1047794.8707 | 70.4383 | 170.0531
0.69076 | 1043538.1914 | 55.7474 | 134.586
0.71448 | 1094522.2258 | 79.9486 | 193.0129
0.66703 | 974684.0035 | 60.2864 | 145.5443
0.68393 | 1042366.3503 | 70.7698 | 170.8533
0.67942 | 1017300.7211 | 78.1424 | 188.6525
0.70281 | 1067130.1329 | 79.2413 | 191.3054
0.69313 | 1052072.4722 | 75.8419 | 183.0985
0.67188 | 997034.1114 | 75.8143 | 183.0319
0.69804 | 1060246.3156 | 77.5991 | 187.3409
0.68158 | 1023116.7933 | 76.6903 | 185.1467
0.69639 | 1046198.2555 | 69.7586 | 168.4121
0.67558 | 1002680.8688 | 57.0044 | 137.6207
0.68822 | 1030074.7932 | 71.3051 | 172.1457
0.69821 | 1066491.1228 | 60.8818 | 146.9818
0.71102 | 1082509.7961 | 72.5318 | 175.1072
0.66675 | 990864.3763 | 70.7994 | 170.925
0.68694 | 1023161.8857 | 65.42 | 157.9378
0.68688 | 1032454.4634 | 80.063 | 193.2891
0.69621 | 1055689.7353 | 77.4693 | 187.0274
0.71602 | 1086103.9403 | 75.7916 | 182.9772
0.68516 | 1030400.946 | 84.7748 | 204.6645
0.68409 | 1033501.0332 | 68.2797 | 164.8417
0.67853 | 1026309.4049 | 68.3605 | 165.0368
0.67331 | 1000017.072 | 65.2345 | 157.49
0.68691 | 1035379.4648 | 73.582 | 177.6427
0.70885 | 1097872.4864 | 68.0267 | 164.2311
0.68623 | 1015922.482 | 80.4176 | 194.1453
0.6696 | 986471.2065 | 68.4078 | 165.151
0.69189 | 1044320.4502 | 73.6839 | 177.8886
0.68545 | 1052512.8166 | 74.9596 | 180.9685
0.69432 | 1048674.9115 | 70.1326 | 169.3152
0.69186 | 1037493.8012 | 78.2607 | 188.9381
0.68797 | 1028102.2226 | 76.3475 | 184.3192
0.68569 | 1020349.0268 | 77.6599 | 187.4875
0.70085 | 1058472.9714 | 75.4621 | 182.1817
0.7023 | 1054492.933 | 71.5317 | 172.6929
0.72571 | 1128608.8446 | 74.0897 | 178.8683
0.70654 | 1100142.1582 | 81.2615 | 196.1827
0.69278 | 1054795.4358 | 70.6985 | 170.6813
0.68841 | 1026898.4409 | 71.0654 | 171.5672
0.69356 | 1057067.5563 | 66.9428 | 161.6143
0.67798 | 1019122.2975 | 76.7959 | 185.4017
0.69732 | 1065075.7752 | 63.0159 | 152.1338
0.69538 | 1064341.139 | 92.2381 | 222.6826
0.67998 | 1003223.4632 | 62.8664 | 151.773
0.68896 | 1027961.5104 | 69.2818 | 167.261
0.6814 | 1008976.359 | 66.9233 | 161.5671
0.6832 | 1030968.6797 | 68.2014 | 164.6527
0.6864 | 1040172.0137 | 70.3666 | 169.8799
0.68131 | 1034080.464 | 79.2106 | 191.2312
0.70641 | 1078583.74 | 67.0627 | 161.9036
0.71795 | 1108904.4112 | 69.8062 | 168.5271
0.71052 | 1079032.3052 | 62.6554 | 151.2636
0.71319 | 1094931.5866 | 71.2887 | 172.1061
0.68488 | 1032694.2254 | 84.9986 | 205.2049
0.6733 | 1008187.7717 | 69.5294 | 167.8589
0.70783 | 1078617.0073 | 79.527 | 191.9951
0.68648 | 1034134.9963 | 71.7678 | 173.2627
0.67152 | 992449.8075 | 70.4087 | 169.9817
0.6795 | 1013463.7107 | 79.4433 | 191.7932
0.68369 | 1028867.4321 | 57.7998 | 139.5411
0.6928 | 1038628.342 | 81.4347 | 196.6008
0.6841 | 1032320.1233 | 73.6661 | 177.8456
0.69603 | 1029608.6338 | 68.6886 | 165.8291
0.69031 | 1048267.9407 | 84.0585 | 202.9353
0.70053 | 1077991.7217 | 71.4543 | 172.506
0.69927 | 1061160.304 | 90.0114 | 217.3067
0.68945 | 1030816.1934 | 84.2814 | 203.4733
0.67365 | 1012038.7911 | 67.8978 | 163.9197
0.70582 | 1079051.4857 | 70.9885 | 171.3814
0.66084 | 978785.2234 | 78.8914 | 190.4608
0.70928 | 1101484.5885 | 68.2089 | 164.671
0.73738 | 1168937.1915 | 82.7319 | 199.7325
0.70051 | 1071949.8835 | 65.2047 | 157.418
0.68593 | 1032068.1818 | 66.086 | 159.5457
0.68299 | 1018803.1465 | 72.8199 | 175.8028
0.69553 | 1063917.5957 | 79.8303 | 192.7275
0.69126 | 1038682.739 | 78.4174 | 189.3163
0.69298 | 1045902.5442 | 75.4128 | 182.0626
0.69334 | 1053644.5506 | 68.6817 | 165.8124
0.69768 | 1075223.7225 | 73.4206 | 177.2529
0.69177 | 1027899.5414 | 67.156 | 162.129
0.68851 | 1028407.9366 | 67.507 | 162.9764
0.69056 | 1040906.1191 | 69.9921 | 168.9759
0.6785 | 1018545.9013 | 75.0076 | 181.0844
0.68135 | 1037452.5419 | 66.8774 | 161.4564
0.69607 | 1069570.611 | 59.0175 | 142.4808
0.66964 | 993524.9732 | 75.8002 | 182.9979
0.68073 | 1026790.0064 | 55.6605 | 134.3764
0.6956 | 1050137.0844 | 88.0893 | 212.6664
0.69529 | 1057095.2687 | 70.5969 | 170.436
0.69838 | 1055007.1536 | 53.3718 | 128.8509
0.70719 | 1074619.6781 | 84.1149 | 203.0714
0.70049 | 1061351.5253 | 75.7062 | 182.771
0.69042 | 1034617.3649 | 79.8185 | 192.6988
0.67747 | 1028210.4298 | 72.305 | 174.5596
0.67542 | 994932.5221 | 68.4005 | 165.1333
0.69753 | 1047587.6784 | 76.7287 | 185.2395
0.68632 | 1034168.9512 | 71.7864 | 173.3076
0.69026 | 1031609.7633 | 72.9588 | 176.138
0.69617 | 1065796.2806 | 69.192 | 167.0442
0.68664 | 1024359.7235 | 69.2722 | 167.238
0.70812 | 1086043.2733 | 63.2759 | 152.7616
0.69508 | 1041344.9187 | 78.1863 | 188.7585
0.69543 | 1049926.9055 | 70.1037 | 169.2454
0.69921 | 1059974.1769 | 88.455 | 213.5492
0.68568 | 1031004.0497 | 67.9576 | 164.0641
0.67189 | 997811.9559 | 72.4973 | 175.0239
0.68226 | 1009276.3684 | 68.6911 | 165.8351
0.68867 | 1034090.476 | 76.7146 | 185.2053
0.67901 | 1015941.9823 | 79.6956 | 192.4022
0.67063 | 984409.7535 | 71.6586 | 172.9991
0.69878 | 1066408.0918 | 70.8628 | 171.078
0.69422 | 1042471.3477 | 78.3685 | 189.1984
0.68593 | 1026109.0166 | 65.7046 | 158.6249
0.68822 | 1020432.8556 | 74.0856 | 178.8584
0.69341 | 1051050.852 | 79.1052 | 190.9768
0.68543 | 1036195.8284 | 68.3237 | 164.9481
0.69191 | 1050069.7458 | 77.5353 | 187.1868
0.68866 | 1044637.594 | 79.1651 | 191.1214
0.71914 | 1088587.1004 | 64.5014 | 155.7203
0.67401 | 995994.2791 | 68.6164 | 165.6546
0.71442 | 1101234.1027 | 73.435 | 177.2878
0.69598 | 1059591.051 | 79.4536 | 191.8179
0.68406 | 1022206.0261 | 74.6652 | 180.2577
0.68023 | 1005347.9292 | 70.4105 | 169.986
0.72294 | 1126313.4852 | 72.3429 | 174.6512
0.69872 | 1044541.1453 | 50.3584 | 121.5759
0.68601 | 1015002.0571 | 54.5985 | 131.8124
0.68917 | 1046235.761 | 68.6121 | 165.6443
0.68561 | 1031891.5152 | 64.1319 | 154.8281
0.69592 | 1057491.426 | 64.7487 | 156.3173
0.68447 | 1021078.0568 | 74.5386 | 179.9521
0.70347 | 1082745.3647 | 76.9059 | 185.6672
0.66034 | 981104.2719 | 59.22 | 142.9696
0.72373 | 1139887.8687 | 78.6972 | 189.9919
0.69357 | 1067456.4648 | 65.5809 | 158.3264
0.70764 | 1074338.1642 | 70.4959 | 170.1923
0.69157 | 1041060.2621 | 90.7176 | 219.0117
0.69202 | 1055154.3193 | 74.4336 | 179.6986
0.67527 | 995651.1966 | 56.6307 | 136.7186
0.67056 | 981534.9218 | 67.9686 | 164.0908
0.69109 | 1068297.8098 | 78.9751 | 190.6627
0.68643 | 1025826.1587 | 65.6385 | 158.4654
0.6841 | 1032920.3021 | 88.1011 | 212.695
0.68149 | 1031823.5227 | 72.1308 | 174.1392
0.70381 | 1069334.9223 | 56.0664 | 135.3563
0.70467 | 1069770.3398 | 74.1105 | 178.9185
0.69214 | 1036146.8076 | 76.4818 | 184.6434
0.68394 | 1031808.7019 | 71.8721 | 173.5147
0.70885 | 1082328.4556 | 86.0712 | 207.7942
0.69753 | 1057777.0792 | 83.9426 | 202.6555
0.68951 | 1048427.7468 | 64.7913 | 156.4199
0.67842 | 1004622.8356 | 61.0948 | 147.496
0.67636 | 1019836.7575 | 70.0148 | 169.0306
0.6725 | 999363.3548 | 74.2524 | 179.2611
0.69446 | 1038660.1386 | 74.7318 | 180.4185
0.69978 | 1051035.2672 | 67.8819 | 163.8813
0.67473 | 1000550.9762 | 85.0718 | 205.3816
0.69269 | 1041784.0641 | 69.8355 | 168.5979
0.70768 | 1092766.2484 | 82.2492 | 198.5673
0.68018 | 1025254.7089 | 80.287 | 193.83
0.71084 | 1107046.3163 | 64.2532 | 155.121
0.68932 | 1046725.251 | 68.6289 | 165.6847
0.72635 | 1137894.2082 | 68.8081 | 166.1176
0.69139 | 1042343.4344 | 84.9842 | 205.1701
0.6928 | 1042253.9483 | 69.8563 | 168.648
0.71341 | 1087564.5823 | 81.1994 | 196.0326
0.68571 | 1042734.9317 | 74.2351 | 179.2193
0.67383 | 1002882.7962 | 70.9055 | 171.181
0.69079 | 1062924.8962 | 85.1748 | 205.6301
0.67866 | 1018260.3536 | 71.4801 | 172.5683
0.70788 | 1088691.4635 | 80.2641 | 193.7747
0.7013 | 1090093.5215 | 74.0525 | 178.7786
0.70333 | 1068587.0205 | 62.6836 | 151.3317
0.69993 | 1062334.4716 | 60.496 | 146.0502
0.68044 | 1017145.8212 | 65.6566 | 158.509
0.68714 | 1040681.3697 | 79.0395 | 190.8181
0.70911 | 1096804.0278 | 71.8213 | 173.3919
0.68984 | 1033851.1511 | 71.3987 | 172.3716
0.68759 | 1036018.6697 | 62.1639 | 150.077
0.72333 | 1140135.5753 | 78.7579 | 190.1383
0.67665 | 995701.133 | 78.3799 | 189.2258
0.69654 | 1070654.9947 | 76.8032 | 185.4192
0.67945 | 1018309.9916 | 88.8955 | 214.6128
0.68265 | 1026131.4834 | 55.4689 | 133.9137
0.7064 | 1076772.4636 | 56.249 | 135.7971
0.67415 | 1007631.6685 | 72.9653 | 176.1539
0.6876 | 1028809.5489 | 67.6232 | 163.2568
0.67873 | 1003912.8147 | 75.5668 | 182.4345
0.67134 | 988034.405 | 78.3909 | 189.2525
0.70194 | 1071354.5927 | 62.9699 | 152.0229
0.68651 | 1033164.3897 | 83.3059 | 201.1182
0.70726 | 1088779.9328 | 59.1269 | 142.7451
0.67642 | 1006425.1475 | 76.0139 | 183.5138
0.7123 | 1095291.3652 | 73.2566 | 176.857
0.71132 | 1075634.5008 | 77.2561 | 186.5127
0.69286 | 1063910.307 | 63.3523 | 152.946
0.69167 | 1042526.392 | 84.9655 | 205.1249
0.67821 | 1017304.7007 | 84.5547 | 204.1331
0.70515 | 1097912.0971 | 76.008 | 183.4996
0.68446 | 1035730.5549 | 73.8705 | 178.3392
0.70368 | 1056292.6262 | 75.0045 | 181.0768
0.69448 | 1041614.7892 | 74.9412 | 180.9241
0.67747 | 996690.5908 | 71.8197 | 173.3881
0.6877 | 1040928.4065 | 75.631 | 182.5894
0.68523 | 1024305.5425 | 81.1994 | 196.0327
0.71946 | 1100446.9424 | 76.4259 | 184.5085
0.70443 | 1063488.406 | 66.3097 | 160.0859
0.69294 | 1041030.1046 | 83.1608 | 200.7679
0.71215 | 1082783.3066 | 66.7811 | 161.2239
0.71251 | 1097395.2369 | 67.6482 | 163.3171
0.7 | 1066394.7351 | 96.0146 | 231.7998
0.6914 | 1042523.2375 | 66.2277 | 159.8879
0.68485 | 1014265.3409 | 61.102 | 147.5134
0.6796 | 1018722.8547 | 74.8991 | 180.8225
0.69479 | 1045205.1378 | 80.1283 | 193.4469
0.68794 | 1040090.8256 | 60.2615 | 145.484
0.70783 | 1092116.6162 | 75.321 | 181.841
0.66606 | 981083.7407 | 58.7101 | 141.7388
0.67048 | 996877.8869 | 70.8322 | 171.004
0.68803 | 1029386.4948 | 69.4716 | 167.7194
0.6973 | 1049808.9879 | 64.1516 | 154.8755
0.67875 | 1006013.0342 | 83.5504 | 201.7086
0.66438 | 981526.8877 | 73.8177 | 178.2116
0.68795 | 1023315.5546 | 80.7992 | 195.0666
0.68112 | 1016758.205 | 82.6317 | 199.4906
0.69953 | 1058554.8774 | 85.3018 | 205.9367
0.70727 | 1084702.9063 | 80.038 | 193.2289
0.69579 | 1042831.6216 | 69.5946 | 168.0162
0.67671 | 1027816.8277 | 80.9004 | 195.3108
0.71025 | 1091063.4783 | 74.9912 | 181.0448
0.66908 | 982029.8106 | 66.3352 | 160.1473
0.68254 | 1026046.7681 | 71.6168 | 172.8984
0.67556 | 1022128.4147 | 68.3812 | 165.0867
0.71465 | 1105392.3471 | 73.6886 | 177.8999
0.69689 | 1057904.9491 | 79.2236 | 191.2628
0.68994 | 1038009.3308 | 73.445 | 177.312
0.70803 | 1076388.1621 | 68.7091 | 165.8783
0.66441 | 965288.8913 | 59.9411 | 144.7107
0.67566 | 995762.2304 | 80.0544 | 193.2684
0.68293 | 1042006.9111 | 72.6895 | 175.4881
0.68271 | 1009331.9946 | 70.4327 | 170.0396
0.69454 | 1052123.4369 | 79.0572 | 190.861
0.67537 | 1005008.7019 | 80.8324 | 195.1467
0.68715 | 1050316.7389 | 74.618 | 180.1437
0.66613 | 988172.0045 | 74.1229 | 178.9484
0.70612 | 1055597.6083 | 67.5645 | 163.1151
0.7123 | 1081962.6955 | 85.6252 | 206.7174
0.68797 | 1044870.1007 | 75.887 | 183.2074
0.65927 | 976489.0455 | 77.2996 | 186.6177
0.69901 | 1066123.0528 | 64.9892 | 156.8979
0.6885 | 1022440.3519 | 74.2012 | 179.1376
0.70329 | 1070221.8521 | 70.6608 | 170.5903
0.65959 | 969333.7357 | 70.0126 | 169.0255
0.70958 | 1086226.9997 | 78.2078 | 188.8105
0.6842 | 1024117.411 | 79.7267 | 192.4774
0.69267 | 1054755.2121 | 70.7066 | 170.7009
0.66714 | 995393.8272 | 81.0468 | 195.6643
0.69442 | 1050464.7492 | 66.4421 | 160.4053
0.6911 | 1053923.6757 | 76.1977 | 183.9575
0.70066 | 1049077.344 | 63.7998 | 154.0264
0.68157 | 1024913.4299 | 67.9629 | 164.0769
0.69021 | 1059403.514 | 84.9636 | 205.1203
0.69335 | 1063159.448 | 54.1737 | 130.787
0.70791 | 1088987.0199 | 85.1528 | 205.5771
0.70134 | 1068821.0195 | 70.764 | 170.8395
0.68076 | 1012612.6618 | 75.9936 | 183.4647
0.7002 | 1064457.7187 | 62.7205 | 151.4206
0.63866 | 917223.5121 | 73.3735 | 177.1393
0.68654 | 1036439.2701 | 83.567 | 201.7487
0.71147 | 1104624.674 | 75.7935 | 182.9817
0.68483 | 1011313.6534 | 68.5712 | 165.5456
0.65696 | 959526.5543 | 77.7964 | 187.817
0.70462 | 1080976.5034 | 78.3771 | 189.2191
0.70717 | 1068401.0225 | 64.3763 | 155.4182
0.69421 | 1049326.4119 | 77.2242 | 186.4357
0.71436 | 1094812.4649 | 78.2989 | 189.0303
0.70535 | 1076168.6859 | 79.6248 | 192.2313
0.68037 | 1018267.9758 | 69.4955 | 167.777
0.69534 | 1036833.576 | 71.0376 | 171.4999
0.6932 | 1038545.5087 | 85.9495 | 207.5003
0.71162 | 1104574.5021 | 77.5574 | 187.2401
0.70832 | 1088206.6554 | 74.9991 | 181.0638
0.69929 | 1055051.8671 | 90.9224 | 219.5062
0.67456 | 997048.911 | 66.8024 | 161.2752
0.69931 | 1057227.361 | 72.0237 | 173.8807
0.68774 | 1033792.0144 | 78.4482 | 189.3907
0.70225 | 1062774.8473 | 77.3182 | 186.6627
0.69215 | 1040121.6712 | 87.684 | 211.688
0.6675 | 980450.5381 | 75.0424 | 181.1684
0.68535 | 1028475.9107 | 69.6821 | 168.2275
0.69126 | 1052810.0691 | 72.5576 | 175.1697
0.68014 | 1021630.7042 | 61.9725 | 149.6148
0.6866 | 1017241.6026 | 83.5169 | 201.6276
0.68431 | 1024406.4873 | 63.788 | 153.9979
0.68721 | 1067433.9969 | 75.8268 | 183.062
0.7109 | 1092946.4292 | 73.2431 | 176.8245
0.69237 | 1046131.5199 | 64.3818 | 155.4314
0.70341 | 1077013.3067 | 69.2001 | 167.0639
0.70481 | 1081150.4591 | 76.6462 | 185.0404
0.68824 | 1034247.688 | 57.7173 | 139.3419
0.68633 | 1046927.9735 | 89.0066 | 214.8808
0.68124 | 1020731.4151 | 82.9405 | 200.236
0.68373 | 1025684.0603 | 61.8914 | 149.4191
0.68376 | 1019677.9393 | 81.0062 | 195.5662
0.65904 | 966235.4263 | 68.0779 | 164.3546
0.64553 | 951147.0286 | 58.7282 | 141.7825
0.69614 | 1058728.5459 | 70.6196 | 170.4908
0.68206 | 1029970.4989 | 80.4338 | 194.1844
0.68455 | 1038813.1906 | 77.5289 | 187.1713
0.71653 | 1102088.5271 | 69.5144 | 167.8227
0.6943 | 1045324.8329 | 78.5742 | 189.695
0.71181 | 1105115.0075 | 79.3189 | 191.4928
0.70732 | 1078010.5924 | 79.5393 | 192.0247
0.67591 | 1017297.6157 | 73.4738 | 177.3814
0.70064 | 1057619.9592 | 80.4037 | 194.1118
0.69803 | 1072099.1104 | 78.4678 | 189.4379
0.67266 | 1000971.6719 | 63.4081 | 153.0808
0.68999 | 1028400.0256 | 79.3884 | 191.6606
0.6811 | 1026157.4797 | 69.9584 | 168.8946
0.68687 | 1035070.2727 | 76.9879 | 185.8653
0.70097 | 1080147.1225 | 78.725 | 190.059
0.70904 | 1088945.5753 | 71.3134 | 172.1659
0.68022 | 1027311.0648 | 64.3528 | 155.3615
0.67207 | 1006916.0736 | 76.5442 | 184.7941
0.68745 | 1056950.3504 | 83.5157 | 201.6247
0.67569 | 1002093.8936 | 80.5603 | 194.4899
0.69164 | 1052426.9393 | 76.0301 | 183.5528
0.69993 | 1053515.4586 | 68.0271 | 164.2319
0.69498 | 1031692.9883 | 68.9309 | 166.414
0.68299 | 1029475.4087 | 78.6644 | 189.9127
0.69024 | 1052638.7964 | 72.6537 | 175.4016
0.68076 | 1008426.1378 | 77.3079 | 186.6377
0.69058 | 1037465.0205 | 69.7581 | 168.4109
0.69105 | 1054651.9157 | 53.9198 | 130.1738
0.68305 | 1016891.0222 | 74.3868 | 179.5856
0.68873 | 1030609.4029 | 74.7686 | 180.5074
0.69221 | 1061421.3016 | 76.6177 | 184.9714
0.65035 | 959099.3324 | 66.6516 | 160.9111
0.68537 | 1031135.5846 | 80.1365 | 193.4667
0.67537 | 1005520.0844 | 72.6808 | 175.467
0.70472 | 1096303.137 | 81.1868 | 196.0024
0.70308 | 1085934.0432 | 77.175 | 186.317
0.6741 | 1018711.3245 | 71.9954 | 173.8123
0.67942 | 1009134.8149 | 75.882 | 183.1954
0.67575 | 1002151.7524 | 74.1012 | 178.896
0.70879 | 1070290.3876 | 73.7151 | 177.9639
0.69624 | 1040358.29 | 76.3407 | 184.3027
0.67426 | 987461.1791 | 60.4218 | 145.871
0.69603 | 1052844.6374 | 83.732 | 202.1468
0.68978 | 1032147.8346 | 69.4816 | 167.7434
0.67363 | 1001777.1253 | 65.318 | 157.6915
0.69182 | 1039294.7806 | 78.4007 | 189.276
0.67825 | 1024775.4819 | 68.2867 | 164.8587
0.65561 | 948002.7262 | 89.822 | 216.8496
0.69704 | 1040251.1306 | 71.0984 | 171.6467
0.6841 | 1013423.8594 | 66.9525 | 161.6377
0.68248 | 1033858.8232 | 60.3799 | 145.77
0.69816 | 1052441.3354 | 57.8981 | 139.7784
0.66964 | 996954.395 | 61.4235 | 148.2894
0.70531 | 1080989.4995 | 51.6287 | 124.6426
0.69876 | 1073415.7677 | 71.3313 | 172.2091
0.68377 | 1039919.5433 | 70.6198 | 170.4913
0.70223 | 1074036.5537 | 72.8644 | 175.9102
0.6892 | 1041409.0227 | 68.292 | 164.8715
0.70411 | 1054509.7748 | 59.6933 | 144.1123
0.68018 | 1016975.3228 | 59.5945 | 143.8739
0.70289 | 1068512.7908 | 66.7692 | 161.1952
0.71163 | 1077145.6362 | 70.2844 | 169.6816
0.67382 | 1006023.4256 | 63.618 | 153.5875
0.70164 | 1095448.7419 | 72.4958 | 175.0203
0.70362 | 1082403.2104 | 86.7169 | 209.3531
0.69324 | 1036990.2354 | 67.6059 | 163.2152
0.68539 | 1027704.8147 | 65.8585 | 158.9964
0.69496 | 1066294.6001 | 78.3872 | 189.2434
0.6791 | 1020525.4742 | 73.3609 | 177.1089
0.69177 | 1048793.9523 | 60.4293 | 145.8892
0.68872 | 1042545.2444 | 80.6288 | 194.6551
0.71027 | 1068656.0698 | 76.8677 | 185.5751
0.6711 | 993914.1004 | 87.8455 | 212.0778
0.68371 | 1018811.2053 | 83.9996 | 202.7931
0.68861 | 1052097.0008 | 76.3412 | 184.3039
0.67857 | 1033238.7465 | 77.7918 | 187.8061
0.67925 | 1007950.7565 | 70.6952 | 170.6732
0.67895 | 1029223.2225 | 83.9128 | 202.5835
0.68488 | 1034285.7213 | 73.7227 | 177.9824
0.67671 | 986379.8295 | 67.6827 | 163.4006
0.67631 | 1019437.0844 | 53.2553 | 128.5696
0.69758 | 1070269.7203 | 76.6611 | 185.0764
0.69458 | 1036212.9221 | 72.8284 | 175.8233
0.67882 | 1013010.8419 | 79.4879 | 191.9007
0.68897 | 1044586.7358 | 66.5924 | 160.7682
0.70093 | 1076107.2855 | 81.7749 | 197.4222
0.6983 | 1076722.7641 | 56.8989 | 137.3661
0.66982 | 1005008.6478 | 74.9704 | 180.9944
0.68447 | 1023153.3976 | 71.2138 | 171.9253
0.70211 | 1063826.3389 | 70.3561 | 169.8546
0.69409 | 1044631.3028 | 71.7005 | 173.1002
0.68747 | 1027001.0147 | 65.6145 | 158.4074
0.69957 | 1061339.2421 | 66.6351 | 160.8715
0.70165 | 1066950.6738 | 80.9961 | 195.542
0.68409 | 1017899.5716 | 67.2491 | 162.3536
0.70728 | 1082839.109 | 79.9 | 192.8957
0.70202 | 1084503.699 | 78.9247 | 190.541
0.69944 | 1065190.3144 | 58.4928 | 141.2142
0.68695 | 1021408.3181 | 80.7668 | 194.9882
0.70345 | 1068660.235 | 80.9187 | 195.3549
0.67387 | 1017864.0993 | 80.4554 | 194.2365
0.66087 | 980433.3716 | 79.0203 | 190.7719
0.70981 | 1082308.6362 | 76.1879 | 183.9338
0.66649 | 996617.468 | 91.5108 | 220.9267
0.6765 | 1014860.3686 | 65.8901 | 159.0729
0.69093 | 1059894.3035 | 63.2893 | 152.7938
0.70463 | 1062030.9319 | 77.0027 | 185.9011
0.68025 | 1011227.3984 | 67.1853 | 162.1996
0.7056 | 1089108.9444 | 77.1237 | 186.1931
0.71444 | 1107535.8883 | 73.8531 | 178.297
0.6915 | 1060222.4063 | 71.1942 | 171.8779
0.68041 | 1027150.4453 | 83.5635 | 201.7401
0.70993 | 1074506.5009 | 64.4925 | 155.6987
0.6931 | 1054662.2037 | 66.8345 | 161.3526
0.68554 | 1040767.6195 | 77.5628 | 187.2531
0.68507 | 1018768.1649 | 79.2098 | 191.2294
0.68281 | 1024632.3241 | 79.9297 | 192.9675
0.67476 | 994549.0907 | 66.6155 | 160.824
0.69795 | 1064831.0355 | 82.8737 | 200.0747
0.67082 | 991540.7966 | 73.8551 | 178.3019
0.68137 | 1019756.383 | 79.7178 | 192.4558
0.68711 | 1044758.6374 | 60.6045 | 146.3122
0.68945 | 1035372.6802 | 69.7317 | 168.3473
0.69766 | 1057237.9329 | 73.978 | 178.5987
0.70016 | 1064208.0124 | 75.9465 | 183.351
0.68812 | 1039003.5548 | 62.2949 | 150.3932
0.70485 | 1074272.1207 | 74.4032 | 179.6251
0.69712 | 1051449.8389 | 77.0573 | 186.0328
0.69534 | 1063143.1935 | 70.9961 | 171.3998
0.67181 | 992926.5335 | 72.1515 | 174.189
0.69998 | 1065295.2842 | 68.7828 | 166.0564
0.68664 | 1039145.0047 | 82.2068 | 198.4647
0.69684 | 1052894.6205 | 75.497 | 182.2658
0.6598 | 954969.6033 | 70.6565 | 170.5799
0.67318 | 1012963.2836 | 86.3244 | 208.4055
0.69206 | 1050380.1524 | 77.5411 | 187.2009
0.7135 | 1100751.2365 | 69.7256 | 168.3325
0.67341 | 1000839.6744 | 75.9305 | 183.3125
0.68437 | 1043414.4117 | 87.466 | 211.1616
0.67645 | 1013820.993 | 55.3054 | 133.5192
0.70816 | 1062702.9309 | 74.9816 | 181.0216
0.68454 | 1027374.789 | 63.951 | 154.3914
0.67883 | 1014301.1936 | 79.5873 | 192.1407
0.68904 | 1038857.787 | 77.0268 | 185.9592
0.71073 | 1075024.7654 | 85.2639 | 205.8454
0.70538 | 1068894.4324 | 77.4178 | 186.9031
0.69498 | 1037778.9259 | 55.1026 | 133.0294
0.69099 | 1062119.3638 | 75.0374 | 181.1564
0.69065 | 1043066.5079 | 83.5999 | 201.8281
0.68894 | 1040328.6875 | 68.2477 | 164.7644
0.69761 | 1058486.2237 | 62.6002 | 151.1302
0.7016 | 1055603.3811 | 78.5205 | 189.5654
0.67597 | 997934.027 | 72.0685 | 173.9887
0.66983 | 983968.1888 | 58.5514 | 141.3555
0.70343 | 1094137.9875 | 70.2987 | 169.7162
0.7173 | 1087365.9924 | 70.3904 | 169.9375
0.71065 | 1091818.5983 | 78.4924 | 189.4975
0.71031 | 1086707.5211 | 74.6858 | 180.3076
0.71557 | 1116933.1861 | 76.1986 | 183.9598
0.68743 | 1022571.5503 | 79.2323 | 191.2837
0.72552 | 1122097.7803 | 69.9 | 168.7536
0.67775 | 1020516.2323 | 76.9273 | 185.7188
0.69446 | 1047321.0022 | 70.0279 | 169.0624
0.69701 | 1061562.7522 | 72.5576 | 175.1694
0.67794 | 1009588.5676 | 72.6991 | 175.5111
0.69279 | 1055509.6669 | 77.8383 | 187.9184
0.66997 | 988646.0749 | 84.657 | 204.38
0.69067 | 1035458.3533 | 67.7579 | 163.582
0.69122 | 1035175.4691 | 61.7869 | 149.1668
0.66776 | 993908.4902 | 70.9094 | 171.1905
0.66742 | 979954.9791 | 71.6769 | 173.0432
0.66077 | 970940.6476 | 66.411 | 160.3303
0.68809 | 1048245.7267 | 84.3281 | 203.586
0.70405 | 1076744.4723 | 81.4604 | 196.6629
0.69076 | 1046886.673 | 73.9853 | 178.6164
0.6975 | 1054577.7305 | 72.0299 | 173.8955
0.69591 | 1052988.4843 | 58.1162 | 140.305
0.68505 | 1040859.0943 | 74.1723 | 179.0677
0.71829 | 1120684.2372 | 77.9574 | 188.2058
0.69688 | 1056521.8662 | 83.9218 | 202.6051
0.70496 | 1079177.6836 | 74.7311 | 180.4167
0.69587 | 1053534.8547 | 64.5115 | 155.7446
0.67989 | 1021715.8769 | 66.0045 | 159.349
0.67181 | 990381.5177 | 62.8164 | 151.6523
0.67853 | 1018307.377 | 84.5984 | 204.2385
0.68382 | 1041552.8083 | 81.1295 | 195.8639
0.69829 | 1063296.9774 | 81.1069 | 195.8093
0.68885 | 1026702.2893 | 80.9619 | 195.4592
0.67259 | 996742.7947 | 61.1409 | 147.6072
0.67929 | 1029115.8673 | 77.3998 | 186.8597
0.68664 | 1032937.908 | 76.856 | 185.5468
0.69164 | 1035460.6047 | 65.7563 | 158.7497
0.69551 | 1053341.6988 | 73.9331 | 178.4904
0.68453 | 1040406.0579 | 77.65 | 187.4637
0.67851 | 1006877.2282 | 82.144 | 198.3132
0.68616 | 1044728.6191 | 62.8176 | 151.6552
0.71491 | 1101817.6292 | 85.3122 | 205.9618
0.70123 | 1093568.8089 | 87.9094 | 212.2321
0.68041 | 1012428.5674 | 83.8773 | 202.4976
0.69256 | 1039491.1894 | 50.5609 | 122.0648
0.66869 | 967495.8708 | 63.549 | 153.4208
0.68539 | 1028758.543 | 76.6848 | 185.1334
0.70499 | 1076292.4373 | 70.8904 | 171.1446
0.70823 | 1102838.7654 | 62.0666 | 149.842
0.68683 | 1041765.9218 | 79.1581 | 191.1045
0.69275 | 1060910.3542 | 58.82 | 142.0041
0.68626 | 1041847.8767 | 72.3256 | 174.6095
0.68196 | 1020582.8928 | 85.8548 | 207.2717
0.66452 | 1016294.2566 | 72.8598 | 175.8991
0.70608 | 1080350.6249 | 84.697 | 204.4765
0.70686 | 1087695.6675 | 73.3078 | 176.9806
0.6812 | 1037382.7035 | 70.7135 | 170.7176
0.69066 | 1051796.1531 | 62.8299 | 151.6849
0.71897 | 1116112.3186 | 72.0832 | 174.0242
0.70032 | 1072766.8582 | 71.1295 | 171.7217
0.6755 | 999351.3874 | 80.3625 | 194.0122
0.68323 | 1020697.2218 | 84.2141 | 203.3108
0.68793 | 1039723.691 | 55.7716 | 134.6444
0.70678 | 1073429.681 | 62.6602 | 151.2752
0.67961 | 1008752.9195 | 75.833 | 183.077
0.67697 | 992188.6956 | 71.5484 | 172.7332
0.67678 | 1008136.0927 | 67.3235 | 162.5333
0.69064 | 1043986.7856 | 68.3186 | 164.9357
0.70701 | 1095172.48 | 68.0362 | 164.254
0.69815 | 1053342.5391 | 73.6541 | 177.8167
0.6611 | 966465.3967 | 53.9461 | 130.2374
0.6996 | 1073123.637 | 80.0523 | 193.2633
0.68326 | 1033705.7345 | 68.376 | 165.0742
0.68591 | 1039845.0203 | 72.2232 | 174.3621
0.66371 | 969144.928 | 71.2831 | 172.0926
0.70073 | 1052374.7843 | 65.9628 | 159.2483
0.68411 | 1027435.9928 | 83.1348 | 200.705
0.67241 | 997006.4015 | 75.7521 | 182.8818
0.69037 | 1043781.8063 | 82.1264 | 198.2708
0.69199 | 1049501.4427 | 74.0385 | 178.7447
0.68509 | 1033801.3447 | 71.89 | 173.5578
0.69139 | 1049746.8799 | 68.6212 | 165.6662
0.68208 | 1018149.4182 | 72.8697 | 175.9231
0.69138 | 1042743.7119 | 75.3592 | 181.9332
0.7085 | 1082870.8437 | 72.8142 | 175.789
0.67687 | 1016158.8765 | 63.9694 | 154.4359
0.68786 | 1026006.1264 | 67.4133 | 162.7501
0.67114 | 1011218.5667 | 55.6 | 134.2303
0.68151 | 1019769.5411 | 69.4808 | 167.7414
0.70196 | 1070964.5446 | 76.2324 | 184.0413
0.69726 | 1057274.3879 | 86.9899 | 210.0123
0.66956 | 998345.5039 | 65.4895 | 158.1056
0.6943 | 1052625.5061 | 60.4285 | 145.8873
0.69103 | 1039398.1845 | 67.98 | 164.1182
0.70639 | 1090622.2002 | 65.7882 | 158.8267
0.6751 | 998305.8601 | 75.5488 | 182.3909
0.69636 | 1065800.1895 | 51.5577 | 124.4712
0.70306 | 1060288.8687 | 89.4518 | 215.9558
0.69215 | 1059493.4222 | 83.1486 | 200.7386
0.69893 | 1057126.6744 | 73.8353 | 178.2541
0.69962 | 1053967.8933 | 84.9214 | 205.0185
0.70915 | 1086767.2748 | 76.5726 | 184.8625
0.71077 | 1089523.145 | 93.809 | 226.475
0.68503 | 1032105.1079 | 68.0947 | 164.395
0.69509 | 1042203.1945 | 94.4993 | 228.1414
0.69454 | 1055714.3932 | 66.7624 | 161.1786
0.68189 | 1018160.2652 | 69.3333 | 167.3854
0.69941 | 1065469.1311 | 85.8737 | 207.3175
0.69759 | 1050160.2716 | 69.903 | 168.7607
0.68542 | 1029693.1381 | 65.9832 | 159.2976
0.7072 | 1078145.3965 | 60.7647 | 146.6989
0.6967 | 1043091.3957 | 63.7957 | 154.0165
0.69818 | 1067727.4322 | 71.0119 | 171.4379
0.69889 | 1055548.3178 | 71.9529 | 173.7097
0.67383 | 1003990.2766 | 72.7366 | 175.6018
0.68228 | 1014333.1341 | 79.1882 | 191.1773
0.69376 | 1050301.9804 | 55.8659 | 134.8721
0.68024 | 1038218.037 | 82.9061 | 200.1531
0.68555 | 1040147.6036 | 75.8116 | 183.0253
0.65393 | 967327.8829 | 77.1992 | 186.3754
0.68435 | 1037749.4353 | 79.1384 | 191.057
0.68808 | 1035779.6049 | 83.3269 | 201.169
0.67567 | 1013507.2396 | 61.8864 | 149.4071
0.68316 | 1022319.7659 | 74.7719 | 180.5154
0.69995 | 1054247.1792 | 83.3739 | 201.2823
0.68682 | 1035267.3617 | 78.6651 | 189.9143
0.68228 | 1008072.7763 | 60.9581 | 147.1659
0.68886 | 1036180.9117 | 75.1591 | 181.45
0.68706 | 1045526.8475 | 81.0546 | 195.6832
0.66931 | 998923.5869 | 67.4534 | 162.8468
0.68897 | 1036575.6786 | 75.1759 | 181.4907
0.66847 | 1003238.9535 | 81.5231 | 196.8142
0.66644 | 987052.0212 | 72.3018 | 174.5521
0.68725 | 1040694.6339 | 66.5103 | 160.5702
0.68038 | 1013250.3419 | 72.7003 | 175.5142
0.68391 | 1054920.6774 | 77.0717 | 186.0676
0.69296 | 1056636.4995 | 72.6434 | 175.3766
0.69732 | 1058163.0091 | 77.5265 | 187.1656
0.70538 | 1070679.0872 | 72.313 | 174.5791
0.6798 | 1020195.6339 | 81.5195 | 196.8055
0.66444 | 977955.8324 | 78.3974 | 189.2681
0.6877 | 1034795.9901 | 80.1488 | 193.4964
0.70178 | 1076478.05 | 74.8816 | 180.7801
0.69616 | 1042597.7292 | 65.7792 | 158.8049
0.67603 | 1000893.848 | 77.2676 | 186.5404
0.68412 | 1029751.2692 | 59.995 | 144.8407
0.68963 | 1030467.1576 | 72.9448 | 176.1043
0.67251 | 990843.5843 | 58.7255 | 141.776
0.72243 | 1132264.1244 | 76.5117 | 184.7156
0.68175 | 1009221.5761 | 75.6299 | 182.5868
0.6868 | 1043282.1353 | 72.3681 | 174.712
0.66864 | 993574.7252 | 70.7926 | 170.9084
0.67632 | 1013576.1086 | 63.859 | 154.1692
0.69473 | 1042359.7372 | 69.6687 | 168.1951
0.69024 | 1052402.0972 | 72.6369 | 175.361
0.67379 | 1002548.7932 | 80.8235 | 195.1252
0.68954 | 1056538.9646 | 84.8165 | 204.765
0.68624 | 1037619.0667 | 86.7438 | 209.4182
0.70697 | 1083264.0946 | 63.2173 | 152.6201
0.67851 | 999657.8148 | 66.5877 | 160.7569
0.68914 | 1042728.8882 | 72.0844 | 174.0271
0.66743 | 993905.3838 | 76.0804 | 183.6744
0.66425 | 984750.7891 | 75.7403 | 182.8533
0.68007 | 1007309.117 | 78.8953 | 190.47
0.69884 | 1079771.5443 | 75.4271 | 182.0972
0.69764 | 1054918.9119 | 73.8174 | 178.2109
0.71964 | 1127203.4669 | 89.8595 | 216.9399
0.70791 | 1078062.2638 | 68.8273 | 166.1639
0.68071 | 997334.4012 | 78.0399 | 188.4049
0.67329 | 991848.0374 | 75.2615 | 181.6973
0.69399 | 1043500.612 | 64.0706 | 154.6801
0.71076 | 1085898.9804 | 86.1225 | 207.9182
0.70492 | 1066464.3464 | 82.18 | 198.4
0.67117 | 985399.0101 | 76.6216 | 184.981
0.72265 | 1102906.9716 | 65.0318 | 157.0007
0.68852 | 1020858.5883 | 63.1777 | 152.5245
0.68672 | 1051641.3749 | 78.2382 | 188.8836
0.6863 | 1034806.6776 | 62.2075 | 150.1823
0.69722 | 1058993.8971 | 64.7201 | 156.2481
0.6919 | 1048700.8774 | 72.2083 | 174.3262
0.66814 | 982910.9117 | 49.5557 | 119.6382
0.70105 | 1068068.8726 | 79.5098 | 191.9537
0.6774 | 1006441.3779 | 72.5304 | 175.1038
0.68461 | 1032338.6876 | 62.213 | 150.1954
0.65785 | 969866.8354 | 71.9709 | 173.7531
0.70832 | 1065587.7987 | 59.9618 | 144.7606
0.68962 | 1033298.5123 | 59.7217 | 144.1811
0.69939 | 1054892.6163 | 73.6312 | 177.7615
0.69833 | 1056800.7578 | 73.8765 | 178.3538
0.69326 | 1042668.9562 | 69.8377 | 168.6032
0.69833 | 1067109.68 | 76.3528 | 184.332
0.68977 | 1040432.0484 | 71.86 | 173.4855
0.69295 | 1039780.5366 | 74.9747 | 181.005
0.66445 | 996199.7921 | 69.6176 | 168.0718
0.6775 | 993590.1907 | 74.9251 | 180.8851
0.69704 | 1065584.1761 | 76.3846 | 184.4088
0.7023 | 1065080.2997 | 64.6212 | 156.0095
0.71021 | 1088468.1035 | 78.2991 | 189.0308
0.70817 | 1063235.7896 | 79.9712 | 193.0677
0.67271 | 1001836.9305 | 59.3592 | 143.3059
0.70616 | 1090440.2654 | 77.1569 | 186.2732
0.68474 | 1038844.5888 | 80.8096 | 195.0917
0.68171 | 1013032.041 | 63.3199 | 152.8677
0.6785 | 1018962.1531 | 83.1307 | 200.6952
0.67243 | 1006503.0123 | 71.3095 | 172.1564
0.65876 | 972365.2696 | 77.0325 | 185.9729
0.70029 | 1068608.3816 | 84.4689 | 203.9259
0.7011 | 1080884.4873 | 92.908 | 224.2996
0.69548 | 1076588.5842 | 77.3472 | 186.7327
0.69397 | 1050499.896 | 77.3371 | 186.7084
0.6983 | 1054864.6041 | 68.2733 | 164.8264
0.69654 | 1059653.7829 | 68.7567 | 165.9935
0.6907 | 1038628.1988 | 81.6827 | 197.1994
0.68699 | 1028296.1945 | 71.3089 | 172.155
0.68488 | 1045012.446 | 81.2727 | 196.2098
0.69866 | 1059788.9997 | 60.4606 | 145.9648
0.70704 | 1080635.1317 | 67.1303 | 162.0669
0.68601 | 1025979.9193 | 65.6892 | 158.5877
0.70474 | 1081736.6521 | 74.9323 | 180.9025
0.68155 | 1025650.4918 | 62.0876 | 149.8927
0.70334 | 1088244.3903 | 71.845 | 173.4492
0.69945 | 1054936.5223 | 90.8912 | 219.4308
0.67193 | 1008386.3387 | 79.9365 | 192.9837
0.69652 | 1049281.9984 | 78.6201 | 189.8057
0.68731 | 1021633.7924 | 63.3375 | 152.9102
0.65776 | 952373.3894 | 67.2318 | 162.312
0.68524 | 1027540.4015 | 76.6106 | 184.9544
0.70539 | 1070558.898 | 56.0044 | 135.2066
0.67317 | 1002514.9922 | 70.9489 | 171.2858
0.70113 | 1063515.1638 | 85.506 | 206.4298
0.6987 | 1071499.6957 | 63.2708 | 152.7492
0.67553 | 1009860.0823 | 82.7501 | 199.7765
0.68446 | 1033074.3373 | 66.2923 | 160.0437
0.71237 | 1100394.264 | 67.8461 | 163.795
0.67435 | 1004180.7353 | 65.4049 | 157.9013
0.69079 | 1036864.7386 | 84.6114 | 204.27
0.69923 | 1059087.9461 | 75.6812 | 182.7105
0.68799 | 1031332.7807 | 78.2327 | 188.8705
0.71337 | 1108285.2342 | 72.1393 | 174.1596
0.68151 | 1023109.3299 | 81.5475 | 196.873
0.70436 | 1072337.9929 | 76.0695 | 183.6481
0.68246 | 1022397.1725 | 74.8546 | 180.7149
0.69588 | 1068432.2347 | 80.1286 | 193.4474
0.71025 | 1062603.899 | 78.0794 | 188.5003
0.68816 | 1037610.8239 | 77.0323 | 185.9725
0.68781 | 1039326.5802 | 64.3871 | 155.4442
0.69467 | 1062667.5907 | 83.0957 | 200.6107
0.70916 | 1089188.001 | 69.4155 | 167.5838
0.69569 | 1065079.6211 | 95.1016 | 229.5955
0.70073 | 1068689.024 | 72.657 | 175.4094
0.68801 | 1037854.2787 | 64.2774 | 155.1794
0.69898 | 1067138.8274 | 70.8918 | 171.1478
0.68838 | 1040204.7293 | 75.503 | 182.2803
0.66779 | 986960.4835 | 70.4109 | 169.9869
0.6982 | 1080910.8394 | 82.4613 | 199.0792
0.67441 | 1008245.9697 | 76.3789 | 184.395
0.69022 | 1048638.0717 | 66.7296 | 161.0995
0.71459 | 1119089.0424 | 76.4511 | 184.5693
0.67783 | 1009002.093 | 79.289 | 191.4206
0.67598 | 1011833.1395 | 82.8564 | 200.0331
0.70348 | 1076724.0189 | 72.9174 | 176.0382
0.6849 | 1026956.3516 | 84.5235 | 204.0578
0.69028 | 1050344.5572 | 58.4426 | 141.0929
0.68207 | 1038956.041 | 77.8218 | 187.8784
0.69777 | 1064504.1068 | 79.6177 | 192.2142
0.70079 | 1074820.0128 | 73.0675 | 176.4007
0.70466 | 1080821.4253 | 54.19 | 130.8262
0.67331 | 999961.244 | 75.8752 | 183.1789
0.70103 | 1078066.4474 | 71.4085 | 172.3953
0.68497 | 1027590.7424 | 63.9242 | 154.3267
0.70802 | 1070520.1904 | 64.4271 | 155.5407
0.66054 | 963111.246 | 76.0942 | 183.7078
0.70286 | 1076486.8298 | 67.2705 | 162.4054
0.70269 | 1077881.8379 | 74.8826 | 180.7825
0.67402 | 991548.9784 | 85.8375 | 207.2302
0.69272 | 1051945.0877 | 79.2521 | 191.3314
0.70959 | 1081421.5546 | 74.3975 | 179.6114
0.68589 | 1049448.483 | 74.693 | 180.3248
0.70257 | 1069859.4223 | 66.8838 | 161.4717
0.65385 | 958961.4101 | 69.8716 | 168.685
0.7145 | 1100700.5648 | 65.6423 | 158.4745
0.67763 | 1023683.3447 | 62.327 | 150.4706
0.66904 | 994084.6222 | 61.3019 | 147.9959
0.6834 | 1012324.6881 | 74.015 | 178.6879
0.70319 | 1061470.8094 | 66.3011 | 160.065
0.67447 | 1005821.4761 | 63.5443 | 153.4095
0.68396 | 1040542.0746 | 82.6395 | 199.5094
0.67912 | 1024521.5817 | 64.2822 | 155.191
0.67691 | 1016324.6888 | 76.7067 | 185.1863
0.70168 | 1080030.015 | 73.3287 | 177.0311
0.6843 | 1022884.3556 | 74.6884 | 180.3139
0.6767 | 1019861.5915 | 55.6139 | 134.2639
0.69323 | 1041858.9589 | 35.7981 | 86.4242
0.70362 | 1061604.1256 | 80.3644 | 194.0167
0.69045 | 1047759.7675 | 74.0807 | 178.8466
0.66813 | 979474.7338 | 58.4316 | 141.0663
0.68468 | 1027895.1507 | 69.7103 | 168.2956
0.69417 | 1036743.3902 | 72.9087 | 176.0173
0.69884 | 1081096.7972 | 65.9182 | 159.1405
0.67501 | 1006309.0858 | 72.7765 | 175.6981
0.68126 | 1026162.6236 | 73.5317 | 177.5213
0.69719 | 1055080.8199 | 75.6158 | 182.5526
0.69792 | 1062542.0649 | 81.092 | 195.7735
0.70791 | 1079620.6696 | 72.6554 | 175.4057
0.68369 | 1021892.6642 | 82.7367 | 199.744
0.71574 | 1102394.2458 | 75.1892 | 181.5229
0.68185 | 1021710.3115 | 71.3869 | 172.3432
0.71077 | 1103949.1137 | 79.3508 | 191.5698
0.69224 | 1049861.7207 | 75.7792 | 182.9473
0.67565 | 1004110.1357 | 52.5059 | 126.7604
0.71184 | 1114091.2689 | 92.8323 | 224.117
0.72226 | 1119599.357 | 87.3822 | 210.9592
0.67379 | 994274.8215 | 66.6926 | 161.0102
0.69933 | 1062335.4048 | 75.9022 | 183.2441
0.6947 | 1044469.758 | 78.4645 | 189.43
0.69564 | 1065617.267 | 86.3514 | 208.4708
0.69578 | 1060386.6767 | 76.1495 | 183.8411
0.69867 | 1076405.0459 | 56.8518 | 137.2523
0.67702 | 1007817.0435 | 76.0521 | 183.6061
0.67665 | 1021932.4535 | 62.3105 | 150.4308
0.70691 | 1073202.4463 | 77.5619 | 187.2509
0.68815 | 1024084.8051 | 74.3584 | 179.5171
0.71642 | 1091788.976 | 77.8571 | 187.9638
0.69339 | 1030006.4236 | 65.2448 | 157.5149
0.70517 | 1045724.0627 | 72.3933 | 174.773
0.68226 | 1026391.5224 | 75.1697 | 181.4757
0.6894 | 1050825.2469 | 68.8898 | 166.3146
0.68091 | 1010908.3423 | 65.2086 | 157.4274
0.69959 | 1073807.6771 | 67.416 | 162.7566
0.68099 | 1015977.9946 | 81.0401 | 195.6482
0.71207 | 1106263.3639 | 74.0137 | 178.6849
0.68898 | 1037851.3171 | 76.9593 | 185.7963
0.72899 | 1146055.8667 | 76.7468 | 185.2832
0.69212 | 1041091.8026 | 80.1215 | 193.4304
0.70121 | 1070304.9311 | 74.5127 | 179.8896
0.6843 | 1023129.4653 | 61.6856 | 148.9222
0.68518 | 1018070.4188 | 65.2614 | 157.555
0.69697 | 1065194.6719 | 56.397 | 136.1543
0.69058 | 1059019.5342 | 81.6822 | 197.1983
0.68915 | 1027744.784 | 76.1905 | 183.94
0.70975 | 1098808.4299 | 80.7963 | 195.0596
0.69568 | 1064177.2867 | 78.1509 | 188.6729
0.69665 | 1060034.4191 | 81.0486 | 195.6687
0.69373 | 1052791.5205 | 91.4552 | 220.7924
0.67958 | 1015171.495 | 76.3184 | 184.2489
0.68529 | 1013238.8583 | 66.8242 | 161.3278
0.6839 | 1025948.9611 | 61.8919 | 149.4202
0.70019 | 1043418.3638 | 74.4595 | 179.761
0.71279 | 1100993.3173 | 80.6111 | 194.6125
0.68369 | 1029066.8474 | 75.2078 | 181.5676
0.69322 | 1037536.3733 | 69.4981 | 167.7832
0.71032 | 1071383.7494 | 78.7164 | 190.0381
0.6827 | 1040823.7341 | 72.6316 | 175.3481
0.69317 | 1041611.6984 | 67.9145 | 163.9601
0.68977 | 1039090.4126 | 81.8291 | 197.5528
0.71369 | 1103574.3501 | 77.2622 | 186.5274
0.69174 | 1043702.3265 | 68.191 | 164.6277
0.6628 | 972800.503 | 68.1523 | 164.5342
0.66699 | 976500.4414 | 73.1525 | 176.6058
0.71351 | 1092362.867 | 72.9364 | 176.0841
0.67563 | 998226.1351 | 62.9559 | 151.989
0.70897 | 1092351.0795 | 76.9853 | 185.859
0.6935 | 1055318.0256 | 77.7439 | 187.6903
0.70182 | 1061694.4355 | 39.4723 | 95.2945
0.66921 | 981861.6816 | 71.6241 | 172.9158
0.7002 | 1071186.369 | 76.4398 | 184.542
0.72762 | 1144592.3043 | 73.0441 | 176.344
0.66288 | 981679.3932 | 68.7662 | 166.0163
0.70652 | 1075495.7085 | 70.6845 | 170.6475
0.68783 | 1039010.8209 | 71.9646 | 173.7379
0.66332 | 989159.4775 | 79.4293 | 191.7593
0.69426 | 1046527.6777 | 71.8138 | 173.3739
0.69892 | 1069726.5677 | 93.6666 | 226.1311
0.67464 | 1003394.2625 | 69.0619 | 166.7302
0.68386 | 1045768.435 | 78.1261 | 188.6131
0.68601 | 1011714.3501 | 71.9944 | 173.81
0.68966 | 1042044.9394 | 78.8543 | 190.371
0.68997 | 1053048.0362 | 84.4833 | 203.9607
0.69715 | 1049401.5257 | 73.4831 | 177.4039
0.69782 | 1067237.8706 | 74.6604 | 180.2461
0.71035 | 1087518.8337 | 86.3006 | 208.3482
0.68177 | 1007747.9939 | 83.3231 | 201.1597
0.70926 | 1072802.8722 | 95.3308 | 230.149
0.70386 | 1070770.7168 | 64.8225 | 156.4953
0.69566 | 1046315.4573 | 50.666 | 122.3186
0.69234 | 1056319.6788 | 72.6861 | 175.4798
0.67918 | 1034101.7299 | 79.4607 | 191.835
0.70049 | 1062977.2661 | 62.6071 | 151.147
0.68903 | 1033648.9468 | 69.6578 | 168.1687
0.69552 | 1036386.3744 | 74.4544 | 179.7488
0.69247 | 1055647.4516 | 66.9579 | 161.6506
0.65976 | 979281.0038 | 77.4626 | 187.0112
0.68665 | 1037157.2954 | 68.2665 | 164.81
0.68669 | 1027437.387 | 81.2741 | 196.213
0.70257 | 1075089.6467 | 71.0552 | 171.5424
0.68893 | 1046338.4418 | 75.8879 | 183.2096
0.72432 | 1126566.7821 | 75.9273 | 183.3048
0.67149 | 1000283.0679 | 67.3259 | 162.5391
0.70737 | 1071859.4655 | 72.1888 | 174.2792
0.68819 | 1043951.3235 | 82.4003 | 198.9319
0.68225 | 1022712.7559 | 73.1783 | 176.6681
0.68012 | 1018750.4246 | 59.7032 | 144.1364
0.67883 | 1021913.7777 | 76.9098 | 185.6766
0.6985 | 1053726.1953 | 66.259 | 159.9634
0.68105 | 1013870.9595 | 72.6735 | 175.4494
0.71333 | 1114596.6467 | 74.4401 | 179.7143
0.68562 | 1032009.7201 | 73.4805 | 177.3976
0.68998 | 1048967.7976 | 58.3914 | 140.9693
0.70789 | 1086219.7506 | 69.3497 | 167.4251
0.68428 | 1031892.7052 | 60.1061 | 145.109
0.71231 | 1103460.3738 | 73.8018 | 178.1734
0.69746 | 1053401.8054 | 75.4656 | 182.1902
0.70548 | 1102432.4589 | 79.8364 | 192.742
0.69085 | 1045453.0409 | 60.317 | 145.6182
0.70899 | 1073763.4834 | 64.5336 | 155.7979
0.68219 | 1008877.9412 | 75.5547 | 182.4053
0.66688 | 1000105.6653 | 73.2105 | 176.7458
0.70198 | 1066614.4258 | 80.9326 | 195.3886
0.73136 | 1128400.9474 | 67.7662 | 163.6021
0.67751 | 1021588.5755 | 64.2504 | 155.1141
0.70255 | 1066399.3484 | 70.4823 | 170.1593
0.71547 | 1094119.0323 | 76.4101 | 184.4702
0.67059 | 995624.5618 | 72.4791 | 174.9799
0.71364 | 1106605.6568 | 63.9759 | 154.4516
0.65846 | 959666.3443 | 58.0909 | 140.2438
0.69747 | 1065423.371 | 78.0423 | 188.4107
0.67358 | 1007733.0354 | 67.6167 | 163.2411
0.69538 | 1050403.5443 | 77.1732 | 186.3126
0.71784 | 1099129.1922 | 69.2022 | 167.0688
0.68622 | 1026443.6596 | 69.4565 | 167.6827
0.68698 | 1031143.13 | 65.9263 | 159.1601
0.68883 | 1038221.6568 | 89.3999 | 215.8305
0.67706 | 1016661.464 | 55.4101 | 133.7719
0.69056 | 1040411.9069 | 64.198 | 154.9876
0.69245 | 1039590.799 | 79.9674 | 193.0585
0.67628 | 1001633.2682 | 63.1099 | 152.3609
0.69149 | 1046390.5767 | 68.3921 | 165.113
0.68468 | 1030218.6407 | 68.0045 | 164.1773
0.68705 | 1049339.6595 | 78.0295 | 188.3799
0.6894 | 1037062.9801 | 71.6658 | 173.0165
0.66719 | 977441.1757 | 79.5785 | 192.1194
0.67054 | 974621.0979 | 75.8549 | 183.1299
0.70012 | 1050181.2245 | 73.2265 | 176.7844
0.70773 | 1090304.3061 | 69.6264 | 168.093
0.68747 | 1028856.6642 | 79.7155 | 192.4501
0.66423 | 987452.4861 | 80.7524 | 194.9535
0.69556 | 1037534.951 | 79.4524 | 191.815
0.67781 | 1006129.5035 | 59.1649 | 142.8366
0.68346 | 1032170.8353 | 77.1563 | 186.2717
0.69177 | 1057770.4622 | 78.1198 | 188.5979
0.69355 | 1055084.7405 | 85.7559 | 207.0332
0.67676 | 1016510.2432 | 67.8845 | 163.8876
0.66804 | 979089.7279 | 78.066 | 188.4679
0.70667 | 1079880.8088 | 78.0705 | 188.4788
0.64497 | 942221.7947 | 79.355 | 191.5798
0.69846 | 1058855.6797 | 77.4509 | 186.9829
0.68919 | 1041556.1388 | 82.0918 | 198.1872
0.67896 | 1019841.8951 | 69.7471 | 168.3843
0.66777 | 1004204.5801 | 64.0782 | 154.6984
0.69221 | 1044467.6643 | 70.0635 | 169.1483
0.69564 | 1052441.0917 | 80.5853 | 194.5501
0.68285 | 1029102.9592 | 64.2175 | 155.0347
0.66797 | 979955.0397 | 73.3033 | 176.9699
0.6871 | 1025747.061 | 89.9654 | 217.1957
0.67672 | 1032997.0004 | 84.4972 | 203.9944
0.6848 | 1051420.7005 | 78.4411 | 189.3737
0.70285 | 1068825.1403 | 80.9997 | 195.5506
0.69405 | 1052838.5858 | 64.6375 | 156.0488
0.68976 | 1047694.8324 | 74.1502 | 179.0144
0.73212 | 1138313.9834 | 75.9016 | 183.2427
0.66759 | 972759.7381 | 77.6127 | 187.3737
0.69302 | 1054463.3614 | 68.7674 | 166.0193
0.69277 | 1039758.9031 | 61.9219 | 149.4927
0.68124 | 1026973.0875 | 77.0177 | 185.9372
0.69698 | 1062487.4349 | 74.4178 | 179.6605
0.67486 | 1013322.8913 | 81.5632 | 196.9109
0.67986 | 1031771.8569 | 74.0867 | 178.8611
0.69184 | 1036180.9826 | 81.3551 | 196.4087
0.6894 | 1051117.1352 | 74.4309 | 179.6922
0.70877 | 1066198.1176 | 80.1259 | 193.4411
0.68459 | 1023257.8539 | 78.0063 | 188.3238
0.65833 | 965205.5661 | 66.2627 | 159.9722
0.67402 | 996737.1822 | 68.6664 | 165.7754
0.68047 | 1020685.951 | 83.0303 | 200.4528
0.67022 | 993636.2813 | 67.0045 | 161.7632
0.69416 | 1061509.3274 | 78.1256 | 188.6118
0.70212 | 1063492.3143 | 64.5413 | 155.8165
0.69128 | 1041042.5583 | 78.5941 | 189.7429
0.69139 | 1038048.1904 | 65.9446 | 159.2043
0.67125 | 1004589.7235 | 74.7787 | 180.5317
0.67623 | 1013485.5753 | 74.8863 | 180.7915
0.71571 | 1107355.8121 | 67.9567 | 164.0621
0.67418 | 1014070.0637 | 73.4283 | 177.2715
0.68894 | 1050836.593 | 82.8536 | 200.0263
0.6927 | 1023144.5703 | 75.4609 | 182.1788
0.6935 | 1041664.3375 | 91.9223 | 221.92
0.69537 | 1064707.8068 | 81.5557 | 196.8928
0.68705 | 1028669.8029 | 76.4924 | 184.6691
0.69802 | 1065417.5337 | 76.2658 | 184.1219
0.69037 | 1048587.8222 | 85.6203 | 206.7056
0.70247 | 1068651.6869 | 74.2645 | 179.2904
0.70715 | 1078785.9139 | 72.6775 | 175.459
0.66197 | 965119.5655 | 77.6578 | 187.4824
0.69873 | 1054898.3085 | 85.1051 | 205.4619
0.69251 | 1046577.3382 | 73.7762 | 178.1115
0.67405 | 997576.6289 | 63.0627 | 152.2469
0.69341 | 1031600.8208 | 74.3354 | 179.4614
0.69884 | 1073881.1836 | 80.1958 | 193.6098
0.69962 | 1072722.1975 | 71.1974 | 171.8857
0.70329 | 1072771.891 | 79.606 | 192.1859
0.66419 | 983822.1156 | 72.31 | 174.5718
0.70823 | 1119717.1688 | 71.4549 | 172.5075
0.69908 | 1082243.5051 | 77.3521 | 186.7446
0.69949 | 1052890.2248 | 71.7097 | 173.1224
0.69793 | 1054547.6516 | 80.4814 | 194.2994
0.69728 | 1059842.9108 | 51.2018 | 123.612
0.67917 | 1027304.9911 | 80.5678 | 194.5078
0.68654 | 1033052.8113 | 69.8705 | 168.6822
0.70071 | 1065192.002 | 87.2093 | 210.5419
0.70072 | 1054569.205 | 78.542 | 189.6172
0.65437 | 955370.6066 | 65.9844 | 159.3005
0.68281 | 1034525.5097 | 89.8676 | 216.9596
0.72018 | 1140923.3267 | 77.3371 | 186.7083
0.70437 | 1068507.3614 | 74.054 | 178.7822
0.69459 | 1062054.8324 | 84.6399 | 204.3387
0.65605 | 971839.9311 | 83.5049 | 201.5986
0.70162 | 1052686.6111 | 76.6527 | 185.0561
0.69737 | 1053279.9826 | 89.586 | 216.2798
0.71142 | 1101070.3439 | 71.8665 | 173.501
0.68828 | 1041451.83 | 68.3862 | 165.0989
0.67771 | 1016746.8029 | 57.6224 | 139.1129
0.68606 | 1046779.7196 | 70.8488 | 171.0441
0.69843 | 1055965.6085 | 75.5441 | 182.3796
0.69148 | 1049853.373 | 70.4999 | 170.2019
0.68122 | 1015885.965 | 74.513 | 179.8904
0.69055 | 1054932.1574 | 60.8901 | 147.0018
0.70039 | 1094437.0465 | 72.5584 | 175.1715
0.69454 | 1050378.2027 | 78.9269 | 190.5464
0.68312 | 1028812.681 | 62.1354 | 150.0082
0.69799 | 1064380.1215 | 77.9068 | 188.0838
0.70462 | 1071943.4713 | 79.394 | 191.6741
0.71212 | 1098399.0808 | 77.0694 | 186.062
0.6979 | 1041373.0006 | 82.5973 | 199.4076
0.69941 | 1064357.0084 | 87.3036 | 210.7696
0.68139 | 1007530.3655 | 56.7921 | 137.1084
0.69707 | 1048089.3627 | 76.8158 | 185.4498
0.68154 | 1013607.6201 | 67.2163 | 162.2745
0.68353 | 1021441.1753 | 71.1921 | 171.873
0.71416 | 1101498.0254 | 77.0659 | 186.0535
0.70328 | 1085577.4388 | 78.8608 | 190.3868
0.66309 | 980152.02 | 80.2932 | 193.8448
0.69869 | 1063045.7456 | 77.7318 | 187.6611
0.68269 | 1032968.9895 | 78.9656 | 190.6398
0.69123 | 1047988.3393 | 75.7544 | 182.8873
0.7129 | 1101422.4948 | 65.8325 | 158.9337
0.70175 | 1070984.8049 | 56.3177 | 135.9629
0.70322 | 1078380.0981 | 74.8604 | 180.729
0.68154 | 1017666.2649 | 76.2864 | 184.1715
0.70031 | 1072582.0659 | 63.1886 | 152.5508
0.66313 | 991903.5292 | 71.26 | 172.0368
0.68346 | 1027271.5462 | 67.9305 | 163.9988
0.70459 | 1090882.231 | 77.3601 | 186.7639
0.71947 | 1128144.4635 | 74.7012 | 180.3445
0.69622 | 1033231.766 | 65.1827 | 157.3651
0.6855 | 1028226.7571 | 50.528 | 121.9855
0.69681 | 1055452.8424 | 70.3339 | 169.801
0.69007 | 1040939.7685 | 73.6384 | 177.7788
0.70951 | 1094341.6738 | 86.9459 | 209.9061
0.6906 | 1057799.1244 | 69.2265 | 167.1275
0.67516 | 1013079.9828 | 73.2502 | 176.8417
0.66224 | 978958.0428 | 73.8624 | 178.3196
0.69902 | 1038935.7604 | 75.9324 | 183.3171
0.68158 | 1036317.6388 | 70.6017 | 170.4477
0.67876 | 1008243.6835 | 61.2844 | 147.9536
0.69047 | 1036827.5488 | 84.6478 | 204.3577
0.68691 | 1028403.9773 | 85.8888 | 207.354
0.6696 | 991463.5098 | 70.5402 | 170.2991
0.68847 | 1029269.005 | 74.6999 | 180.3415
0.69888 | 1052069.715 | 57.8184 | 139.586
0.67163 | 990667.4672 | 58.7915 | 141.9352
0.69186 | 1030843.3234 | 61.4717 | 148.4057
0.70789 | 1083673.6559 | 79.8926 | 192.8778
0.69342 | 1051911.1343 | 68.5614 | 165.5218
0.66767 | 967506.5842 | 63.6464 | 153.6561
0.67485 | 1007883.4124 | 52.7658 | 127.3878
0.69967 | 1061057.0617 | 55.9615 | 135.1031
0.67445 | 1001719.5698 | 70.758 | 170.8249
0.69722 | 1048961.2268 | 73.4448 | 177.3115
0.705 | 1066817.9024 | 75.6097 | 182.5379
0.67966 | 1007904.7204 | 70.6484 | 170.5604
0.70613 | 1083420.7912 | 73.0958 | 176.4688
0.69685 | 1073339.9346 | 80.3847 | 194.0659
0.6886 | 1039415.9413 | 68.2212 | 164.7006
0.68526 | 1034530.5681 | 68.2316 | 164.7256
0.67981 | 1015843.7961 | 81.1699 | 195.9615
0.70452 | 1076085.359 | 78.0336 | 188.3899
0.6935 | 1034118.4379 | 72.794 | 175.7402
0.72334 | 1113330.1347 | 69.5631 | 167.9402
0.71896 | 1109556.5692 | 65.7597 | 158.758
0.70037 | 1067958.5585 | 58.5664 | 141.3918
0.67911 | 1016530.6806 | 81.3087 | 196.2967
0.68108 | 1027618.2304 | 75.3893 | 182.006
0.70171 | 1045114.8002 | 59.364 | 143.3174
0.7131 | 1104535.7262 | 66.0002 | 159.3385
0.68507 | 1035250.0683 | 68.2974 | 164.8846
0.65058 | 945994.3941 | 65.7177 | 158.6565
0.69011 | 1044562.6785 | 73.8907 | 178.388
0.70441 | 1084820.4591 | 73.3687 | 177.1277
0.68743 | 1050661.815 | 83.0223 | 200.4336
0.71246 | 1097875.8733 | 80.421 | 194.1534
0.7016 | 1059752.1665 | 83.6117 | 201.8566
0.68198 | 1017981.4205 | 81.0819 | 195.749
0.67286 | 993514.2611 | 83.2343 | 200.9455
0.68958 | 1059368.7666 | 91.1214 | 219.9864
0.71114 | 1086598.7938 | 70.9502 | 171.289
0.73562 | 1152990.4457 | 68.8423 | 166.2001
0.67903 | 1017461.8392 | 65.8128 | 158.8861
0.69019 | 1033079.8022 | 74.7899 | 180.5588
0.69353 | 1045995.8716 | 77.054 | 186.0248
0.68933 | 1040966.7093 | 62.3029 | 150.4124
0.6946 | 1052738.6158 | 80.7287 | 194.8962
0.69301 | 1042036.0057 | 80.0906 | 193.3558
0.69222 | 1034315.1312 | 74.6595 | 180.2441
0.69033 | 1031197.8886 | 71.1647 | 171.8068
0.6888 | 1042411.5477 | 64.4858 | 155.6824
0.69986 | 1075952.0446 | 75.0511 | 181.1894
0.66697 | 978052.392 | 62.4842 | 150.8501
0.68837 | 1048513.4292 | 76.1075 | 183.7399
0.71523 | 1105273.3739 | 71.7568 | 173.2362
0.69144 | 1053271.9355 | 68.8462 | 166.2095
0.69191 | 1050104.2467 | 72.304 | 174.5574
0.69863 | 1072027.6947 | 76.6714 | 185.1012
0.7031 | 1082627.3132 | 71.4991 | 172.614
0.6953 | 1031706.6969 | 86.5801 | 209.023
0.68717 | 1039671.4832 | 50.9661 | 123.043
0.70957 | 1081999.2065 | 84.4082 | 203.7795
0.68994 | 1035930.4015 | 45.2387 | 109.2159
0.68574 | 1028337.1997 | 65.8656 | 159.0135
0.70176 | 1093758.9583 | 96.8183 | 233.74
0.69416 | 1052609.7989 | 79.1375 | 191.0549
0.6631 | 982204.2703 | 75.6754 | 182.6965
0.67921 | 1023998.609 | 78.6804 | 189.9512
0.68056 | 1022594.7393 | 71.5839 | 172.8187
0.68234 | 1020255.563 | 79.6433 | 192.276
0.67305 | 1005056.64 | 63.6299 | 153.6162
0.72013 | 1126324.9717 | 83.3088 | 201.1252
0.70023 | 1069100.6614 | 86.0662 | 207.7822
0.7152 | 1098945.5673 | 83.4656 | 201.5038
0.69334 | 1054541.459 | 79.3763 | 191.6314
0.70418 | 1055487.9581 | 76.5778 | 184.8751
0.69771 | 1056643.7595 | 73.9314 | 178.4862
0.68055 | 1018785.3416 | 71.3182 | 172.1773
0.69975 | 1068227.3918 | 82.9237 | 200.1954
0.66145 | 980920.5501 | 75.7649 | 182.9127
0.70976 | 1116070.9477 | 74.9127 | 180.8552
0.66908 | 1003613.8579 | 79.8964 | 192.8869
0.7125 | 1093780.1334 | 61.1872 | 147.7189
0.68469 | 1017147.7465 | 79.128 | 191.032
0.6779 | 1026120.5027 | 80.6258 | 194.6479
0.68962 | 1056449.4049 | 69.236 | 167.1504
0.68516 | 1011429.6687 | 56.6981 | 136.8812
0.69884 | 1074089.8115 | 78.5229 | 189.5711
0.70123 | 1079928.2204 | 67.6418 | 163.3018
0.68039 | 1019932.1509 | 93.1181 | 224.807
0.697 | 1061507.9924 | 53.8065 | 129.9004
0.6677 | 982163.1673 | 66.4812 | 160.4999
0.67574 | 1024219.7827 | 58.2151 | 140.5437
0.68746 | 1029956.8893 | 79.8014 | 192.6575
0.70271 | 1052644.8671 | 74.848 | 180.6991
0.70591 | 1084488.8108 | 83.4519 | 201.4706
0.6775 | 1005897.6654 | 86.6322 | 209.1486
0.67405 | 1009589.7808 | 74.8388 | 180.677
0.7023 | 1066095.872 | 68.6076 | 165.6334
0.70667 | 1090494.9171 | 63.6663 | 153.704
0.68721 | 1042462.7971 | 82.4003 | 198.9319
0.69323 | 1040025.415 | 67.8056 | 163.6971
0.68443 | 1008502.0153 | 81.5646 | 196.9144
0.69666 | 1050201.1836 | 81.8859 | 197.69
0.66402 | 992544.6686 | 74.6867 | 180.3096
0.68547 | 1024179.4608 | 74.6558 | 180.235
0.66801 | 980145.6665 | 84.1745 | 203.2151
0.7151 | 1084612.7569 | 79.1551 | 191.0973
0.69313 | 1032414.2281 | 76.0971 | 183.7147
0.67179 | 990446.9505 | 70.2383 | 169.5702
0.70475 | 1089296.3849 | 75.4547 | 182.1637
0.67305 | 1014254.9151 | 71.7918 | 173.3208
0.71946 | 1096387.0548 | 87.4239 | 211.0599
0.67595 | 980378.6331 | 72.4693 | 174.9563
0.70678 | 1083885.6889 | 80.0342 | 193.2196
0.70796 | 1088462.9038 | 59.0513 | 142.5624
0.70229 | 1084725.7085 | 83.9002 | 202.5531
0.67903 | 1016202.0625 | 71.2963 | 172.1245
0.69765 | 1050166.1879 | 78.559 | 189.6583
0.6951 | 1041909.041 | 57.5713 | 138.9893
0.66262 | 976141.2091 | 81.5456 | 196.8686
0.69755 | 1075336.8993 | 78.7585 | 190.1399
0.68648 | 1045520.9342 | 72.5578 | 175.17
0.6697 | 994599.1428 | 68.9609 | 166.4863
0.70938 | 1084164.2243 | 70.9667 | 171.3288
0.64283 | 920798.841 | 66.513 | 160.5765
0.69468 | 1048163.236 | 64.2925 | 155.2158
0.70099 | 1075582.2236 | 73.0989 | 176.4763
0.7111 | 1086035.8255 | 72.7596 | 175.6572
0.68615 | 1028956.1936 | 71.4427 | 172.4778
0.68756 | 1034562.4722 | 76.8737 | 185.5895
0.70212 | 1086048.9134 | 73.6908 | 177.9054
0.67659 | 998760.9145 | 60.9031 | 147.0331
0.68448 | 1022555.3134 | 81.1488 | 195.9106
0.67243 | 1005220.9174 | 71.2187 | 171.9372
0.67502 | 994991.194 | 75.307 | 181.8072
0.71431 | 1083886.7051 | 85.5365 | 206.5034
0.65097 | 952932.3739 | 76.3966 | 184.4378
0.6765 | 994178.4138 | 71.741 | 173.198
0.68571 | 1014051.4491 | 72.1929 | 174.289
0.68336 | 1024580.5614 | 81.4429 | 196.6206
0.69546 | 1049426.8088 | 80.3752 | 194.043
0.67495 | 998717.7749 | 64.9911 | 156.9025
0.67123 | 993927.3453 | 73.988 | 178.6228
0.68302 | 1028724.3482 | 81.8135 | 197.5153
0.68272 | 1024945.5524 | 79.9645 | 193.0513
0.6573 | 970356.7598 | 62.9048 | 151.8656
0.72701 | 1133162.7312 | 74.2553 | 179.2681
0.67151 | 987648.8586 | 71.7995 | 173.3393
0.68946 | 1040370.8515 | 68.1987 | 164.6461
0.67071 | 983571.3126 | 74.9207 | 180.8746
0.70104 | 1073746.7418 | 74.7712 | 180.5136
0.6835 | 1029655.6994 | 68.0851 | 164.3719
0.68559 | 1028818.8465 | 69.6807 | 168.2242
0.68139 | 1021424.8524 | 81.4264 | 196.5807
0.69341 | 1048920.3572 | 82.169 | 198.3736
0.69933 | 1057550.018 | 80.5481 | 194.4604
0.68563 | 1017744.7258 | 82.1677 | 198.3703
0.68092 | 1015850.2487 | 85.587 | 206.6252
0.69169 | 1043494.3608 | 73.6469 | 177.7992
0.67866 | 1026585.1213 | 62.4088 | 150.6681
0.66218 | 981619.7526 | 78.9596 | 190.6253
0.68642 | 1039619.5536 | 75.6769 | 182.7003
0.7168 | 1115192.1762 | 72.2418 | 174.4071
0.68863 | 1022912.1375 | 86.6736 | 209.2485
0.69712 | 1061583.4838 | 70.7102 | 170.7095
0.68178 | 1029471.8418 | 83.5845 | 201.7909
0.66722 | 983164.2547 | 73.6267 | 177.7505
0.70347 | 1083094.0377 | 59.1349 | 142.7642
0.66854 | 999330.3217 | 81.1973 | 196.0276
0.67767 | 1029015.0703 | 79.2459 | 191.3164
0.68909 | 1054077.0286 | 78.1498 | 188.6703
0.71268 | 1075744.0954 | 75.421 | 182.0823
0.70803 | 1080810.9911 | 67.0469 | 161.8655
0.69236 | 1037935.1181 | 65.8451 | 158.9642
0.68616 | 1037511.7766 | 58.0756 | 140.2069
0.68522 | 1027718.7836 | 63.9088 | 154.2894
0.67534 | 997139.2823 | 68.5596 | 165.5176
0.69467 | 1047393.7523 | 73.6041 | 177.6961
0.68187 | 1016682.2489 | 63.5335 | 153.3835
0.68191 | 1028477.4724 | 73.6274 | 177.7523
0.7014 | 1054085.7738 | 67.9209 | 163.9756
0.69514 | 1046202.0903 | 68.0201 | 164.215
0.70583 | 1056641.3079 | 85.8817 | 207.3368
0.71044 | 1100751.0929 | 94.8197 | 228.9149
0.69967 | 1076693.2804 | 81.9668 | 197.8855
0.69843 | 1066941.7811 | 60.8883 | 146.9975
0.69146 | 1050897.6666 | 79.3167 | 191.4875
0.69269 | 1049785.1384 | 66.5643 | 160.7004
0.67148 | 999298.7324 | 74.36 | 179.5209
0.70477 | 1059820.5988 | 83.2412 | 200.9619
0.69006 | 1045250.2679 | 70.3348 | 169.8033
0.69006 | 1032736.8728 | 84.3641 | 203.6731
0.68579 | 1036711.4257 | 77.2321 | 186.4548
0.67117 | 992151.3907 | 66.8101 | 161.2938
0.6524 | 961238.3501 | 60.7879 | 146.7549
0.68667 | 1035123.0794 | 80.2942 | 193.8473
0.67325 | 990516.5014 | 70.0915 | 169.2158
0.70472 | 1065420.7996 | 62.3739 | 150.584
0.69575 | 1047090.2756 | 62.0918 | 149.9028
0.69482 | 1052735.2658 | 88.4065 | 213.4321
0.70534 | 1093344.6261 | 70.0571 | 169.1328
0.67456 | 998330.9733 | 86.3561 | 208.4822
0.6683 | 990697.8142 | 75.7358 | 182.8424
0.67829 | 999254.5083 | 66.5817 | 160.7426
0.69269 | 1053785.2279 | 70.6401 | 170.5404
0.69253 | 1045632.3933 | 78.7973 | 190.2335
0.69024 | 1045165.549 | 72.1372 | 174.1546
0.68385 | 1040398.764 | 62.7686 | 151.5368
0.70344 | 1075779.6039 | 69.1186 | 166.8671
0.68343 | 1012904.7379 | 80.5068 | 194.3606
0.69407 | 1052874.075 | 77.5 | 187.1017
0.6687 | 989386.5707 | 72.2286 | 174.3752
0.67697 | 1016192.6027 | 67.8426 | 163.7866
0.7291 | 1141582.5848 | 76.4357 | 184.5322
0.69162 | 1059578.4018 | 81.6028 | 197.0065
0.67764 | 1018350.0823 | 62.0009 | 149.6833
0.68224 | 1022571.2351 | 93.1067 | 224.7795
0.67823 | 1011451.6176 | 81.0074 | 195.5691
0.67496 | 993751.4245 | 68.3651 | 165.048
0.6764 | 1003754.4714 | 80.609 | 194.6073
0.68468 | 1012924.7584 | 75.5823 | 182.4718
0.68641 | 1030556.9328 | 76.7041 | 185.1801
0.69536 | 1037421.3204 | 63.5717 | 153.4757
0.68892 | 1039511.4196 | 83.5242 | 201.6453
0.71844 | 1101463.4412 | 73.0395 | 176.3331
0.68641 | 1035859.988 | 80.3816 | 194.0583
0.69033 | 1042218.4002 | 68.2338 | 164.731
0.6909 | 1038005.5678 | 69.7627 | 168.422
0.69619 | 1043002.1157 | 87.4161 | 211.0412
0.69808 | 1060879.1775 | 74.1879 | 179.1055
0.67303 | 997678.8393 | 59.0847 | 142.643
0.65638 | 975246.6755 | 79.141 | 191.0634
0.68993 | 1036845.0766 | 67.9209 | 163.9755
0.65681 | 965757.101 | 79.8698 | 192.8228
0.68565 | 1023289.32 | 82.6128 | 199.445
0.69858 | 1052815.4992 | 62.1782 | 150.1115
0.66791 | 990767.424 | 74.1193 | 178.9398
0.69543 | 1050030.1402 | 89.5589 | 216.2143
0.70393 | 1075414.2279 | 86.1188 | 207.9091
0.69375 | 1028631.3489 | 72.3817 | 174.7449
0.70528 | 1072840.6032 | 77.7146 | 187.6195
0.67627 | 1001679.8036 | 74.0092 | 178.674
0.71848 | 1102070.2868 | 67.3718 | 162.6499
0.66708 | 980818.9687 | 75.1171 | 181.3486
0.70921 | 1072902.3582 | 70.246 | 169.5888
0.68904 | 1039807.9259 | 70.0722 | 169.1691
0.67608 | 1031875.7593 | 72.7124 | 175.5432
0.70829 | 1083516.4356 | 83.0959 | 200.6112
0.68689 | 1029047.452 | 67.7083 | 163.4622
0.71772 | 1116568.0923 | 70.3117 | 169.7474
0.70165 | 1059893.1453 | 53.3689 | 128.8439
0.69793 | 1076142.912 | 65.7019 | 158.6184
0.67482 | 993204.6543 | 87.2986 | 210.7574
0.71176 | 1105148.1238 | 73.9714 | 178.5827
0.65534 | 976314.6345 | 65.4339 | 157.9714
0.7059 | 1094701.9588 | 90.4873 | 218.4556
0.69525 | 1035616.011 | 57.2111 | 138.1198
0.7041 | 1068787.7596 | 74.1018 | 178.8975
0.68694 | 1020408.537 | 79.1213 | 191.0157
0.67538 | 1008995.3803 | 77.9678 | 188.231
0.67746 | 1003624.1075 | 65.0685 | 157.0892
0.68712 | 1023951.3537 | 76.1342 | 183.8041
0.68101 | 1016332.4514 | 74.5686 | 180.0245
0.69178 | 1045128.6085 | 77.1844 | 186.3395
0.71106 | 1106293.4139 | 70.3173 | 169.7609
0.68526 | 1040472.0089 | 80.8755 | 195.2507
0.67651 | 999877.6585 | 73.85 | 178.2896
0.69791 | 1045838.836 | 78.2057 | 188.8053
0.67195 | 999104.2857 | 83.8165 | 202.351
0.68771 | 1032619.1173 | 81.5625 | 196.9093
0.69977 | 1078199.3869 | 82.07 | 198.1346
0.65875 | 975629.9671 | 66.9365 | 161.5989
0.69099 | 1047203.559 | 60.4055 | 145.8318
0.6747 | 1003794.2942 | 74.3375 | 179.4667
0.65976 | 977454.3416 | 56.9034 | 137.3768
0.6514 | 947751.294 | 74.3322 | 179.4537
0.7241 | 1128543.0049 | 81.3376 | 196.3663
0.69458 | 1054144.7812 | 75.8322 | 183.0751
0.66675 | 978482.896 | 64.4573 | 155.6137
0.70331 | 1068144.6008 | 82.4979 | 199.1675
0.66154 | 966539.1181 | 73.0027 | 176.2441
0.681 | 1035964.706 | 74.2627 | 179.2859
0.6868 | 1011772.5594 | 73.6084 | 177.7063
0.69838 | 1055686.4158 | 62.3656 | 150.5638
0.71591 | 1101841.9163 | 80.3222 | 193.915
0.72787 | 1140975.0273 | 62.4795 | 150.8388
0.69163 | 1038269.9647 | 75.009 | 181.0877
0.71145 | 1109384.1846 | 68.4894 | 165.348
0.69462 | 1063260.1022 | 67.2314 | 162.3108
0.68742 | 1048479.8636 | 68.9345 | 166.4226
0.68132 | 1020361.4939 | 55.2642 | 133.4195
0.69788 | 1073536.4257 | 85.151 | 205.5727
0.69492 | 1045456.4267 | 80.1329 | 193.458
0.69208 | 1045729.9664 | 58.035 | 140.1089
0.68285 | 1005550.4861 | 58.6946 | 141.7014
0.65601 | 961296.6536 | 78.0522 | 188.4346
0.68872 | 1042019.3872 | 53.4536 | 129.0483
0.71111 | 1095649.3535 | 73.4023 | 177.2088
0.68154 | 1017091.133 | 61.5703 | 148.6439
0.68027 | 1024159.0378 | 75.2246 | 181.6082
0.69645 | 1069396.0609 | 98.1484 | 236.9512
0.68526 | 1047317.2654 | 67.1285 | 162.0626
0.67059 | 996503.328 | 85.2506 | 205.8131
0.7015 | 1068834.0434 | 82.7641 | 199.8102
0.67735 | 1004284.5079 | 89.3507 | 215.7117
0.6822 | 1026647.9112 | 78.5386 | 189.6089
0.69363 | 1042175.6056 | 81.615 | 197.0359
0.6665 | 987058.5239 | 73.9979 | 178.6468
0.69621 | 1048915.3493 | 81.8388 | 197.5764
0.6775 | 1009193.6147 | 63.472 | 153.2349
0.67877 | 1004419.5655 | 77.2269 | 186.4422
0.68687 | 1028589.1354 | 74.8245 | 180.6423
0.66485 | 982036.3401 | 58.8739 | 142.1341
0.70113 | 1078225.4734 | 69.5038 | 167.7971
0.67364 | 997194.1766 | 65.0186 | 156.9687
0.70746 | 1063304.5152 | 71.6037 | 172.8666
0.67936 | 998230.2387 | 64.5374 | 155.807
0.69528 | 1037633.7723 | 67.4497 | 162.8381
0.66144 | 982393.3946 | 67.1255 | 162.0553
0.68265 | 1032527.5576 | 75.575 | 182.4542
0.69392 | 1042859.9263 | 64.0376 | 154.6005
0.71127 | 1102806.0246 | 70.0744 | 169.1747
0.6968 | 1082928.9192 | 72.1657 | 174.2235
0.70256 | 1082665.257 | 73.4152 | 177.2399
0.67582 | 993369.4739 | 56.4551 | 136.2947
0.69671 | 1055589.0593 | 98.1821 | 237.0326
0.64342 | 924794.7028 | 64.9603 | 156.8281
0.70442 | 1082226.8164 | 69.4357 | 167.6326
0.70471 | 1077471.8055 | 74.6395 | 180.1957
0.69467 | 1057710.9363 | 82.7082 | 199.6753
0.69961 | 1058348.3043 | 72.0696 | 173.9915
0.67404 | 1003965.694 | 67.3172 | 162.5181
0.69477 | 1034458.6973 | 74.3962 | 179.6084
0.68242 | 1017628.2842 | 72.7967 | 175.7468
0.69103 | 1025994.3618 | 70.7337 | 170.7662
0.68535 | 1031164.6714 | 89.2435 | 215.4529
0.69546 | 1046122.3861 | 93.4398 | 225.5838
0.6757 | 998498.4306 | 73.8367 | 178.2576
0.67427 | 1011557.8802 | 59.7961 | 144.3607
0.69934 | 1071820.1709 | 76.5791 | 184.8784
0.69397 | 1057125.4896 | 87.3895 | 210.977
0.70327 | 1073294.0171 | 86.0296 | 207.6939
0.68875 | 1037564.2069 | 66.1659 | 159.7387
0.6867 | 1025048.2177 | 65.5636 | 158.2845
0.69548 | 1061690.6992 | 65.0471 | 157.0375
0.69694 | 1053575.5609 | 80.5215 | 194.396
0.70292 | 1069000.8347 | 74.2415 | 179.2348
0.70216 | 1078816.7808 | 71.3429 | 172.237
0.69912 | 1053364.5435 | 69.9624 | 168.9043
0.69653 | 1043890.0236 | 67.7344 | 163.5254
0.69795 | 1074262.7779 | 85.2001 | 205.6912
0.70552 | 1078728.3696 | 72.8412 | 175.8542
0.65894 | 966520.0774 | 66.2922 | 160.0436
0.70257 | 1053220.7151 | 76.5878 | 184.8994
0.68957 | 1032565.7395 | 74.8201 | 180.6316
0.68388 | 1019309.0247 | 76.1482 | 183.8379
0.68186 | 1034334.3656 | 80.7983 | 195.0643
0.68409 | 1013832.3703 | 61.1441 | 147.6149
0.68581 | 1039672.1316 | 48.4461 | 116.9592
0.70644 | 1099410.9852 | 62.0301 | 149.7538
0.68617 | 1036874.5206 | 77.2013 | 186.3804
0.69443 | 1052550.096 | 85.4406 | 206.2718
0.69795 | 1076168.0889 | 71.5967 | 172.8497
0.65937 | 984611.2407 | 77.9307 | 188.1413
0.7139 | 1104871.2002 | 80.7697 | 194.9953
0.6857 | 1023553.1584 | 88.5391 | 213.7522
0.68857 | 1023763.167 | 70.8322 | 171.0041
0.67368 | 986389.6742 | 64.31 | 155.258
0.68406 | 1033600.8365 | 78.8557 | 190.3745
0.69429 | 1065091.0914 | 81.7123 | 197.2709
0.7158 | 1106099.938 | 78.9462 | 190.5929
0.71596 | 1098965.6681 | 83.3784 | 201.2933
0.69825 | 1056317.5618 | 70.2458 | 169.5885
0.67543 | 993947.2037 | 75.1818 | 181.5049
0.71305 | 1095001.3913 | 69.405 | 167.5585
0.68676 | 1037650.8915 | 71.9821 | 173.7802
0.67973 | 995483.6472 | 66.1902 | 159.7972
0.66878 | 980930.3061 | 62.4989 | 150.8858
0.69358 | 1050179.6103 | 58.1555 | 140.3997
0.72201 | 1117922.7662 | 79.1046 | 190.9753
0.67429 | 1007751.3388 | 81.1823 | 195.9914
0.69617 | 1058540.7029 | 70.6043 | 170.4537
0.68405 | 1023659.8397 | 71.2929 | 172.1163
0.68513 | 1021239.9101 | 50.2118 | 121.2219
0.70393 | 1073157.2543 | 72.6295 | 175.3432
0.72227 | 1130240.2697 | 70.7242 | 170.7432
0.68921 | 1036862.4678 | 75.1705 | 181.4777
0.67764 | 1013792.4648 | 82.8139 | 199.9303
0.67072 | 1007658.7722 | 73.3403 | 177.0591
0.68715 | 1027019.3708 | 76.3584 | 184.3455
0.68168 | 1013497.5688 | 75.9582 | 183.3794
0.70131 | 1066581.0068 | 64.8039 | 156.4504
0.67171 | 1002429.6717 | 69.2967 | 167.297
0.67714 | 1010099.2119 | 63.5623 | 153.4529
0.68017 | 1022101.156 | 71.5897 | 172.8329
0.69776 | 1066227.955 | 84.5863 | 204.2094
0.69094 | 1039438.3206 | 73.4402 | 177.3002
0.69631 | 1051474.0114 | 71.9403 | 173.6792
0.6937 | 1050185.9239 | 66.4928 | 160.5278
0.67636 | 1021087.6915 | 70.1011 | 169.239
0.68153 | 1020922.9987 | 84.4429 | 203.8632
0.71123 | 1119324.426 | 73.0775 | 176.4246
0.70163 | 1063637.5594 | 79.1151 | 191.0007
0.67533 | 994345.0268 | 56.5518 | 136.5282
0.68586 | 1026241.402 | 63.7573 | 153.9237
0.69314 | 1049337.563 | 70.2958 | 169.709
0.70698 | 1082615.9943 | 84.7663 | 204.644
0.67577 | 1004864.9611 | 65.312 | 157.677
0.64233 | 925060.7117 | 75.1599 | 181.452
0.70816 | 1081668.6162 | 67.0878 | 161.9642
0.71417 | 1113196.3981 | 77.8839 | 188.0283
0.6654 | 987685.9728 | 74.1676 | 179.0564
0.70407 | 1067861.637 | 75.7833 | 182.9571
0.68097 | 1013582.1358 | 67.2705 | 162.4054
0.70465 | 1068337.4056 | 64.6026 | 155.9645
0.66538 | 975842.3741 | 81.1827 | 195.9924
0.67671 | 1012997.9388 | 74.7964 | 180.5745
0.70159 | 1070557.6287 | 67.021 | 161.8031
0.69974 | 1079774.6867 | 73.5144 | 177.4794
0.70154 | 1073649.9159 | 60.9995 | 147.2657
0.69931 | 1055374.6715 | 73.6736 | 177.8638
0.67629 | 1002035.0333 | 70.5872 | 170.4126
0.71455 | 1107504.1109 | 66.0436 | 159.4433
0.69925 | 1065792.3835 | 72.6134 | 175.3042
0.67691 | 1011687.9987 | 74.6781 | 180.2888
0.68103 | 1006491.9859 | 83.3104 | 201.1292
0.67402 | 998507.3886 | 75.6847 | 182.719
0.71708 | 1084925.0144 | 73.8591 | 178.3116
0.69629 | 1049156.3665 | 78.6366 | 189.8455
0.69964 | 1068522.6396 | 69.0244 | 166.6395
0.66858 | 979687.7873 | 74.863 | 180.7352
0.69005 | 1035676.2469 | 50.5581 | 122.058
0.70658 | 1072770.7591 | 62.6391 | 151.2241
0.66715 | 984462.5909 | 73.7316 | 178.0039
0.69434 | 1045248.7139 | 66.1198 | 159.6274
0.6906 | 1036362.0196 | 76.6687 | 185.0946
0.66263 | 975550.9573 | 73.5625 | 177.5956
0.68747 | 1035759.6109 | 60.051 | 144.976
0.67178 | 995512.9031 | 70.599 | 170.441
0.69598 | 1057552.6209 | 75.9247 | 183.2985
0.67552 | 1015947.0685 | 76.8352 | 185.4967
0.68611 | 1033415.9373 | 80.2272 | 193.6855
0.68357 | 999461.169 | 67.8921 | 163.9061
0.69015 | 1035767.2032 | 79.9386 | 192.9889
0.68794 | 1039206.9238 | 70.1438 | 169.342
0.67637 | 994296.7561 | 51.9372 | 125.3876
0.7024 | 1063485.5954 | 75.6525 | 182.6412
0.70217 | 1058226.827 | 68.1129 | 164.4392
0.65819 | 970238.6861 | 62.8119 | 151.6413
0.70062 | 1067857.2444 | 68.8855 | 166.3043
0.6954 | 1056057.1371 | 80.8891 | 195.2834
0.7049 | 1063569.9684 | 83.52 | 201.6351
0.69388 | 1064137.9551 | 87.9809 | 212.4046
0.70764 | 1096546.2629 | 71.9535 | 173.7111
0.68308 | 1019925.6692 | 71.1338 | 171.7322
0.67877 | 1014935.3123 | 79.6441 | 192.2779
0.72144 | 1116761.5068 | 71.8781 | 173.529
0.69094 | 1057275.7166 | 72.8999 | 175.996
0.69596 | 1053905.99 | 68.4402 | 165.2292
0.70634 | 1081595.7182 | 76.5117 | 184.7155
0.70344 | 1085651.7738 | 78.8484 | 190.3569
0.69902 | 1052077.3102 | 69.8868 | 168.7216
0.67642 | 995139.9348 | 54.2887 | 131.0646
0.69035 | 1027688.0218 | 77.6902 | 187.5607
0.6827 | 1032824.9511 | 75.592 | 182.4952
0.70222 | 1060193.6471 | 66.3125 | 160.0925
0.70614 | 1081159.9398 | 71.0949 | 171.6383
0.66046 | 977568.5835 | 84.913 | 204.998
0.67652 | 997559.6972 | 73.677 | 177.872
0.67189 | 1005077.9147 | 74.7441 | 180.4482
0.68726 | 1028124.3791 | 74.7482 | 180.4581
0.67816 | 1005198.2828 | 85.0336 | 205.2892
0.68168 | 1016641.7773 | 69.2508 | 167.1862
0.70481 | 1074323.4291 | 68.8904 | 166.3162
0.66595 | 970934.5917 | 67.6999 | 163.4421
0.67373 | 1008284.4654 | 79.7144 | 192.4477
0.69388 | 1045388.9448 | 64.1962 | 154.9833
0.70371 | 1088364.1211 | 67.9303 | 163.9982
0.71462 | 1075433.6664 | 68.0151 | 164.203
0.68923 | 1026214.6563 | 74.396 | 179.6079
0.69507 | 1040749.0301 | 76.498 | 184.6826
0.69772 | 1061050.3813 | 70.6146 | 170.4787
0.68224 | 1034257.5669 | 68.5152 | 165.4103
0.68943 | 1040555.4443 | 78.7683 | 190.1634
0.69051 | 1051130.0672 | 58.4673 | 141.1527
0.70698 | 1079508.3072 | 79.6876 | 192.3829
0.69231 | 1038071.3761 | 65.8582 | 158.9958
0.66965 | 993399.8599 | 82.1159 | 198.2452
0.70591 | 1085193.8168 | 73.2381 | 176.8123
0.67365 | 993738.217 | 75.3642 | 181.9453
0.67029 | 994391.4813 | 63.2141 | 152.6123
0.6836 | 1036334.1251 | 68.5166 | 165.4137
0.68582 | 1034215.886 | 77.043 | 185.9983
0.69338 | 1032665.1999 | 72.7043 | 175.5236
0.68332 | 1026666.0797 | 78.4112 | 189.3014
0.70231 | 1047158.3204 | 74.5001 | 179.8592
0.6835 | 1016649.8758 | 74.321 | 179.4267
0.70918 | 1077283.2844 | 82.5148 | 199.2083
0.71631 | 1112436.5648 | 73.9859 | 178.6178
0.65916 | 959719.1391 | 80.5947 | 194.5729
0.67863 | 1011767.8135 | 69.2285 | 167.1324
0.68609 | 1030560.5722 | 69.7477 | 168.3858
0.68225 | 1010325.6026 | 77.2847 | 186.5818
0.68519 | 1039737.7711 | 53.6117 | 129.4301
0.67912 | 1016898.6055 | 65.7679 | 158.7777
0.6894 | 1036009.7818 | 62.0007 | 149.6829
0.69909 | 1048751.4403 | 73.234 | 176.8025
0.67898 | 1026297.2308 | 75.5258 | 182.3354
0.6856 | 1030136.4098 | 53.0844 | 128.1571
0.68732 | 1014308.4324 | 62.8821 | 151.8108
0.69708 | 1058037.1244 | 70.4788 | 170.1509
0.71354 | 1102226.1508 | 69.8152 | 168.5487
0.71408 | 1109015.7445 | 64.0755 | 154.6921
0.68713 | 1031977.1678 | 65.9649 | 159.2534
0.67879 | 1014394.1738 | 81.1767 | 195.9779
0.68235 | 1024541.9392 | 76.7101 | 185.1946
0.6974 | 1074977.9871 | 69.6644 | 168.1847
0.67139 | 998694.3388 | 67.2283 | 162.3035
0.68555 | 1036986.3752 | 62.4068 | 150.6633
0.66718 | 1001821.2685 | 69.7241 | 168.3289
0.69439 | 1062166.0047 | 72.8726 | 175.93
0.6955 | 1039285.6376 | 82.7166 | 199.6955
0.67411 | 1015058.1362 | 78.5844 | 189.7195
0.67675 | 1006422.4268 | 72.5979 | 175.2668
0.68241 | 1028618.7624 | 64.2279 | 155.06
0.67839 | 1023554.0134 | 77.0837 | 186.0966
0.67829 | 1006430.0338 | 83.6409 | 201.9271
0.69079 | 1038122.091 | 62.0015 | 149.6849
0.69274 | 1029983.8434 | 80.764 | 194.9816
0.66544 | 973149.6333 | 64.2326 | 155.0713
0.69672 | 1058655.4191 | 77.6299 | 187.4151
0.71023 | 1088772.8739 | 67.3321 | 162.554
0.6828 | 1038631.6136 | 76.005 | 183.4922
0.6889 | 1034344.5991 | 67.8587 | 163.8254
0.69954 | 1086138.3867 | 75.7956 | 182.9868
0.68844 | 1037860.942 | 75.3263 | 181.8538
0.70766 | 1090839.1919 | 77.0216 | 185.9465
0.69385 | 1064940.0025 | 78.4138 | 189.3076
0.68294 | 1024523.4536 | 59.7937 | 144.3548
0.70876 | 1087395.9265 | 78.383 | 189.2333
0.69239 | 1048154.5509 | 64.5044 | 155.7275
0.72048 | 1109120.0338 | 78.6481 | 189.8732
0.69529 | 1059122.0361 | 64.9076 | 156.7008
0.71323 | 1113871.5183 | 76.2396 | 184.0587
0.67534 | 995348.1693 | 70.2152 | 169.5144
0.65547 | 957995.1583 | 85.28 | 205.884
0.6896 | 1045669.7051 | 75.7657 | 182.9145
0.6836 | 1026100.1253 | 61.9278 | 149.5069
0.70632 | 1077822.3156 | 77.9612 | 188.215
0.68246 | 1030006.0666 | 66.29 | 160.0382
0.67117 | 990056.5047 | 66.6685 | 160.9519
0.69049 | 1035887.3127 | 69.6614 | 168.1776
0.67549 | 1001992.57 | 77.414 | 186.894
0.69783 | 1061266.7292 | 87.2465 | 210.6317
0.71195 | 1102002.1936 | 87.2543 | 210.6505
0.68632 | 1058372.022 | 67.7318 | 163.519
0.69919 | 1061492.0182 | 60.5113 | 146.0873
0.67354 | 1007269.0934 | 79.6568 | 192.3085
0.6762 | 1004024.1982 | 67.107 | 162.0107
0.69834 | 1088328.6848 | 77.8699 | 187.9946
0.68195 | 1022987.5392 | 88.9773 | 214.8103
0.68219 | 1009292.1789 | 64.9826 | 156.8819
0.69911 | 1053730.6537 | 69.9879 | 168.9657
0.69128 | 1031024.547 | 81.0164 | 195.5908
0.67713 | 1019973.5598 | 76.9565 | 185.7894
0.67271 | 1002053.7648 | 83.9687 | 202.7184
0.68823 | 1032570.7977 | 74.9653 | 180.9822
0.68381 | 1024480.5015 | 69.567 | 167.9497
0.71816 | 1106173.9657 | 82.043 | 198.0694
0.68188 | 1010189.9115 | 77.3159 | 186.657
0.7065 | 1080658.6972 | 86.2236 | 208.1621
0.71858 | 1121051.2076 | 76.1592 | 183.8645
0.67823 | 1020501.4364 | 73.4529 | 177.331
0.68905 | 1030466.3934 | 69.4421 | 167.648
0.70556 | 1085733.1685 | 69.5479 | 167.9035
0.68292 | 1012826.9616 | 92.1274 | 222.4151
0.70749 | 1084971.5843 | 90.9609 | 219.5991
0.68523 | 1037935.2671 | 73.9446 | 178.518
0.68078 | 1017213.521 | 85.7198 | 206.946
0.67639 | 1003866.4041 | 68.9152 | 166.3759
0.69168 | 1053367.6214 | 82.7241 | 199.7136
0.6738 | 1009791.1282 | 74.8819 | 180.781
0.69206 | 1046813.7535 | 78.9408 | 190.58
0.69939 | 1073450.2283 | 80.1006 | 193.3798
0.70374 | 1056053.4785 | 74.9805 | 181.019
0.69836 | 1053668.6424 | 81.9564 | 197.8602
0.73031 | 1137887.3387 | 74.228 | 179.2022
0.67776 | 1014339.0712 | 81.2949 | 196.2633
0.68007 | 1029608.9795 | 66.4973 | 160.5388
0.70114 | 1057812.6012 | 66.2658 | 159.9798
0.68189 | 1024086.5564 | 78.3783 | 189.2219
0.70817 | 1085122.1408 | 79.9679 | 193.0597
0.6815 | 1022347.8687 | 83.0395 | 200.4751
0.68146 | 1013369.2184 | 67.2086 | 162.2559
0.71249 | 1088724.1772 | 60.9059 | 147.0399
0.67493 | 1013090.1553 | 73.2761 | 176.9042
0.68785 | 1050089.7253 | 82.9261 | 200.2014
0.69346 | 1044707.2679 | 87.9047 | 212.2208
0.70996 | 1090683.2822 | 83.4496 | 201.4652
0.67162 | 1007002.9736 | 73.195 | 176.7083
0.6736 | 1004812.9104 | 81.0288 | 195.6209
0.67043 | 983605.1232 | 64.4388 | 155.5691
0.67262 | 1000374.7919 | 83.8388 | 202.4047
0.69079 | 1029398.6847 | 74.4579 | 179.7574
0.68179 | 1031365.9701 | 72.0676 | 173.9867
0.71781 | 1119717.7646 | 77.9432 | 188.1716
0.72659 | 1119541.2358 | 82.0713 | 198.1377
0.68663 | 1047686.5119 | 68.9608 | 166.4862
0.69022 | 1046025.0687 | 62.5255 | 150.9499
0.70313 | 1072146.5601 | 79.5781 | 192.1186
0.68596 | 1046417.4894 | 87.5151 | 211.2802
0.70214 | 1065088.5489 | 75.7944 | 182.9839
0.67554 | 1002724.5437 | 63.2476 | 152.6933
0.68695 | 1037483.8606 | 77.1589 | 186.2782
0.70455 | 1057515.6703 | 63.9578 | 154.4078
0.67457 | 1003672.0246 | 69.0878 | 166.7927
0.70558 | 1088052.3976 | 59.2281 | 142.9892
0.65904 | 975198.9587 | 72.236 | 174.3932
0.66459 | 967104.1152 | 67.5704 | 163.1293
0.68528 | 1032533.581 | 62.1637 | 150.0765
0.6923 | 1043587.1553 | 69.9958 | 168.9848
0.72733 | 1153975.3397 | 81.0581 | 195.6917
0.71009 | 1100896.1503 | 70.0695 | 169.1627
0.68968 | 1039870.176 | 64.2465 | 155.1047
0.69206 | 1048055.8385 | 73.9287 | 178.4796
0.67822 | 1020526.575 | 78.528 | 189.5833
0.6949 | 1041667.7481 | 79.8444 | 192.7614
0.7043 | 1087387.1751 | 75.37 | 181.9592
0.69644 | 1047237.4399 | 76.823 | 185.4672
0.70045 | 1066364.3973 | 81.0899 | 195.7682
0.708 | 1086133.057 | 78.3759 | 189.2161
0.67279 | 997111.6918 | 72.3501 | 174.6686
0.71237 | 1106396.0085 | 61.905 | 149.4518
0.67123 | 1005061.8327 | 67.6731 | 163.3774
0.68242 | 1044211.0447 | 71.0512 | 171.5327
0.67133 | 1022351.6657 | 72.5503 | 175.1518
0.66786 | 997022.152 | 85.6431 | 206.7608
0.70119 | 1059883.9138 | 66.3899 | 160.2795
0.68867 | 1035809.6636 | 73.4244 | 177.2622
0.69766 | 1058973.4651 | 92.8915 | 224.2599
0.69432 | 1045643.8863 | 75.2488 | 181.6667
0.68271 | 1015720.0725 | 65.3462 | 157.7598
0.69376 | 1038750.054 | 85.896 | 207.3713
0.68779 | 1037346.5796 | 70.0339 | 169.0767
0.7031 | 1074700.4177 | 60.9236 | 147.0827
0.69241 | 1042583.0607 | 71.734 | 173.1813
0.69481 | 1057400.3385 | 68.7808 | 166.0516
0.68667 | 1041779.334 | 80.8101 | 195.0929
0.70491 | 1076735.6591 | 67.0898 | 161.9692
0.69821 | 1062208.596 | 81.0332 | 195.6315
0.6893 | 1027978.2854 | 57.2796 | 138.2852
0.67555 | 1025201.4796 | 85.5472 | 206.5293
0.6898 | 1032184.3486 | 57.4726 | 138.7512
0.69475 | 1049764.3925 | 85.1756 | 205.6321
0.69708 | 1065825.2652 | 63.0819 | 152.2932
0.7149 | 1106852.0751 | 84.1016 | 203.0393
0.66861 | 999140.5943 | 79.5959 | 192.1615
0.6774 | 1003474.8789 | 74.0183 | 178.696
0.66869 | 991455.0954 | 56.9475 | 137.4833
0.66721 | 994337.1537 | 74.4642 | 179.7726
0.68605 | 1028773.2179 | 74.9273 | 180.8905
0.66228 | 956644.4713 | 70.5156 | 170.2397
0.69614 | 1042824.8904 | 74.8501 | 180.7041
0.68712 | 1037151.3149 | 88.0742 | 212.6299
0.7004 | 1061594.4531 | 94.1531 | 227.3056
0.70714 | 1088780.6415 | 75.1636 | 181.4611
0.6951 | 1057868.6553 | 62.7896 | 151.5876
0.6964 | 1059459.6213 | 64.8249 | 156.5011
0.67817 | 1015964.0658 | 78.1829 | 188.7503
0.6871 | 1032726.0446 | 89.1507 | 215.2289
0.68479 | 1022260.3286 | 71.1185 | 171.6953
0.69151 | 1048085.1908 | 62.5317 | 150.9649
0.67027 | 986437.1555 | 68.3372 | 164.9807
0.70141 | 1072516.2664 | 72.8468 | 175.8677
0.67087 | 1000930.0676 | 92.6799 | 223.7491
0.68601 | 1042690.1666 | 77.652 | 187.4684
0.71256 | 1108371.1303 | 66.2793 | 160.0124
0.69237 | 1053004.5163 | 70.62 | 170.4918
0.67528 | 1006133.0631 | 59.387 | 143.3729
0.68986 | 1036886.2727 | 76.7889 | 185.3848
0.66559 | 983631.536 | 75.502 | 182.278
0.68245 | 1036403.0282 | 60.5309 | 146.1346
0.69751 | 1057415.431 | 70.3937 | 169.9455
0.67191 | 1012432.2289 | 83.408 | 201.3647
0.68804 | 1039757.7848 | 80.4927 | 194.3266
0.70859 | 1100866.3073 | 77.6274 | 187.4092
0.70659 | 1077503.1728 | 74.4431 | 179.7215
0.68893 | 1032119.7273 | 91.6832 | 221.3429
0.72099 | 1107098.2192 | 80.1367 | 193.4671
0.6842 | 1040696.9884 | 87.2602 | 210.6648
0.66921 | 992885.3477 | 61.2121 | 147.7791
0.68533 | 1036884.8968 | 66.4524 | 160.4302
0.67116 | 995254.1279 | 56.9554 | 137.5025
0.68707 | 1035562.9846 | 81.8721 | 197.6567
0.67956 | 1002238.7909 | 75.3488 | 181.9081
0.6884 | 1040778.1467 | 66.4053 | 160.3167
0.71404 | 1111217.1151 | 90.806 | 219.2251
0.68024 | 998584.4886 | 84.2169 | 203.3177
0.69662 | 1053913.7429 | 66.4496 | 160.4235
0.70442 | 1069930.9073 | 77.5987 | 187.3397
0.69536 | 1060352.2924 | 77.9065 | 188.0829
0.67103 | 998938.4272 | 50.1468 | 121.065
0.70037 | 1074880.0383 | 78.408 | 189.2937
0.68953 | 1012120.584 | 64.4704 | 155.6453
0.67228 | 998168.9181 | 82.1882 | 198.4199
0.69225 | 1029974.6886 | 80.8201 | 195.1169
0.66665 | 986183.2049 | 81.8864 | 197.6912
0.68267 | 1010284.3404 | 77.234 | 186.4593
0.68116 | 1023910.6166 | 66.023 | 159.3936
0.69557 | 1042483.5955 | 69.5932 | 168.0129
0.69329 | 1040049.5981 | 76.6424 | 185.0312
0.69628 | 1068077.226 | 71.2292 | 171.9626
0.68876 | 1037095.4984 | 76.9273 | 185.7189
0.70649 | 1082159.3273 | 71.1248 | 171.7105
0.69461 | 1043090.0771 | 65.9571 | 159.2346
0.67192 | 1005603.8721 | 73.0608 | 176.3844
0.69952 | 1068819.6125 | 78.0608 | 188.4554
0.71519 | 1108077.4225 | 75.6348 | 182.5986
0.68202 | 1010569.6459 | 75.7003 | 182.7567
0.68918 | 1040130.8593 | 71.9008 | 173.5838
0.64376 | 949986.4758 | 77.0136 | 185.9272
0.68711 | 1027458.808 | 76.3951 | 184.4341
0.6832 | 1034661.7183 | 75.67 | 182.6836
0.71962 | 1094959.2063 | 72.489 | 175.0039
0.70376 | 1083227.1875 | 75.1399 | 181.4039
0.67716 | 1013108.3769 | 59.6325 | 143.9656
0.6749 | 1001740.9217 | 75.8307 | 183.0716
0.69246 | 1050844.4997 | 66.6544 | 160.9179
0.66514 | 983137.6289 | 80.2896 | 193.8362
0.68995 | 1044716.2044 | 80.6525 | 194.7123
0.67114 | 990848.6736 | 72.0727 | 173.9988
0.69737 | 1054897.0788 | 82.1679 | 198.3709
0.6806 | 1010266.3102 | 75.8354 | 183.0829
0.68494 | 1014496.9437 | 63.1126 | 152.3673